## Collocation Data
National Diffusion Tube Bias Adjustment Factor Spreadsheet | [7 empty] | Spreadsheet Version Number: 03/23
Follow the steps below in the correct order to show the results of relevant co-location studies | [8 empty] | This spreadsheet will be updated at the end of June 2023
Data only apply to tubes exposed monthly and are not suitable for correcting individual short-term monitoring periods
Whenever presenting adjusted data, you should state the adjustment factor used and the version of the spreadsheet
This spreadhseet will be updated every few months: the factors may therefore be subject to change. This should not discourage their immediate use. | [8 empty] | LAQM Helpdesk Website
The LAQM Helpdesk is operated on behalf of Defra and the Devolved Administrations by Bureau Veritas, in conjunction with contract partners AECOM and the National Physical Laboratory. | [5 empty] | Spreadsheet maintained by the National Physical Laboratory. Original compiled by Air Quality Consultants Ltd.
Step 1: | Step 2: | Step 3: | Step 4:
Select the Laboratory that Analyses Your Tubes from the Drop-Down List | Select a Preparation Method from the Drop-Down List | Select a Year from the Drop-Down List | Where there is only one study for a chosen combination, you should use the adjustment factor shown with caution. Where there is more than one study, use the overall factor3 shown in blue at the foot of the final column.
If a laboratory is not shown, we have no data for this laboratory. | If a preparation method is not shown, we have no data for this method at this laboratory. | If a year is not shown, we have no data2 | If you have your own co-location study then see footnote4. If uncertain what to do then contact the Local Air Quality Management Helpdesk at LAQMHelpdesk@bureauveritas.com or 0800 0327953
Analysed By1 | Method To undo your selection, choose (All) from the pop-up list | Year5 To undo your selection, choose (All) | Site Type | Local Authority |  | Length of Study (months) | Diffusion Tube Mean Conc. (Dm) (mg/m3) | Automatic Monitor Mean Conc. (Cm) (mg/m3) | Bias (B) | Tube Precision6 | Bias Adjustment Factor (A) (Cm/Dm)
Aberdeen CC | 50% TEA in Acetone | 2001 | R | Aberdeen CC |  | 12 | 45.6462 | 61.2385 | -25.46% | na | 1.3416
Aberdeen CC | 50% TEA in Acetone | 2001 | R | Aberdeen CC |  | 12 | 40 | 53.1917 | -24.8% | na | 1.3298
Gradko | 50% TEA in Acetone | 2001 | UB | RB Kensington and Chelsea |  | 11 | 31.4524 | 42.1184 | -25.32% | na | 1.3391
Gradko | 50% TEA in Acetone | 2001 | Su | LB Hillingdon |  | 12 | 28.926 | 46.2756 | -37.49% | na | 1.5998
Gradko | 50% TEA in Acetone | 2001 | UB | LB Brent |  | 10 | 28.9886 | 36.9286 | -21.5% | na | 1.2739
Gradko | 50% TEA in Water | 2000 | UB | Birmingham CC |  | 12 | 29.1417 | 33.3417 | -12.6% | na | 1.1441
Gradko | 50% TEA in Water | 2001 | UB | Birmingham CC |  | 12 | 29.1917 | 36.425 | -19.86% | na | 1.2478
Gradko | 50% TEA in Water | 2001 | Int | New Forest DC |  | 11 | 20.9291 | 26.0318 | -19.6% | na | 1.2438
Gradko | 50% TEA in Water | 2001 | Su | New Forest DC |  | 11 | 13.4545 | 23.2727 | -42.19% | na | 1.7297
Gradko | 50% TEA in Water | 2000 | UC | Taunton Deane BC |  | 12 | 18.3333 | 28.5 | -35.67% | na | 1.5545
Gradko | 50% TEA in Water | 2001 | UB | Thurrock BC |  | 12 | 25.2789 | 35.267 | -28.32% | na | 1.3951
Gradko | 50% TEA in Water | 2001 | R | Dartford BC |  | 9 | 52.1378 | 65.8244 | -20.79% | na | 1.2625
Environmental Scientific Groups | 10% TEA in Water | 2001 | B | Derby CC |  | 12 | 35.1667 | 37.25 | -5.593% | na | 1.0592
Edinburgh Scientific Services | 50% TEA in Acetone | 2001 | R | Edinburgh CC |  | 12 | 43.0833 | 38.9583 | 10.59% | na | 0.90426
Edinburgh Scientific Services | 50% TEA in Acetone | 2001 | R | Edinburgh CC |  | 12 | 43.75 | 41.825 | 4.603% | na | 0.956
Harwell Scientific Services | 50% TEA in Acetone | 2001 | Int | Blackburn with Darwen B C |  | 9 | 31.1967 | 30.6667 | 1.728% | na | 0.98301
Harwell Scientific Services | 50% TEA in Acetone | 2001 | R | St Edmundsbury BC |  | 10 | 53.0475 | 38.0687 | 39.35% | na | 0.71763
Gradko | 50% TEA in Acetone | 2000 | R | LB Croydon |  | 9 | 35.1322 | 33.8383 | 3.824% | na | 0.96317
Gradko | 50% TEA in Acetone | 2000 | R | LB Croydon |  | 11 | 34.5127 | 52.3573 | -34.08% | na | 1.517
Gradko | 20% TEA in Water | 2002 | R | Exeter CC |  | 12 | 30.75 | 37.9 | -18.87% | na | 1.2325
Gradko | 50% TEA in Acetone | 2002 | B | Sandwell MBC |  | 9.033 | 25.64 | 21.18 | 21.06% | na | 0.82605
Gradko | 50% TEA in Water | 2003 | R | Taunton Deane BC |  | 9.8022 | 21.71 | 24.95 | -12.99% | na | 1.1492
West Yorkshire Analytical Services | 50% TEA in Acetone | 2002 | UB | Wakefield MDC |  | 10.033 | 34.83 | 29.8 | 16.88% | na | 0.85558
Harwell Scientific Services | 50% TEA in Acetone | 2002 | R | Hambleton DC |  | 12 | 31.1 | 26.4 | 17.8% | na | 0.84887
Gradko | 50% TEA in Acetone | 2002 | R | North West Leicestershire DC |  | 12 | 33.36 | 37.25 | -10.44% | na | 1.1166
Walsall MBC | 50% TEA in Acetone | 2002 | R | Gedling BC |  | 12 | 39 | 34 | 14.71% | na | 0.87179
Walsall MBC | 50% TEA in Acetone | 2001 | R | Gedling BC |  | 9.978 | 36 | 51 | -29.41% | na | 1.4167
Gradko | 50% TEA in Acetone | 2001 | UB | LB Camden |  | 12 | 40.75 | 50.45 | -19.23% | na | 1.238
Gradko | 50% TEA in Acetone | 2002 | K | LB Camden |  | 10 | 36.5 | 65.9 | -44.61% | na | 1.8055
Gradko | 50% TEA in Acetone | 2001 | K | LB Camden |  | 11 | 40.86 | 65.74 | -37.85% | na | 1.6089
Environmental Scientific Groups | 10% TEA in Water | 2002 | R | South Buckinghamshire DC |  | 10 | 31.7 | 28.9 | 9.689% | na | 0.91167
Jesmond Dene Laboratory | 20% TEA in Water | 2000 | UC | Newcastle upon Tyne CC |  | 9 | 28.3 | 30.5 | -7.213% | na | 1.0777
Jesmond Dene Laboratory | 20% TEA in Water | 2001 | R | Newcastle upon Tyne CC |  | 11.033 | 29.5 | 30.9 | -4.531% | na | 1.0475
Jesmond Dene Laboratory | 50% TEA in Acetone | 2002 | R | Newcastle upon Tyne CC |  | 11.967 | 33.4 | 30 | 11.33% | na | 0.8982
Tayside SS | 20% TEA in Water | 2002 | R | Dundee CC |  | 12 | 52.1 | 42.6 | 22.3% | na | 0.81766
Harwell Scientific Services | 50% TEA in Acetone | 2002 | R | Suffolk Coastal DC |  | 10.6374 | 36.7 | 32.5 | 12.92% | na | 0.88556
Gradko | 20% TEA in Water | 2002 | R | Southampton CC |  | 12 | 39.5417 | 27.26 | 45.05% | na | 0.6894
Gradko | 20% TEA in Water | 2002 | R | Southampton CC |  | 12 | 35.0653 | 36.7105 | -4.482% | na | 1.0469
Gradko | 20% TEA in Water | 2002 | UC | Southampton CC |  | 10 | 33.8601 | 29.8428 | 13.46% | na | 0.88136
University of Essex | 50% TEA in Acetone | 2002 | Su | Colchester BC |  | 10.033 | 37 | 32.88 | 12.53% | na | 0.88865
Gradko | 20% TEA in Water | 2002 | R | Dartford BC |  | 11.6703 | 52.2249 | 61.216 | -14.69% | na | 1.1722
Gradko | 20% TEA in Water | 2002 | UB | LB Ealing |  | 11 | 23.7 | 28.2 | -15.96% | na | 1.1899
Gradko | 20% TEA in Water | 2002 | R | LB Ealing |  | 11 | 45.2 | 52.3 | -13.58% | na | 1.1571
Jesmond Dene Laboratory | 20% TEA in Water | 2000 | K | Gateshead Council |  | 12 | 28.9 | 31.1 | -7.074% | na | 1.0761
Jesmond Dene Laboratory | 50% TEA in Acetone | 2002 | K | Gateshead Council |  | 9.956 | 34.2 | 27.1 | 26.2% | na | 0.7924
Kent Scientific Services | 50% TEA in Acetone | 2001 | R | Eastleigh BC |  | 12.033 | 38.8 | 34.4 | 12.79% | na | 0.8866
Kent Scientific Services | 50% TEA in Acetone | 2002 | R | Eastleigh BC |  | 13.044 | 36.68 | 34.38 | 6.69% | na | 0.9373
Gradko | 20% TEA in Water | 2002 | K | South Lakeland DC |  | 12 | 35.76 | 35.95 | -0.5285% | na | 1.0053
Gradko | 50% TEA in Acetone | 2001 | Su | Slough BC |  | 12.033 | 21.5 | 36.2 | -40.61% | na | 1.6837
Gradko | 50% TEA in Acetone | 2002 | Su | Slough BC |  | 11.967 | 24.8 | 35.8 | -30.73% | na | 1.4435
Gradko | 50% TEA in Acetone | 2001 | UC | Slough BC |  | 12.033 | 23.1 | 38.4 | -39.84% | na | 1.6623
Gradko | 50% TEA in Acetone | 2002 | UC | Slough BC |  | 10.967 | 23.5 | 33.2 | -29.22% | na | 1.4128
Harwell Scientific Services | 50% TEA in Acetone | 2000 | K | Adur DC |  | >9 | 41.5121 | 36.347 | 14.21% | na | 0.87558
Harwell Scientific Services | 50% TEA in Acetone | 2001 | K | Adur DC |  | >9 | 47.4424 | 38.26 | 24% | na | 0.80645
Harwell Scientific Services | 50% TEA in Acetone | 2002 | K | Adur DC |  | >9 | 42.086 | 34.434 | 22.22% | na | 0.81818
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | K | Belfast CC |  | 12 | 49.71 | 50.17 | -0.9169% | na | 1.0093
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | UC | Belfast CC |  | 12 | 28.76 | 30.47 | -5.612% | na | 1.0595
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | K | Belfast CC |  | 11 | 37.71 | 58.57 | -35.62% | na | 1.5532
Bristol Scientific Services | 20% TEA in Water | 2002 | UB | LB Waltham Forest |  | 12 | 35 | 38 | -7.895% | na | 1.0857
Environmental Scientific Groups | 10% TEA in Water | 2002 | R | East Lothian Council |  | 8.8901 | 48.5 | 33.7 | 43.92% | na | 0.69485
Gradko | 20% TEA in Water | 2002 | R | Coventry CC |  | 11.9341 | 35.66 | 41.58 | -14.24% | na | 1.166
Gradko | 20% TEA in Water | 2002 | R | Coventry CC |  | 10.9341 | 41.86 | 35.19 | 18.95% | na | 0.84066
Gradko | 20% TEA in Water | 2002 | UB | Coventry CC |  | 10.9341 | 21.62 | 14.35 | 50.66% | na | 0.66374
Gradko | 50% TEA in Acetone | 2002 | R | LB Richmond |  | 12 | 30.55 | 43.999 | -30.57% | na | 1.4402
Glasgow Scientific Services | 20% TEA in Water | 2002 | K | Glasgow CC |  | 12 | 88.1 | 74.5 | 18.26% | na | 0.84563
Glasgow Scientific Services | 20% TEA in Water | 2002 | UB | Glasgow CC |  | 11 | 40 | 32.7 | 22.32% | na | 0.8175
Glasgow Scientific Services | 20% TEA in Water | 2002 | R | Glasgow CC |  | 12 | 41.7 | 47.1 | -11.46% | na | 1.1295
Gradko | 20% TEA in Water | 2002 | R | Ellesmere Port & Neston BC |  | 9.9341 | 35.55 | 36.39 | -2.308% | na | 1.0236
Lambeth Scientific Services | 50% TEA in Acetone | 2000 | B | York CC |  | 11 | 14.24 | 19.25 | -26.03% | na | 1.3518
Lambeth Scientific Services | 50% TEA in Acetone | 2001 | B | York CC |  | 10.0989 | 15.85 | 14.66 | 8.117% | na | 0.92492
Lambeth Scientific Services | 50% TEA in Acetone | 2000 | R | York CC |  | 10 | 37.82 | 33.67 | 12.33% | na | 0.89027
Lambeth Scientific Services | 50% TEA in Acetone | 2001 | R | York CC |  | 10.0989 | 33.23 | 33.29 | -0.1802% | na | 1.0018
Lambeth Scientific Services | 50% TEA in Acetone | 2000 | R | York CC |  | 9 | 36.29 | 27.78 | 30.63% | na | 0.7655
Lambeth Scientific Services | 50% TEA in Acetone | 2001 | R | York CC |  | 12.0989 | 22.6 | 37.74 | -40.12% | na | 1.6699
Gradko | 20% TEA in Water | 2003 | R | Dudley MBC |  | 9 | 52.1256 | 51.8292 | 0.5718% | na | 0.99431
Gradko | 20% TEA in Water | 2002 | UB | Dudley MBC |  | 11.044 | 26.9733 | 29.6493 | -9.025% | na | 1.0992
Kent Scientific Services | 50% TEA in Acetone | 2002 | R | Tonbridge & Malling BC |  | 11.9341 | 37.26 | 42.59 | -12.51% | na | 1.143
Kent Scientific Services | 50% TEA in Acetone | 2001 | R | Tonbridge & Malling BC |  | 9.0659 | 42.43 | 50.15 | -15.39% | na | 1.1819
Environmental Scientific Groups | 10% TEA in Water | 2002 | R | Bury MBC |  | 12.8901 | 56 | 68 | -17.65% | na | 1.2143
Environmental Scientific Groups | 10% TEA in Water | 2001 | R | Kingston upon Hull CC |  | 9.1099 | 19 | 19 | 0% | na | 1
South Yorkshire Labs | 50% TEA in Acetone | 2002 | R | Kingston upon Hull CC |  | 8.967 | 40 | 40 | 0% | na | 1
South Yorkshire Labs | 50% TEA in Acetone | 2002 | R | Kingston upon Hull CC |  | 17.1978 | 33 | 29 | 13.79% | na | 0.87879
Gradko | 50% TEA in Acetone | 2002 | K | LB Hounslow |  | 12 | 39.7761 | 57 | -30.22% | na | 1.433
Gradko | 50% TEA in Acetone | 2002 | R | LB Hounslow |  | 12 | 31.7694 | 58.9 | -46.06% | na | 1.854
Gradko | 50% TEA in Acetone | 2002 | B | LB Hounslow |  | 12 | 26.9056 | 42.7 | -36.99% | na | 1.587
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | R | Waverly BC |  | 12.033 | 31.1819 | 37.227 | -16.24% | na | 1.1939
Gradko | 20% TEA in Water | 2002 | R | Rushmoor BC |  | 12 | 36.46 | 41.9 | -12.98% | na | 1.1492
Gradko | 50% TEA in Acetone | 2002 | R | Wolverhampton CC |  | 12 | 37 | 34 | 8.824% | na | 0.91892
Gradko | 50% TEA in Acetone | 2001 | UB | Wolverhampton CC |  | 9.033 | 22 | 28 | -21.43% | na | 1.2727
Gradko | 50% TEA in Acetone | 2002 | UB | Wolverhampton CC |  | 12 | 27 | 29 | -6.897% | na | 1.0741
Environmental Scientific Groups | 10% TEA in Water | 2002 | B | Basingstoke and Deane BC |  | 11.011 | 36.2 | 30.59 | 18.34% | na | 0.84503
Kent Scientific Services | 50% TEA in Acetone | 2002 | R | Dover DC |  | 9.0989 | 58.6 | 40.53 | 44.58% | na | 0.69164
South Yorkshire Labs | 50% TEA in Acetone | 2003 | R | Barnsley MBC |  | 11.044 | 44.5 | 38.2 | 16.49% | na | 0.85843
Lambeth Scientific Services | 50% TEA in Acetone | 2001 | K | North Hertfordshire DC |  | 11.011 | 37.5644 | 36.1731 | 3.846% | na | 0.96296
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | K | North Hertfordshire DC |  | 10.033 | 23.9125 | 34.434 | -30.56% | na | 1.44
Gradko | 50% TEA in Acetone | 2002 | K | Boston BC |  | 13.8462 | 44.77 | 44.68 | 0.2014% | na | 0.99799
Edinburgh Scientific Services | 50% TEA in Acetone | 2002 | R | West Lothian Council |  | 12 | 18 | 22 | -18.18% | na | 1.2222
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | Su | Reigate and Banstead DC |  | 10.6593 | 29.3636 | 32.1228 | -8.589% | na | 1.094
South Yorkshire Labs | 50% TEA in Acetone | 2002 | UB | Bournemouth BC |  | 9.967 | 21.64 | 17.27 | 25.3% | na | 0.79806
Harwell Scientific Services | 50% TEA in Acetone | 2000 | R | Cambridge CC |  | 12 | 53.7 | 38.6 | 39.12% | na | 0.71881
Harwell Scientific Services | 50% TEA in Acetone | 2001 | R | Cambridge CC |  | 12 | 52 | 37.9 | 37.2% | na | 0.72885
Harwell Scientific Services | 50% TEA in Acetone | 2002 | R | Cambridge CC |  | 12 | 48.5 | 40.1 | 20.95% | na | 0.8268
Kent Scientific Services | 50% TEA in Acetone | 2003 | B | Medway Council |  | 12 | 35.63 | 22.0383 | 61.67% | na | 0.61853
Kent Scientific Services | 50% TEA in Acetone | 2003 | UB | Medway Council |  | 12 | 30.29 | 25.1 | 20.68% | na | 0.82866
Kent Scientific Services | 50% TEA in Acetone | 2003 | K | Medway Council |  | 12 | 48.1242 | 35.6342 | 35.05% | na | 0.74046
South Yorkshire Labs | 50% TEA in Acetone | 2003 | UB | Rotherham MBC |  | 11.1099 | 33.2 | 29.9 | 11.04% | na | 0.9006
Harwell Scientific Services | 50% TEA in Acetone | 2001 | R | South Cambridgeshire DC |  | 10 | 49 | 35.2 | 39.2% | na | 0.71837
Gradko | 50% TEA in Water | 2003 | R | Mid Devon DC |  | 9.2308 | 50.3 | 59 | -14.75% | na | 1.173
Kirklees Council Scientific Services | 50% TEA in Acetone | 2000 | UC | Harrogate |  | 11 | 24.4 | 23.8 | 2.521% | na | 0.97541
Jesmond Dene Laboratory | 50% TEA in Acetone | 2003 | R | Newcastle upon Tyne CC |  | 12 | 31.8592 | 31.7558 | 0.3255% | na | 0.99676
Gradko | 20% TEA in Water | 2003 | R | Exeter CC |  | 12 | 39.25 | 41.05 | -4.385% | na | 1.0459
Staffordshire Scientific Services | 50% TEA in Water | 2003 | R | South Staffordshire DC |  | 10.5385 | 50.1 | 58.7 | -14.65% | na | 1.1717
Walsall MBC | 50% TEA in Acetone | 2003 | R | Gedling BC |  | 12 | 36 | 39 | -7.692% | na | 1.0833
Kent Scientific Services | 50% TEA in Acetone | 2003 | R | Tonbridge & Malling BC |  | 10.2308 | 40.89 | 48.02 | -14.85% | na | 1.1744
Gradko | 20% TEA in Water | 2003 | B | Charnwood BC |  | 12.033 | 31.9 | 33.4 | -4.491% | na | 1.047
Gradko | 20% TEA in Water | 2003 | UB | LB Ealing |  | 9 | 45.0917 | 40.7467 | 10.66% | na | 0.90364
Gradko | 20% TEA in Water | 2003 | R | LB Ealing |  | 10 | 57.8846 | 59.401 | -2.553% | na | 1.0262
Jesmond Dene Laboratory | 50% TEA in Acetone | 2003 | K | Gateshead Council |  | 9.9231 | 32.47 | 29.22 | 11.12% | na | 0.89991
Harwell Scientific Services | 50% TEA in Acetone | 2003 | Ind | Gravesham BC |  | 10.9341 | 42.9 | 34.83 | 23.17% | na | 0.81189
Harwell Scientific Services | 50% TEA in Acetone | 2003 | R | Gravesham BC |  | 11.9341 | 57.4 | 53.1 | 8.098% | na | 0.92509
Gradko | 20% TEA in Water | 2003 | K | North Warwickshire BC |  | 9.9451 | 49.7 | 47.25 | 5.185% | na | 0.9507
Gradko | 50% TEA in Acetone | 2003 | R | North West Leicestershire DC |  | 9 | 35.84 | 41.28 | -13.18% | na | 1.1518
Gradko | 20% TEA in Water | 2003 | K | South Lakeland DC |  | 11.9011 | 39.43 | 35.28 | 11.76% | na | 0.89475
West Yorkshire Analytical Services | 50% TEA in Acetone | 2003 | UB | Wakefield MDC |  | 10.967 | 34.33 | 34.55 | -0.6368% | na | 1.0064
West Yorkshire Analytical Services | 50% TEA in Acetone | 2003 | UC | Wakefield MDC |  | 11 | 46.82 | 44.28 | 5.736% | na | 0.94575
Gradko | 20% TEA in Water | 2003 | R | Bromsgrove DC |  | 10.0879 | 48.84 | 51.6 | -5.349% | na | 1.0565
Tayside SS | 20% TEA in Water | 2003 | R | Fife Council |  | 11 | 46.3333 | 37.7273 | 22.81% | na | 0.81426
Gradko | 50% TEA in Acetone | 2003 | K | LB Camden |  | 12 | 58.5 | 64.8 | -9.722% | na | 1.1077
West Yorkshire Analytical Services | 50% TEA in Acetone | 2003 | UC | Leeds CC |  | 10.2967 | 51.1 | 44.1 | 15.87% | na | 0.86301
West Yorkshire Analytical Services | 50% TEA in Acetone | 2002 | UC | Leeds CC |  | 10.967 | 47.2 | 39.1 | 20.72% | na | 0.82839
Kent Scientific Services | 50% TEA in Acetone | 2003 | R | Eastleigh BC |  | 10.011 | 38.69 | 42.02 | -7.925% | na | 1.0861
Environmental Scientific Groups | 10% TEA in Water | 2003 | UC | Manchester CC |  | 10 | 40.2 | 42.722 | -5.903% | na | 1.0627
Environmental Scientific Groups | 10% TEA in Water | 2003 | Su | Manchester CC |  | 11 | 27.7879 | 21.6464 | 28.37% | na | 0.77899
Environmental Scientific Groups | 10% TEA in Water | 2003 | UB | Manchester CC |  | 11 | 34.5758 | 44.1991 | -21.77% | na | 1.2783
West Yorkshire Analytical Services | 50% TEA in Acetone | 2003 | B | Leeds CC |  | 11.9341 | 32.9 | 36.8 | -10.6% | na | 1.1185
West Yorkshire Analytical Services | 50% TEA in Acetone | 2002 | B | Leeds CC |  | 12 | 32.9 | 29.6 | 11.15% | na | 0.8997
Harwell Scientific Services | 50% TEA in Acetone | 2003 | R | Babergh DC |  | 11 | 34.7 | 28.8 | 20.49% | na | 0.82997
Environmental Scientific Groups | 10% TEA in Water | 2003 | UB | Wiiral MBC |  | 12 | 26 | 27 | -3.704% | na | 1.0385
Gradko | 50% TEA in Acetone | 2003 | UB | Casella Stanger |  | 11 | 33.55 | 34.8 | -3.592% | na | 1.0373
Gradko | 50% TEA in Acetone | 2003 | Su | Casella Stanger |  | 12 | 38.51 | 44.7 | -13.85% | na | 1.1607
Gradko | 50% TEA in Acetone | 2003 | Su | Slough BC |  | 11.9341 | 32.2 | 42 | -23.33% | na | 1.3043
Gradko | 50% TEA in Acetone | 2003 | UC | Slough BC |  | 8.9341 | 34.3 | 38 | -9.737% | na | 1.1079
South Yorkshire Labs | 50% TEA in Acetone | 2000 | R | Oxford CC |  | 11 | 58.8 | 60.8 | -3.289% | na | 1.034
South Yorkshire Labs | 50% TEA in Acetone | 2001 | R | Oxford CC |  | 9 | 60.5 | 59.6 | 1.51% | na | 0.98512
South Yorkshire Labs | 50% TEA in Acetone | 2002 | R | Oxford CC |  | 12 | 59.6 | 59.6 | 0% | na | 1
South Yorkshire Labs | 50% TEA in Acetone | 2003 | R | Oxford CC |  | 12 | 69.4 | 71.5 | -2.937% | na | 1.0303
South Yorkshire Labs | 50% TEA in Acetone | 2000 | B | Oxford CC |  | 10 | 27.5 | 19.5 | 41.03% | na | 0.70909
South Yorkshire Labs | 50% TEA in Acetone | 2001 | B | Oxford CC |  | 11 | 29.2 | 22.5 | 29.78% | na | 0.77055
South Yorkshire Labs | 50% TEA in Acetone | 2002 | B | Oxford CC |  | 12 | 25.1 | 21.4 | 17.29% | na | 0.85259
South Yorkshire Labs | 50% TEA in Acetone | 2003 | B | Oxford CC |  | 12 | 26.8 | 24.9 | 7.631% | na | 0.9291
Environmental Scientific Groups | 10% TEA in Water | 2003 | UB | Flintshire CC |  | 10.022 | 20 | 20 | 0% | na | 1
Gradko | 20% TEA in Water | 2003 | R | Mid Beds DC |  | 12.1648 | 38.73 | 38.2 | 1.387% | na | 0.98632
Lambeth Scientific Services | 50% TEA in Acetone | 2003 | UB | Derry CC |  | 12 | 15.3 | 16.25 | -5.846% | na | 1.0621
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | UB | Derry CC |  | 8.9011 | 14.6 | 14.5 | 0.6897% | na | 0.99315
Lambeth Scientific Services | 50% TEA in Acetone | 2002 | B | East Hertfordshire DC |  | 8.967 | 22.7 | 23.4 | -2.991% | na | 1.0308
Lambeth Scientific Services | 50% TEA in Acetone | 2003 | B | East Hertfordshire DC |  | 10.967 | 23.4 | 24.4 | -4.098% | na | 1.0427
Edinburgh Scientific Services | 50% TEA in Acetone | 2003 | R | Edinburgh CC |  | 12.033 | 45.1 | 41.1 | 9.732% | na | 0.91131
Gradko | 50% TEA in Water | 2002 | R | Winchester CC |  | 13.1209 | 35.577 | 47 | -24.46% | na | 1.3239
Gradko | 50% TEA in Water | 2003 | UB | New Forest DC |  | 11.011 | 19.68 | 20.1 | -2.09% | na | 1.0213
Gradko | 50% TEA in Water | 2003 | UB | New Forest DC |  | 11.011 | 26.2 | 25.9 | 1.158% | na | 0.98855
West Yorkshire Analytical Services | 50% TEA in Acetone | 2001 | B | Leeds CC |  | 12 | 30.1 | 32.6 | -7.669% | na | 1.0831
West Yorkshire Analytical Services | 50% TEA in Acetone | 2001 | K | Leeds CC |  | 10 | 54.8 | 52.5 | 4.381% | na | 0.95803
West Yorkshire Analytical Services | 50% TEA in Acetone | 2002 | K | Leeds CC |  | 11.9341 | 54.2 | 47.4 | 14.35% | na | 0.87454
West Yorkshire Analytical Services | 50% TEA in Acetone | 2003 | K | Leeds CC |  | 11.033 | 56.5 | 52.5 | 7.619% | na | 0.9292
Harwell Scientific Services | 50% TEA in Acetone | 2003 | R | South Cambridgeshire DC |  | 12 | 52.24 | 48.43 | 7.867% | na | 0.92707
Harwell Scientific Services | 50% TEA in Acetone | 2003 | R | Hambleton DC |  | 11 | 30.5 | 25.9 | 17.76% | na | 0.84918
Staffordshire Scientific Services | 50% TEA in Water | 2003 | R | East Staffordshire BC |  | 9.1319 | 33.3 | 36.8 | -9.511% | na | 1.1051
South Yorkshire Labs | 50% TEA in Acetone | 2003 | UB | Bournemouth BC |  | 10.967 | 23 | 21.58 | 6.58% | na | 0.93826
Gradko | 50% TEA in Water | 2002 | UB | Thurrock Council |  | 12 | 31.43 | 36 | -12.69% | na | 1.1454
Gradko | 50% TEA in Water | 2003 | UB | Thurrock Council |  | 12 | 34.16 | 38 | -10.11% | na | 1.1124
Gradko | 50% TEA in Water | 2003 | R | Thurrock Council |  | 10.022 | 57.57 | 75.3 | -23.55% | na | 1.308
Bristol Scientific Services | 20% TEA in Water | 2004 | K | South Gloucestershire Council |  | 10.1538 | 32.2 | 31.4 | 2.548% | na | 0.97516
Gradko | 50% TEA in Acetone | 2003 | R | LB Croydon |  | 11 | 74.5 | 74.3 | 0.2692% | na | 0.99732
Gradko | 50% TEA in Acetone | 2003 | R | LB Croydon |  | 11 | 62.04 | 55.6 | 11.58% | na | 0.8962
Gradko | 50% TEA in Acetone | 2003 | R | LB Croydon |  | 11 | 57.4 | 54.6 | 5.128% | na | 0.95122
Gradko | 50% TEA in Acetone | 2003 | Ind | LB Croydon |  | 11 | 44.98 | 38.3 | 17.44% | na | 0.85149
Gradko | 20% TEA in Water | 2003 | R | Highland Council |  | 8.967 | 27.12 | 24.35 | 11.38% | na | 0.89786
Jesmond Dene Laboratory | 50% TEA in Acetone | 2003 | K | Sttockton on Tees BC |  | 10 | 44 | 45.2 | -2.655% | na | 1.0273
Jesmond Dene Laboratory | 50% TEA in Acetone | 2003 | Ind | Sttockton on Tees BC |  | 12 | 28.4 | 31.6 | -10.13% | na | 1.1127
Environmental Scientific Groups | 10% TEA in Water | 2004 | R | East Lothian Council |  | 8.9231 | 43 | 32.8 | 31.1% | na | 0.76279
Bristol Scientific Services | 20% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 12 | 37.0556 | 30.7 | 20.7% | na | 0.82849
Environmental Scientific Groups | 10% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 11 | 33.4636 | 30.7 | 9.002% | na | 0.91741
Exova | 20% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 10 | 41.5297 | 30.7 | 35.28% | na | 0.73923
Edinburgh Scientific Services | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 12 | 30.2778 | 30.7 | -1.375% | na | 1.0139
Glasgow Scientific Services | 20% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 12 | 39.1947 | 30.7 | 27.67% | na | 0.78327
Gradko | 20% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 12 | 39.2033 | 30.7 | 27.7% | na | 0.7831
Gradko | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 12 | 35.3778 | 30.7 | 15.24% | na | 0.86778
Gradko | 50% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 12 | 34.5558 | 30.7 | 12.56% | na | 0.88842
Harwell Scientific Services | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 12 | 41.9556 | 30.7 | 36.66% | na | 0.73173
Jesmond Dene Laboratory | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 12 | 31.8394 | 30.7 | 3.712% | na | 0.96421
Kirklees Council Scientific Services | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 12 | 24.5836 | 30.7 | -19.92% | na | 1.2488
Milton Keynes Council | 20% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 11 | 36.5394 | 30.7 | 19.02% | na | 0.84019
Northampton BC | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 10 | 28.7827 | 30.7 | -6.245% | na | 1.0666
South Yorkshire Labs | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 12 | 40 | 30.7 | 30.29% | na | 0.7675
Staffordshire Scientific Services | 50% TEA in Water | 2003 | UB | AEA Tech intercomparison |  | 12 | 28.7389 | 30.7 | -6.388% | na | 1.0682
Walsall MBC | 50% TEA in Acetone | 2003 | UB | AEA Tech intercomparison |  | 10 | 36.8333 | 30.7 | 19.98% | na | 0.83348
Lambeth Scientific Services | 50% TEA in Acetone | 2004 | B | East Hertfordshire DC |  | 11.033 | 24.91 | 22.55 | 10.47% | na | 0.90526
Gradko | 50% TEA in Acetone | 2003 | UB | Sheffield CC |  | 9 | 35.6078 | 41.8522 | -14.92% | G | 1.1754
Gradko | 50% TEA in Acetone | 2003 | UC | Sheffield CC |  | 12 | 37.1744 | 37.5586 | -1.023% | P | 1.0103
Gradko | 50% TEA in Acetone | 2004 | UC | Sheffield CC |  | 12 | 29.7453 | 33.5732 | -11.4% | G | 1.1287
Gradko | 50% TEA in Acetone | 2004 | UB | Sheffield CC |  | 11 | 29.2214 | 42.399 | -31.08% | G | 1.451
Northampton BC | 20% TEA in Water | 2004 | B | Northampton BC |  | 9 | 19.94 | 20.3722 | -2.122% | P | 1.0217
Harwell Scientific Services | 50% TEA in Acetone | 2004 | R | Gravesham BC |  | 11 | 52.7424 | 53.5818 | -1.567% | G | 1.0159
Gradko | 20% TEA in Water | 2004 | R | LB Ealing |  | 10 | 54.9186 | 54.42 | 0.9162% | G | 0.99092
Gradko | 20% TEA in Water | 2004 | UB | LB Ealing |  | 12 | 43.711 | 41.25 | 5.966% | G | 0.9437
Gradko | 20% TEA in Water | 2004 | R | Exeter CC |  | 12 | 34.5592 | 39.2242 | -11.89% | G | 1.135
Gradko | 50% TEA in Acetone | 2004 | B | Sandwell MBC |  | 10 | 30.2258 | 29.384 | 2.865% | G | 0.97215
Gradko | 50% TEA in Acetone | 2003 | B | Sandwell MBC |  | 11 | 33.8688 | 37.0955 | -8.698% | P | 1.0953
West Yorkshire Analytical Services | 50% TEA in Acetone | 2004 | UB | Leeds CC |  | 12 | 24.875 | 26.0667 | -4.572% | G | 1.0479
Jesmond Dene Laboratory | 50% TEA in Acetone | 2004 | R | Gateshead Council |  | 11 | 35.8242 | 41.3627 | -13.39% | P | 1.1546
Staffordshire Scientific Services | 50% TEA in Water | 2004 | R | East Staffordshire BC |  | 11 | 31.7227 | 32.6436 | -2.821% | P | 1.029
Gradko | 20% TEA in Water | 2004 | UB | Dudley MBC |  | 11 | 29.3 | 25.9818 | 12.77% | G | 0.88675
Harwell Scientific Services | 50% TEA in Acetone | 2004 | Ind | Gravesham BC |  | 11 | 35.397 | 31.2455 | 13.29% | G | 0.88272
Harwell Scientific Services | 50% TEA in Acetone | 2004 | R | Adur DC |  | 11 | 40.8909 | 38.4 | 6.487% | S | 0.93908
Gradko | 20% TEA in Water | 2004 | R | Macclesfield BC |  | 12 | 32.6294 | 38.3592 | -14.94% | G | 1.1756
West Yorkshire Analytical Services | 50% TEA in Acetone | 2004 | UB | Leeds CC |  | 9 | 35.1111 | 32.7222 | 7.301% | S | 0.93196
West Yorkshire Analytical Services | 50% TEA in Acetone | 2004 | K | Leeds CC |  | 10 | 48 | 49.13 | -2.3% | S | 1.0235
West Yorkshire Analytical Services | 50% TEA in Acetone | 2004 | Int | Leeds CC |  | 10 | 40.25 | 45.01 | -10.58% | G | 1.1183
Gradko | 50% TEA in Acetone | 2003 | R | LB Hammersmith and Fulham |  | 9 | 83.9644 | 93.2222 | -9.931% | S | 1.1103
Gradko | 50% TEA in Acetone | 2001 | R | LB Hammersmith and Fulham |  | 11 | 44.7691 | 69.5727 | -35.65% | S | 1.554
Jesmond Dene Laboratory | 50% TEA in Acetone | 2004 | UC | Gateshead Council |  | 10 | 35.5433 | 36.354 | -2.23% | G | 1.0228
Gradko | 20% TEA in Water | 2004 | R | Dudley MBC |  | 12 | 46.0364 | 47.225 | -2.517% | G | 1.0258
Environmental Scientific Groups | 10% TEA in Water | 2004 | R | Wrexham CBC |  | 12 | 22.9444 | 20.375 | 12.61% | P | 0.88801
Environmental Scientific Groups | 10% TEA in Water | 2004 | UB | Warrington BC |  | 9 | 28.7407 | 25.7889 | 11.45% | G | 0.89729
Gradko | 20% TEA in Water | 2004 | R | South Lakeland DC |  | 12 | 36.4275 | 27.1842 | 34% | G | 0.74625
Gradko | 20% TEA in Water | 2004 | R | Rushmoor BC |  | 11 | 38.5379 | 30.2605 | 27.35% | P | 0.78521
Gradko | 50% TEA in Water | 2004 | R | Mid Devon DC |  | 11 | 47.1991 | 53.4718 | -11.73% | G | 1.1329
West Yorkshire Analytical Services | 50% TEA in Acetone | 2004 | UC | Leeds CC |  | 11 | 40.8788 | 33.6364 | 21.53% | P | 0.82283
Environmental Scientific Groups | 10% TEA in Water | 2004 | na | Stockport MBC |  | 10 | 47.3667 | 37.3 | 26.99% | P | 0.78747
Gradko | 50% TEA in Acetone | 2003 | UB | LB Hounslow |  | 11 | 36.2951 | 52.2182 | -30.49% | P | 1.4387
Gradko | 50% TEA in Acetone | 2003 | R | LB Hounslow |  | 9 | 64.3486 | 84.9222 | -24.23% | P | 1.3197
Gradko | 50% TEA in Acetone | 2004 | UC | Slough BC |  | 12 | 33.9333 | 34.4167 | -1.404% | G | 1.0142
Gradko | 50% TEA in Acetone | 2004 | Su | Slough BC |  | 12 | 28.2486 | 30.9167 | -8.63% | G | 1.0944
Walsall MBC | 50% TEA in Acetone | 2004 | R | Gedling BC |  | 12 | 32.8611 | 38.6667 | -15.01% | G | 1.1767
Gradko | 50% TEA in Acetone | 2004 | Su | Telford & Wrekin BC |  | 11 | 29.2985 | 38.3273 | -23.56% | G | 1.3082
Gradko | 50% TEA in Acetone | 2004 | R | LB Southwark |  | 12 | 55.8736 | 58.6167 | -4.68% | G | 1.0491
Edinburgh Scientific Services | 50% TEA in Acetone | 2004 | R | Edinburgh CC |  | 12 | 40.75 | 37.4167 | 8.909% | G | 0.9182
Edinburgh Scientific Services | 50% TEA in Acetone | 2004 | R | Edinburgh CC |  | 12 | 41.5417 | 36.8583 | 12.71% | G | 0.88726
Edinburgh Scientific Services | 50% TEA in Acetone | 2004 | B | Edinburgh CC |  | 10 | 11.2 | 10.69 | 4.771% | P | 0.95446
Gradko | 50% TEA in Water | 2004 | R | Thurrock Council |  | 12 | 43.8967 | 39.5 | 11.13% | P | 0.89984
Gradko | 50% TEA in Water | 2004 | UB | Thurrock Council |  | 12 | 35.0915 | 34 | 3.21% | S | 0.96889
Harwell Scientific Services | 50% TEA in Acetone | 2003 | Int | Three Rivers DC |  | 12 | 40.8107 | 41 | -0.4618% | S | 1.0046
Harwell Scientific Services | 50% TEA in Acetone | 2004 | Int | Three Rivers DC |  | 10 | 36.347 | 35.0534 | 3.69% | S | 0.96441
Walsall MBC | 50% TEA in Acetone | 2003 | B | Hertsmere DC |  | 12 | 33.5645 | 31 | 8.272% | S | 0.9236
Walsall MBC | 50% TEA in Acetone | 2004 | B | Hertsmere DC |  | 12 | 31.3036 | 26 | 20.4% | S | 0.83057
Gradko | 20% TEA in Water | 2003 | B | St Albans DC |  | 12 | 31.5114 | 30 | 5.038% | G | 0.95204
Gradko | 20% TEA in Water | 2004 | B | St Albans DC |  | 12 | 25.1878 | 25 | 0.7513% | G | 0.99254
Tayside SS | 20% TEA in Water | 2004 | R | Fife Council |  | 12 | 36.975 | 30.8667 | 19.79% | G | 0.8348
Lambeth Scientific Services | 50% TEA in Acetone | 2004 | R | LB Haringey |  | 11 | 36 | 46.5631 | -22.69% | S | 1.2934
Bristol Scientific Services | 20% TEA in Water | 2004 | R | Brighton and Hove CC |  | 12 | 37.0861 | 37.4172 | -0.8849% | P | 1.0089
Bristol Scientific Services | 20% TEA in Water | 2003 | R | Brighton and Hove CC |  | 12 | 40.1764 | 38.7797 | 3.602% | P | 0.96524
Cardiff Scientific Services | 50% TEA in Acetone | 2002 | UC | Cardiff CC |  | 11 | 34.0909 | 32.0909 | 6.232% | G | 0.94133
Cardiff Scientific Services | 50% TEA in Acetone | 2003 | UC | Cardiff CC |  | 11 | 40.1528 | 35.3636 | 13.54% | P | 0.88073
Cardiff Scientific Services | 50% TEA in Acetone | 2004 | UC | Cardiff CC |  | 12 | 37.4686 | 30.5833 | 22.51% | P | 0.81624
Bristol Scientific Services | 20% TEA in Water | 2004 | B | Cheltenham BC |  | 12 | 22.6111 | 23.3333 | -3.095% | G | 1.0319
Glasgow Scientific Services | 20% TEA in Water | 2004 | R | North Lanarkshire Council |  | 10 | 28.6 | 24.31 | 17.65% | G | 0.85
Jesmond Dene Laboratory | 50% TEA in Acetone | 2004 | UC | Newcastle upon Tyne CC |  | 10 | 27.812 | 28.3685 | -1.962% | P | 1.02
Harwell Scientific Services | 50% TEA in Acetone | 2003 | R | Cambridge CC |  | 11 | 52.9818 | 46.1891 | 14.71% | S | 0.87179
Harwell Scientific Services | 50% TEA in Acetone | 2004 | R | Cambridge CC |  | 12 | 43.9833 | 41.5925 | 5.748% | S | 0.94564
South Yorkshire Labs | 50% TEA in Acetone | 2004 | UB | Bournemouth BC |  | 12 | 22.3846 | 18.3785 | 21.8% | S | 0.82103
Bristol Scientific Services | 20% TEA in Water | 2004 | R | West Wiltshire DC |  | 10 | 32.15 | 26.1 | 23.18% | G | 0.81182
Lambeth Scientific Services | 50% TEA in Acetone | 2004 | R | Newtownabbey BC |  | 10 | 20.2667 | 29.4 | -31.07% | P | 1.4507
Lambeth Scientific Services | 50% TEA in Acetone | 2004 | R | Newtownabbey BC |  | 11 | 24.2727 | 31.2727 | -22.38% | P | 1.2884
Milton Keynes Council | 20% TEA in Water | 2004 | UB | Milton Keynes Council |  | 12 | 28.4153 | 21.2333 | 33.82% | G | 0.74725
Walsall MBC | 50% TEA in Acetone | 2004 | UB | Warwick DC |  | 12 | 31.5278 | 26.0833 | 20.87% | G | 0.82731
Walsall MBC | 50% TEA in Acetone | 2004 | R | Warwick DC |  | 11 | 50.4848 | 46.5455 | 8.464% | P | 0.92197
Bristol Scientific Services | 20% TEA in Water | 2004 | UB | LB Waltham Forest |  | 11 | 37.6636 | 38.0727 | -1.074% | S | 1.0109
Harwell Scientific Services | 50% TEA in Acetone | 2004 | UB | Rugby BC |  | 9 | 26.1204 | 20.6972 | 26.2% | P | 0.79238
Environmental Scientific Groups | 10% TEA in Water | 2003 | R | Dumfries and Galloway Council |  | 10 | 34.2333 | 36.794 | -6.959% | P | 1.0748
Environmental Scientific Groups | 10% TEA in Water | 2004 | R | Dumfries and Galloway Council |  | 12 | 38.0833 | 37.0908 | 2.676% | G | 0.97394
Gradko | 50% TEA in Acetone | 2004 | UB | AEA Tech Intercomparison |  | 11 | 24.2245 | 24.5455 | -1.307% | G | 1.0132
Harwell Scientific Services | 50% TEA in Acetone | 2004 | UB | AEA Tech Intercomparison |  | 11 | 33.7545 | 24.6364 | 37.01% | G | 0.72987
Jesmond Dene Laboratory | 50% TEA in Acetone | 2004 | UB | AEA Tech Intercomparison |  | 12 | 27.5944 | 24.4167 | 13.01% | G | 0.88484
Milton Keynes Council | 20% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 11 | 30.6455 | 25 | 22.58% | G | 0.81578
Northampton BC | 20% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 10 | 26.189 | 24.8 | 5.601% | P | 0.94696
Staffordshire Scientific Services | 50% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 11 | 25.5909 | 25 | 2.364% | P | 0.97691
South Yorkshire Labs | 50% TEA in Acetone | 2004 | UB | AEA Tech Intercomparison |  | 12 | 33.9167 | 24.4167 | 38.91% | G | 0.7199
Walsall MBC | 50% TEA in Acetone | 2004 | UB | AEA Tech Intercomparison |  | 12 | 29.2778 | 24.4167 | 19.91% | G | 0.83397
Glasgow Scientific Services | 20% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 10 | 31.5827 | 25.4 | 24.34% | G | 0.80424
Edinburgh Scientific Services | 50% TEA in Acetone | 2004 | UB | AEA Tech Intercomparison |  | 12 | 24.8104 | 24.4167 | 1.613% | P | 0.98413
Exova | 20% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 12 | 34.6194 | 24.4167 | 41.79% | G | 0.70529
Environmental Scientific Groups | 10% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 12 | 31.1561 | 24.4167 | 27.6% | G | 0.78369
Bristol Scientific Services | 20% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 12 | 27.8833 | 24.4167 | 14.2% | G | 0.87567
Gradko | 20% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 12 | 29.7358 | 24.4167 | 21.78% | G | 0.82112
Gradko | 50% TEA in Water | 2004 | UB | AEA Tech Intercomparison |  | 12 | 24.9578 | 24.4167 | 2.216% | G | 0.97832
Lancashire CC | 50% TEA in Acetone | 2003 | UC | Lancaster CC |  | 12 | 27.25 | 31.9783 | -14.79% | P | 1.1735
Lancashire CC | 50% TEA in Acetone | 2004 | UC | Lancaster CC |  | 12 | 28.1389 | 31.451 | -10.53% | P | 1.1177
Gradko | 50% TEA in Acetone | 2004 | B | LB Barnet |  | 11 | 31.5303 | 36.2727 | -13.07% | G | 1.1504
South Yorkshire Labs | 50% TEA in Acetone | 2004 | R | Barnsley MBC |  | 10 | 47.7333 | 46.93 | 1.712% | G | 0.98317
Gradko | 50% TEA in Acetone | 2004 | UB | RB Kensington and Chelsea |  | 11 | 35.3109 | 38.5897 | -8.496% | S | 1.0929
Gradko | 50% TEA in Acetone | 2004 | R | RB Kensington and Chelsea |  | 12 | 68.9731 | 79.5867 | -13.34% | S | 1.1539
Gradko | 50% TEA in Acetone | 2004 | UB | LB Brent |  | 10 | 30.11 | 30.0742 | 0.1193% | S | 0.99881
Gradko | 50% TEA in Acetone | 2004 | UC | LB Camden |  | 11 | 49.57 | 58.9563 | -15.92% | S | 1.1894
Gradko | 50% TEA in Acetone | 2004 | R | LB Croydon |  | 12 | 51.7264 | 63.6833 | -18.78% | S | 1.2312
Gradko | 50% TEA in Acetone | 2004 | R | LB Hounslow |  | 11 | 65.6482 | 68.6545 | -4.379% | S | 1.0458
Gradko | 50% TEA in Acetone | 2004 | R | LB Hounslow |  | 11 | 52.7145 | 48.5912 | 8.486% | S | 0.92178
Gradko | 50% TEA in Acetone | 2004 | S | LB Hounslow |  | 10 | 33.1988 | 46.35 | -28.37% | S | 1.3961
Gradko | 50% TEA in Acetone | 2004 | R | LB Hillingdon |  | 9 | 47.6789 | 52.2667 | -8.778% | S | 1.0962
Gradko | 50% TEA in Acetone | 2004 | R | LB Hillingdon |  | 9 | 39.9051 | 40.4444 | -1.333% | S | 1.0135
Gradko | 50% TEA in Acetone | 2004 | R | LB Greenwich |  | 11 | 44.9366 | 46.4091 | -3.173% | S | 1.0328
Gradko | 50% TEA in Acetone | 2004 | UB | LB Southwark |  | 9 | 53.8522 | 50.6413 | 6.34% | S | 0.94038
Gradko | 50% TEA in Acetone | 2004 | S | LB Richmond |  | 12 | 27 | 40 | -32.65% | S | 1.45
Gradko | 50% TEA in Acetone | 2004 | R | LB Richmond |  | 12 | 40.5491 | 37.6167 | 7.795% | S | 0.92768
Gradko | 50% TEA in Acetone | 2004 | S | LB Bexley |  | 10 | 32.572 | 34.5068 | -5.607% | S | 1.0594
Staffordshire Scientific Services | 50% TEA in Water | 2004 | 0 | Stoke-on-Trent City Council |  | 12 | 32.1417 | 30.25 | 6.253% | P | 0.94115
Staffordshire Scientific Services | 50% TEA in Water | 2004 | 0 | Stoke-on-Trent City Council |  | 12 | 47.4083 | 41.5833 | 14.01% | S | 0.87713
Staffordshire Scientific Services | 50% TEA in Water | 2004 | 0 | Stoke-on-Trent City Council |  | 12 | 41.4667 | 39.8 | 4.188% | S | 0.95981
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | R | Calderdale MBC |  | 11 | 55.8485 | 37.7 | 48.14% | P | 0.67504
Kent Scientific Services | 50% TEA in Acetone | 2005 | R | Eastleigh BC |  | 12 | 52.3636 | 50.0867 | 4.546% | P | 0.95652
Kent Scientific Services | 50% TEA in Acetone | 2005 | O | Eastleigh BC |  | 9 | 32.6754 | 36.8411 | -11.31% | P | 1.1275
Gradko | 20% TEA in Water | 2005 | R | Exeter CC |  | 10 | 36.9131 | 44.1552 | -16.4% | G | 1.1962
Jesmond Dene Laboratory | 50% TEA in Acetone | 2005 | R | Gateshead Council |  | 12 | 33.0583 | 36.4 | -9.18% | P | 1.1011
Gradko | 50% TEA in Acetone | 2003 | R | LB Richmond |  | 11 | 38.1958 | 48.1045 | -20.6% | P | 1.2594
Kent Scientific Services | 50% TEA in Acetone | 2005 | UB | Medway Council |  | 10 | 26.628 | 25.46 | 4.588% | S | 0.95614
Kent Scientific Services | 50% TEA in Acetone | 2005 | R | Medway Council |  | 10 | 43.478 | 31.92 | 36.21% | S | 0.73416
Kent Scientific Services | 50% TEA in Acetone | 2005 | Rural | Medway Council |  | 12 | 24.9833 | 18.9667 | 31.72% | S | 0.75917
Gradko | 50% TEA in Acetone | 2003 | B | LB Richmond |  | 12 | 27.9294 | 46.83 | -40.36% | P | 1.6767
Gradko | 50% TEA in Acetone | 2005 | UB | Sheffield CC |  | 12 | 30.1196 | 32.4917 | -7.301% | G | 1.0788
Gradko | 50% TEA in Acetone | 2005 | UC | Sheffield CC |  | 12 | 31.2222 | 45.0833 | -30.75% | G | 1.444
Gradko | 20% TEA in Water | 2005 | R | South Lakeland DC |  | 12 | 31.4551 | 25.765 | 22.08% | G | 0.8191
Harwell Scientific Services | 50% TEA in Acetone | 2005 | B | Swale BC |  | 10 | 26.3333 | 25.22 | 4.414% | P | 0.95772
Harwell Scientific Services | 50% TEA in Acetone | 2005 | R | Swale BC |  | 10 | 29.576 | 26.97 | 9.663% | G | 0.91189
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | K | Wakefield MDC |  | 12 | 60.25 | 44.4075 | 35.68% | S | 0.73705
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | UB | Wakefield MDC |  | 9 | 33.3704 | 28.8622 | 15.62% | P | 0.86491
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | UC | Wakefield MDC |  | 12 | 46.0833 | 34.8575 | 32.2% | P | 0.7564
Gradko | 50% TEA in Acetone | 2005 | R | Weymouth & Portland BC |  | 10 | 27.065 | 46.686 | -42.03% | S | 1.725
Environmental Scientific Groups | 10% TEA in Water | 2005 | UB | Wigan MBC |  | 10 | 32.0167 | 23.4 | 36.82% | G | 0.73087
Environmental Scientific Groups | 10% TEA in Water | 2005 | UB | Stockport MBC |  | 11 | 42.4697 | 31.7273 | 33.86% | P | 0.74706
Gradko | 50% TEA in Acetone | 2005 | B | LB Richmond |  | 10 | 28.2767 | 33.539 | -15.69% | P | 1.1861
Jesmond Dene Laboratory | 50% TEA in Acetone | 2005 | UC | Newcastle upon Tyne CC |  | 12 | 28.6731 | 29.7417 | -3.593% | G | 1.0373
South Yorkshire Labs | 50% TEA in Acetone | 2005 | UB | Bournemouth BC |  | 11 | 18.5455 | 16.9664 | 9.307% | S | 0.91485
Harwell Scientific Services | 50% TEA in Acetone | 2005 | R | Cambridge CC |  | 11 | 41.7091 | 43.0373 | -3.086% | S | 1.0318
South Yorkshire Labs | 50% TEA in Acetone | 2005 | R | Barnsley MBC |  | 12 | 44.1944 | 44.0833 | 0.252% | G | 0.99749
Environmental Scientific Groups | 10% TEA in Water | 2005 | R | Dumfries and Galloway Council |  | 12 | 37.3889 | 36.1225 | 3.274% | G | 0.96613
Gradko | 50% TEA in Acetone | 2005 | UB | Darlington BC |  | 12 | 23.1056 | 24.375 | -5.208% | G | 1.0549
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | B | Leeds CC |  | 11 | 41.3182 | 41.3182 | 0% | P | 1
Harwell Scientific Services | 50% TEA in Acetone | 2005 | UC | Vale of White Horse DC |  | 9 | 25.2741 | 24.7778 | 2.003% | G | 0.98036
Gradko | 50% TEA in Acetone | 2005 | R | Hammersmith and Fulham |  | 10 | 81.499 | 71.96 | 13.26% | S | 0.88296
Harwell Scientific Services | 50% TEA in Acetone | 2005 | UB | Rugby BC |  | 12 | 27.1556 | 20.126 | 34.93% | P | 0.74114
Gradko | 50% TEA in Acetone | 2004 | R | LB Hounslow |  | 11 | 52.1754 | 54.6182 | -4.472% | G | 1.0468
Gradko | 50% TEA in Acetone | 2004 | R | LB Hounslow |  | 11 | 65.1085 | 77.0091 | -15.45% | G | 1.1828
Gradko | 50% TEA in Acetone | 2004 | UB | LB Hounslow |  | 11 | 32.5323 | 35.7 | -8.873% | G | 1.0974
Harwell Scientific Services | 50% TEA in Acetone | 2005 | R | Adur DC |  | 12 | 43.1667 | 34.8333 | 23.92% | S | 0.80695
Harwell Scientific Services | 50% TEA in Acetone | 2005 | B | Newport CC |  | 12 | 30.5194 | 24.8417 | 22.86% | G | 0.81396
Gradko | 20% TEA in Water | 2005 | UB | LB Ealing |  | 12 | 42.1639 | 39.2 | 7.561% | G | 0.92971
Gradko | 20% TEA in Water | 2005 | R | LB Ealing |  | 12 | 58.4 | 58.6417 | -0.4121% | G | 1.0041
Gradko | 20% TEA in Water | 2005 | R | LB Ealing |  | 10 | 75.66 | 96.21 | -21.36% | G | 1.2716
Gradko | 20% TEA in Water | 2005 | R | Dartford Council |  | 9 | 61.0193 | 62.7667 | -2.784% | G | 1.0286
University of Essex | 50% TEA in Acetone | 2005 | S | Colchester BC |  | 11 | 46.2879 | 58.0556 | -20.27% | P | 1.2542
Cardiff Scientific Services | 50% TEA in Acetone | 2005 | R | Monmouthshire DC |  | 12 | 50.7083 | 36.977 | 37.13% | G | 0.72921
Cardiff Scientific Services | 50% TEA in Acetone | 2003 | I | Monmouthshire DC |  | 12 | 30.4167 | 33.968 | -10.46% | G | 1.1168
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | B | East Hertfordshire DC |  | 11 | 21.7273 | 20.9636 | 3.643% | S | 0.96485
Harwell Scientific Services | 50% TEA in Acetone | 2005 | I | Huntingdon DC |  | 9 | 25.2987 | 26.6108 | -4.931% | P | 1.0519
Gradko | 50% TEA in Acetone | 2005 | R | LB Southwark |  | 11 | 50.6 | 58.5636 | -13.6% | G | 1.1574
South Yorkshire Labs | 50% TEA in Acetone | 2004 | R | Oxford CC |  | 11 | 66.2273 | 66.7909 | -0.8439% | G | 1.0085
Gradko | 20% TEA in Water | 2004 | UB | Worcester CC |  | 12 | 20 | 14 | 47.5% | G | 0.68
South Yorkshire Labs | 50% TEA in Acetone | 2005 | R | Oxford CC |  | 12 | 59.875 | 67.5833 | -11.3% | G | 1.1273
Harwell Scientific Services | 50% TEA in Acetone | 2005 | R | Gravesham BC |  | 11 | 55.5773 | 54.3909 | 2.181% | G | 0.97865
Harwell Scientific Services | 50% TEA in Acetone | 2005 | B | Gravesham BC |  | 12 | 36.1278 | 29.9917 | 20.46% | G | 0.83016
Bristol Scientific Services | 20% TEA in Water | 2005 | R | Brighton and Hove CC |  | 12 | 32.4306 | 36 | -9.1% | P | 1.1
Harwell Scientific Services | 50% TEA in Acetone | 2005 | R | Hambleton DC |  | 12 | 26.2306 | 25.0164 | 4.853% | P | 0.95371
Gradko | 20% TEA in Water | 2005 | R | North Warwickshire BC |  | 9 | 51.6137 | 43.0867 | 19.79% | G | 0.83479
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | O | Reigate & Banstead BC |  | 12 | 21.5455 | 29.4561 | -26.86% | P | 1.3672
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | O | Reigate & Banstead BC |  | 12 | 24.303 | 34.286 | -29.12% | P | 1.4108
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | O | Reigate & Banstead BC |  | 9 | 15.25 | 19.4324 | -21.52% | P | 1.2743
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | UB | Leeds CC |  | 10 | 42.8 | 32.56 | 31.45% | S | 0.76075
Bristol Scientific Services | 20% TEA in Water | 2005 | UB | Cheltenham BC |  | 12 | 20.975 | 23.15 | -9.395% | G | 1.1037
Environmental Scientific Groups | 10% TEA in Water | 2005 | O | Flintshire CC |  | 12 | 20.1944 | 16.5833 | 21.78% | G | 0.82118
Staffordshire Scientific Services | 50% TEA in Water | 2005 | O | South Staffordshire DC |  | 12 | 32.8194 | 35.335 | -7.119% | G | 1.0766
Staffordshire Scientific Services | 50% TEA in Water | 2005 | UC | Stoke-on-Trent City Council |  | 12 | 31.1083 | 33.8357 | -8.061% | P | 1.0877
Environmental Scientific Groups | 10% TEA in Water | 2005 | R | Wrexham CBC |  | 12 | 23.3611 | 19.6192 | 19.07% | P | 0.83982
Gradko | 20% TEA in Water | 2005 | B | St Albans DC |  | 10 | 26.2767 | 25.9481 | 1.266% | G | 0.9875
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | UB | Elmbridge BC |  | 10 | 19.7333 | 29.73 | -33.62% | P | 1.5066
Environmental Scientific Groups | 10% TEA in Water | 2005 | R | Nottingham CC |  | 12 | 49.4722 | 36.4167 | 35.85% | G | 0.7361
Environmental Scientific Groups | 10% TEA in Water | 2005 | UC | Nottingham CC |  | 10 | 44.1333 | 33.7 | 30.96% | G | 0.7636
Environmental Scientific Groups | 10% TEA in Water | 2005 | O | Nottingham CC |  | 12 | 40.0556 | 35.4167 | 13.1% | G | 0.88419
Glasgow Scientific Services | 20% TEA in Water | 2005 | UB | Glasgow CC |  | 11 | 47.0303 | 34.4273 | 36.61% | P | 0.73202
Glasgow Scientific Services | 20% TEA in Water | 2005 | K | Glasgow CC |  | 12 | 87.0556 | 64.2667 | 35.46% | G | 0.73823
South Yorkshire Labs | 50% TEA in Acetone | 2004 | R | Oxford CC |  | 10 | 67.15 | 67.85 | -1.032% | G | 1.0104
Environmental Scientific Groups | 10% TEA in Water | 2004 | O | Flintshire CC |  | 12 | 20.3056 | 14.5 | 40.04% | G | 0.71409
Milton Keynes Council | 20% TEA in Water | 2005 | UC | Milton Keynes Council |  | 12 | 30.1028 | 22.8683 | 31.64% | G | 0.75968
Milton Keynes Council | 20% TEA in Water | 2005 | R | Milton Keynes Council |  | 11 | 44.8652 | 38.6155 | 16.18% | G | 0.8607
Gradko | 20% TEA in Water | 2005 | R | Rushmoor BC |  | 10 | 40.96 | 38.65 | 5.977% | G | 0.9436
Aberdeen CC | 20% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 10 | 30.486 | 25.2 | 20.98% | G | 0.82661
Bristol Scientific Services | 20% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 12 | 26.4139 | 24.3333 | 8.55% | G | 0.92123
Bristol Scientific Services | 20% TEA in Water | 2005 | UB | LB Waltham Forest |  | 12 | 31.0833 | 36.9167 | -15.8% | S | 1.1877
Environmental Scientific Groups | 10% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 12 | 32.4694 | 24.3333 | 33.44% | G | 0.74942
Environmental Scientific Groups | 10% TEA in Water | 2005 | R | Derby CC |  | 12 | 49.5139 | 37.35 | 32.57% | P | 0.75433
Environmental Scientific Groups | 10% TEA in Water | 2005 | K | Derby CC |  | 10 | 51.4667 | 38.18 | 34.8% | P | 0.74184
Exova | 20% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 12 | 30.8161 | 24.3333 | 26.64% | G | 0.78963
Edinburgh Scientific Services | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 10 | 22.9217 | 22.5 | 1.874% | P | 0.9816
Glasgow Scientific Services | 20% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 12 | 30.7717 | 24.3333 | 26.46% | G | 0.79077
Gradko | 20% TEA in Water | 2005 | R | Macclesfield BC |  | 12 | 34.0775 | 34.0057 | 0.2112% | G | 0.99789
Gradko | 20% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 12 | 30.7833 | 24.3333 | 26.51% | G | 0.79047
Gradko | 20% TEA in Water | 2005 | Rural | South Shropshire DC |  | 10 | 4.9895 | 4.48 | 11.37% | G | 0.89789
Gradko | 50% TEA in Acetone | 2005 | S | Slough BC |  | 12 | 28.9444 | 31.4167 | -7.869% | G | 1.0854
Gradko | 50% TEA in Acetone | 2005 | UC | Slough BC |  | 12 | 33.0389 | 32.5833 | 1.398% | G | 0.98621
Gradko | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 12 | 28.7886 | 24.3333 | 18.31% | G | 0.84524
Gradko | 50% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 12 | 27.5164 | 24.3333 | 13.08% | G | 0.88432
Harwell Scientific Services | 50% TEA in Acetone | 2004 | R | Babergh DC |  | 12 | 30.5167 | 23.8917 | 27.73% | G | 0.78291
Harwell Scientific Services | 50% TEA in Acetone | 2005 | R | Babergh DC |  | 12 | 29.1056 | 23.7583 | 22.51% | G | 0.81628
Harwell Scientific Services | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 11 | 33.5361 | 23.4545 | 42.98% | P | 0.69938
Jesmond Dene Laboratory | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 12 | 25.45 | 24.3333 | 4.589% | G | 0.95612
Kirklees Council Scientific Services | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 11 | 29.8421 | 24.8182 | 20.24% | P | 0.83165
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 10 | 24.4017 | 24.8 | -1.606% | P | 1.0163
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | R | Newtownabbey BC |  | 12 | 27.2778 | 31.3333 | -12.94% | P | 1.1487
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | R | Newtownabbey BC |  | 12 | 21.1944 | 28.9167 | -26.71% | P | 1.3644
Lancashire CC | 50% TEA in Acetone | 2005 | I | Lancaster CC |  | 10 | 30.5667 | 32.5811 | -6.183% | P | 1.0659
Milton Keynes Council | 20% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 11 | 30.1242 | 24.2727 | 24.11% | G | 0.80575
Northampton BC | 20% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 10 | 25.3385 | 25.6 | -1.021% | P | 1.0103
South Yorkshire Labs | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 12 | 33.7222 | 24.3333 | 38.58% | G | 0.72158
Staffordshire Scientific Services | 50% TEA in Water | 2005 | UB | Netcen Intercomparison |  | 11 | 24.9882 | 25 | -0.04727% | G | 1.0005
Walsall MBC | 50% TEA in Acetone | 2005 | UB | Netcen Intercomparison |  | 12 | 28.6389 | 24.3333 | 17.69% | P | 0.84966
Walsall MBC | 50% TEA in Acetone | 2005 | R | Gedling BC |  | 12 | 28.0278 | 33.1667 | -15.49% | P | 1.1833
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | UC | Leeds CC |  | 11 | 53.5455 | 42.4727 | 26.07% | S | 0.79321
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 | UC | Leeds CC |  | 10 | 41.9167 | 32.27 | 29.89% | P | 0.76986
Glasgow Scientific Services | 20% TEA in Water | 2004 | R | Renfrewshire Council |  | 9 | 78.7037 | 82.3667 | -4.447% | P | 1.0465
Bristol Scientific Services | 20% TEA in Water | 2004 | Rural | Pembrokeshire CC |  | 11 | 4.9682 | 5.2155 | -4.741% | P | 1.0498
Lambeth Scientific Services | 50% TEA in Acetone | 2004 | R | Guildford BC |  | 12 | 33 | 37 | -10.81% | S | 1.1212
Bristol Scientific Services | 20% TEA in Water | 2005 | K | South Gloucestershire |  | 12 | 27.1917 | 29.0417 | -6.37% | P | 1.068
Bristol Scientific Services | 20% TEA in Water | 2005 | Rural | Pembrokeshire CC |  | 12 | 5.3389 | 5.055 | 5.616% | G | 0.94683
Cardiff Scientific Services | 50% TEA in Acetone | 2005 | O | Vale of Glamorgan |  | 11 | 15.2424 | 11.4545 | 33.07% | P | 0.75149
Gradko | 50% TEA in Acetone | 2005 | UB | LB Hounslow |  | 10 | 30.1522 | 37.9 | -20.44% | P | 1.257
Gradko | 50% TEA in Acetone | 2005 | R | LB Hounslow |  | 12 | 65.6745 | 72.55 | -9.477% | G | 1.1047
Gradko | 50% TEA in Acetone | 2005 | R | LB Hounslow |  | 12 | 52.6891 | 49.2167 | 7.055% | G | 0.9341
Gradko | 50% TEA in Acetone | 2005 | UB | West Berkshire Council |  | 12 | 21.7 | 23.0167 | -5.72% | G | 1.0607
Gradko | 20% TEA in Water | 2005 | UB | Dudley MBC |  | 12 | 29.6074 | 29.1208 | 1.671% | G | 0.98356
Gradko | 20% TEA in Water | 2005 | R | Dudley MBC |  | 12 | 45.4146 | 46.8353 | -3.033% | G | 1.0313
Harwell Scientific Services | 50% TEA in Acetone | 2005 | I | Watford BC |  | 11 | 42.7273 | 38.2091 | 11.82% | S | 0.89426
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | R | Guildford BC |  | 12 | 35.667 | 38 | -6.139% | S | 1.0654
South Yorkshire Labs | 50% TEA in Acetone | 2005 | UB | Oxford CC |  | 12 | 23.75 | 19.9271 | 19.18% | G | 0.83904
Environmental Scientific Groups | 10% TEA in Water | 2005 | R | Lisburn CC |  | 12 | 26.25 | 27.1667 | -3.374% | G | 1.0349
Environmental Scientific Groups | 10% TEA in Water | 2005 | R | North Down BC |  | 12 | 41.8889 | 27.6667 | 51.41% | G | 0.66048
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | R | Calderdale MBC |  | 12 | 60.7778 | 55.25 | 10.01% | P | 0.90905
Staffordshire Scientific Services | 50% TEA in Water | 2006 | UC | Stoke-on-Trent City Council |  | 11 | 32.7591 | 32.3606 | 1.232% | P | 0.98783
Staffordshire Scientific Services | 50% TEA in Water | 2006 | R | Stoke-on-Trent City Council |  | 12 | 49.8889 | 46.5488 | 7.176% | P | 0.93305
Harwell Scientific Services | 50% TEA in Acetone | 2006 | I | Watford BC |  | 9 | 39.4444 | 36.9111 | 6.863% | S | 0.93577
Cardiff Scientific Services | 50% TEA in Acetone | 2006 | UB | Cardiff CC |  | 12 | 35.8889 | 30.25 | 18.64% | G | 0.84288
Cardiff Scientific Services | 50% TEA in Acetone | 2006 | O | Vale of Glamorgan |  | 11 | 15.1212 | 11.6636 | 29.64% | G | 0.77134
Jesmond Dene Laboratory | 50% TEA in Acetone | 2006 | K | Gateshead Council |  | 12 | 45.9736 | 39.6083 | 16.07% | P | 0.86154
Jesmond Dene Laboratory | 50% TEA in Acetone | 2006 | K | Gateshead Council |  | 12 | 49.9972 | 36.255 | 37.9% | G | 0.72514
Jesmond Dene Laboratory | 50% TEA in Acetone | 2006 | K | Gateshead Council |  | 11 | 41.6455 | 35.6455 | 16.83% | P | 0.85593
Jesmond Dene Laboratory | 50% TEA in Acetone | 2006 | R | Gateshead Council |  | 11 | 42.7606 | 35.6545 | 19.93% | G | 0.83382
Jesmond Dene Laboratory | 50% TEA in Acetone | 2006 | R | Newcastle upon Tyne CC |  | 12 | 48.7278 | 48.7259 | 0.00383% | P | 0.99996
Milton Keynes Council | 20% TEA in Water | 2006 | R | Milton Keynes Council |  | 9 | 31.513 | 21.9622 | 43.49% | P | 0.69693
Milton Keynes Council | 20% TEA in Water | 2006 | S | Milton Keynes Council |  | 12 | 35.1708 | 26.1217 | 34.64% | G | 0.74271
Milton Keynes Council | 20% TEA in Water | 2006 | R | Milton Keynes Council |  | 12 | 49.9972 | 41.3033 | 21.05% | G | 0.82611
Milton Keynes Council | 20% TEA in Water | 2006 | UC | Milton Keynes Council |  | 12 | 30.7611 | 22.0258 | 39.66% | G | 0.71603
Environmental Scientific Groups | 10% TEA in Water | 2006 | R | North Down BC |  | 10 | 43.6667 | 22.7 | 92.36% | G | 0.51985
Environmental Scientific Groups | 10% TEA in Water | 2006 | R | Lisburn CC |  | 11 | 24.697 | 24 | 2.904% | P | 0.97178
South Yorkshire Labs | 50% TEA in Acetone | 2006 | R | Oxford CC |  | 12 | 57.6667 | 65.9304 | -12.53% | G | 1.1433
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | UC | Leeds CC |  | 11 | 46.5758 | 37.5818 | 23.93% | G | 0.8069
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | UC | Belfast CC |  | 11 | 28.3939 | 33.4 | -14.99% | P | 1.1763
Gradko | 50% TEA in Water | 2006 | K | New Forest DC |  | 11 | 48.913 | 41.2636 | 18.54% | G | 0.84361
Gradko | 50% TEA in Water | 2006 | R | New Forest DC |  | 10 | 31.848 | 30.59 | 4.112% | G | 0.9605
Northampton BC | 20% TEA in Water | 2006 | R | Northampton BC |  | 10 | 32.5667 | 30.6 | 6.427% | P | 0.93961
Gradko | 50% TEA in Acetone | 2006 | UC | Norwich CC |  | 12 | 20.9911 | 20.5502 | 2.146% | G | 0.97899
Gradko | 50% TEA in Acetone | 2006 | UB | Reading BC |  | 11 | 22.4148 | 21.9082 | 2.313% | G | 0.9774
Glasgow Scientific Services | 20% TEA in Water | 2006 | R | East Dunbartonshire Council |  | 12 | 39.9722 | 39.075 | 2.296% | G | 0.97755
Glasgow Scientific Services | 20% TEA in Water | 2006 | R | East Dunbartonshire Council |  | 12 | 34.5556 | 33.375 | 3.537% | G | 0.96584
Environmental Scientific Groups | 10% TEA in Water | 2006 | UB | Stockport MBC |  | 10 | 40 | 30.1 | 32.45% | P | 0.755
Edinburgh Scientific Services | 50% TEA in Acetone | 2006 | R | West Lothian Council |  | 10 | 16.7 | 20.7 | -19.32% | P | 1.2395
Gradko | 50% TEA in Acetone | 2006 | R | Boston BC |  | 11 | 38.9697 | 34.0909 | 14.31% | G | 0.87481
Bristol Scientific Services | 20% TEA in Water | 2006 | UB | Cheltenham BC |  | 12 | 22.0111 | 20.8333 | 5.653% | G | 0.94649
Bristol Scientific Services | 20% TEA in Water | 2006 | UB | LB Waltham Forest |  | 10 | 36.01 | 34.43 | 4.589% | S | 0.95612
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | UB | Spelthorne BC |  | 9 | 22.1852 | 34.8983 | -36.43% | P | 1.573
Kirklees Council Scientific Services | 50% TEA in Acetone | 2006 | R | Kirklees MC |  | 12 | 35.1651 | 32.6611 | 7.667% | P | 0.92879
Gradko | 20% TEA in Water | 2006 | R | Dartford Council |  | 12 | 58.2639 | 57.3333 | 1.623% | G | 0.98403
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | UB | Wakefield MDC |  | 10 | 30.8 | 27.695 | 11.21% | S | 0.89919
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | UB | Wakefield MDC |  | 12 | 31.625 | 26.4217 | 19.69% | G | 0.83547
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | K | Wakefield MDC |  | 12 | 57.0556 | 53.7983 | 6.055% | P | 0.94291
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | UC | Wakefield MDC |  | 9 | 49.3333 | 35.8656 | 37.55% | G | 0.727
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | UB | Wakefield MDC |  | 12 | 31.9167 | 28.1725 | 13.29% | G | 0.88269
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | UB | Leeds CC |  | 11 | 33.0909 | 30.8273 | 7.343% | G | 0.93159
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 | UB | Leeds CC |  | 11 | 43.7727 | 40.9455 | 6.905% | G | 0.93541
Bristol Scientific Services | 20% TEA in Water | 2006 | Rural | Pembrokeshire CC |  | 11 | 6.9333 | 5.4336 | 27.6% | G | 0.7837
Gradko | 50% TEA in Acetone | 2006 | R | LB Richmond |  | 12 | 39.0411 | 42.825 | -8.836% | G | 1.0969
Gradko | 50% TEA in Acetone | 2006 | UB | LB Richmond |  | 12 | 26.1967 | 29.9917 | -12.65% | G | 1.1449
Gradko | 20% TEA in Water | 2006 | R | Macclesfield BC |  | 12 | 36.3278 | 32.5 | 11.78% | G | 0.89463
Gradko | 20% TEA in Water | 2006 | R | Macclesfield BC |  | 10 | 34.104 | 31.4 | 8.611% | G | 0.92071
Gradko | 50% TEA in Acetone | 2006 | R | LB Hammersmith and Fulham |  | 10 | 81.959 | 80.04 | 2.398% | S | 0.97659
Gradko | 50% TEA in Acetone | 2006 | R | LB Hounslow |  | 10 | 49.938 | 55.52 | -10.05% | P | 1.1118
Gradko | 50% TEA in Acetone | 2006 | R | LB Hounslow |  | 12 | 64.0349 | 68.1417 | -6.027% | G | 1.0641
Gradko | 50% TEA in Acetone | 2006 | UB | LB Hounslow |  | 10 | 33.0711 | 37.46 | -11.72% | G | 1.1327
Harwell Scientific Services | 50% TEA in Acetone | 2006 | R | Cambridge CC |  | 11 | 47.1091 | 45.2727 | 4.056% | S | 0.96102
Tayside SS | 20% TEA in Water | 2006 | R | Fife Council |  | 12 | 33.9181 | 25.7667 | 31.64% | G | 0.75967
Gradko | 20% TEA in Water | 2006 | R | Dudley MBC |  | 9 | 42.921 | 44.1989 | -2.891% | P | 1.0298
Gradko | 20% TEA in Water | 2006 | UB | Dudley MBC |  | 12 | 27.2339 | 29.147 | -6.564% | G | 1.0702
Cardiff Scientific Services | 50% TEA in Acetone | 2006 | UB | Monmouthshire DC |  | 12 | 28.7 | 25.46 | 12.73% | na | 0.887
Cardiff Scientific Services | 50% TEA in Acetone | 2006 | R | Monmouthshire DC |  | 12 | 53.6 | 34.35 | 56.04% | na | 0.641
Bristol Scientific Services | 20% TEA in Water | 2006 | R | Brighton and Hove CC |  | 11 | 38.1636 | 33.4916 | 13.95% | G | 0.87758
Gradko | 20% TEA in Water | 2006 | UB | LB Ealing |  | 10 | 34.0167 | 33.07 | 2.863% | G | 0.97217
Gradko | 20% TEA in Water | 2006 | K | South Northamptonshire Council |  | 10 | 41.6052 | 36.178 | 15% | G | 0.86956
Gradko | 20% TEA in Water | 2006 | R | Teignbridge DC |  | 10 | 34.7067 | 37.022 | -6.254% | G | 1.0667
Harwell Scientific Services | 50% TEA in Acetone | 2006 | R | Adur DC |  | 11 | 41.9333 | 33.9182 | 23.63% | G | 0.80886
Harwell Scientific Services | 50% TEA in Acetone | 2006 | R | Ashford BC |  | 12 | 43.6667 | 35.3833 | 23.41% | G | 0.81031
Harwell Scientific Services | 50% TEA in Acetone | 2006 | B | Gravesham BC |  | 12 | 33.6361 | 29.0417 | 15.82% | G | 0.86341
Harwell Scientific Services | 50% TEA in Acetone | 2006 | R | Gravesham BC |  | 12 | 57.2611 | 51.8333 | 10.47% | G | 0.90521
Harwell Scientific Services | 50% TEA in Acetone | 2006 | O | Newport CC |  | 12 | 31.9556 | 21.8142 | 46.49% | G | 0.68264
Harwell Scientific Services | 50% TEA in Acetone | 2006 | UB | Rugby BC |  | 11 | 28.8939 | 19.2018 | 50.48% | G | 0.66456
Harwell Scientific Services | 50% TEA in Acetone | 2006 | R | Falkirk Council |  | 12 | 41.2042 | 30.4964 | 35.11% | G | 0.74013
Harwell Scientific Services | 50% TEA in Acetone | 2006 | R | Falkirk Council |  | 11 | 33.7136 | 18.1039 | 86.22% | G | 0.53699
Harwell Scientific Services | 50% TEA in Acetone | 2006 | UB | Falkirk Council |  | 12 | 25.5694 | 18.6677 | 36.97% | P | 0.73008
Kent Scientific Services | 20% TEA in Water | 2006 | R | Eastleigh BC |  | 11 | 49.4745 | 55.6082 | -11.03% | P | 1.124
Lancashire CC | 50% TEA in Acetone | 2006 | I | Lancaster CC |  | 12 | 28.0972 | 31.438 | -10.63% | P | 1.1189
Northampton BC | 20% TEA in Water | 2006 | UB | Northampton BC |  | 12 | 14.3472 | 20.1667 | -28.86% | P | 1.4056
South Yorkshire Labs | 50% TEA in Acetone | 2006 | R | Barnsley MBC |  | 12 | 44.1944 | 45.55 | -2.976% | G | 1.0307
Staffordshire Scientific Services | 50% TEA in Water | 2006 | UC | East Staffordshire BC |  | 10 | 32.8333 | 35.8 | -8.287% | G | 1.0904
Environmental Scientific Groups | 10% TEA in Water | 2006 | R | Dumfries and Galloway Council |  | 12 | 38.6111 | 37.4922 | 2.951% | G | 0.97134
Environmental Scientific Groups | 10% TEA in Water | 2006 | UB | Wigan MBC |  | 12 | 27.1528 | 28 | -3.026% | P | 1.0312
Environmental Scientific Groups | 10% TEA in Water | 2006 | UB | Wigan MBC |  | 11 | 29.2727 | 25.7273 | 13.78% | P | 0.87888
Environmental Scientific Groups | 10% TEA in Water | 2006 | UC | Nottingham CC |  | 12 | 39.1944 | 33.1 | 18.41% | G | 0.84451
Environmental Scientific Groups | 10% TEA in Water | 2006 | R | Nottingham CC |  | 11 | 44.5455 | 39.6364 | 12.39% | G | 0.8898
Environmental Scientific Groups | 10% TEA in Water | 2006 | O | Nottingham CC |  | 12 | 37.4167 | 35.2 | 6.297% | P | 0.94076
Bristol Scientific Services | 20% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 12 | 115.9167 | 111.1667 | 4.273% | G | 0.95902
Environmental Scientific Groups | 10% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 10 | 111.19 | 109.8 | 1.266% | G | 0.9875
Exova | 20% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 11 | 126.1242 | 110.1818 | 14.47% | G | 0.8736
Tayside SS | 20% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 11 | 123.4423 | 112.7273 | 9.505% | G | 0.9132
Edinburgh Scientific Services | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 11 | 110.6909 | 108.2727 | 2.233% | G | 0.97815
Glasgow Scientific Services | 20% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 12 | 119.0356 | 111.1667 | 7.078% | G | 0.93389
Gradko | 20% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 11 | 107.7167 | 111.0909 | -3.037% | G | 1.0313
Gradko | 20% TEA in Water | 2006 | R | Rushmoor BC |  | 12 | 34.9022 | 34.6667 | 0.6795% | G | 0.99325
Gradko | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 11 | 99.0233 | 111.0909 | -10.86% | G | 1.1219
Gradko | 50% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 11 | 101.6245 | 111.0909 | -8.521% | G | 1.0932
Harwell Scientific Services | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 10 | 131.4133 | 107.9 | 21.79% | G | 0.82107
Jesmond Dene Laboratory | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 12 | 120.6056 | 111.1667 | 8.491% | G | 0.92174
Kent Scientific Services | 20% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 9 | 112.4594 | 111.2222 | 1.112% | P | 0.989
Kirklees Council Scientific Services | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 9 | 128.6726 | 107.4444 | 19.76% | G | 0.83502
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | UB | City of York Council |  | 12 | 17.5358 | 19.5317 | -10.22% | P | 1.1138
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | R | City of York Council |  | 12 | 24.0188 | 35.7906 | -32.89% | P | 1.4901
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | R | City of York Council |  | 12 | 21.6807 | 31.7271 | -31.67% | P | 1.4634
Lambeth Scientific Services | 50% TEA in Acetone | 2005 | R | City of York Council |  | 10 | 22.8285 | 25.4544 | -10.32% | P | 1.115
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | O | Reigate & Banstead BC |  | 12 | 19.4167 | 29.3544 | -33.85% | P | 1.5118
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | O | Reigate & Banstead BC |  | 12 | 22.2222 | 32.4012 | -31.42% | P | 1.4581
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | O | Reigate & Banstead BC |  | 12 | 13.75 | 19.3936 | -29.1% | P | 1.4104
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | UB | City of York Council |  | 11 | 17.8547 | 20 | -9.7% | P | 1.1073
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | R | City of York Council |  | 11 | 25.8255 | 31.7129 | -18.56% | P | 1.228
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | B | East Hertfordshire DC |  | 10 | 20.3833 | 18.23 | 11.81% | P | 0.89436
Lancashire CC | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 9 | 101.5667 | 111.8889 | -9.225% | G | 1.1016
Milton Keynes Council | 20% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 12 | 121.6333 | 111.1667 | 9.415% | G | 0.91395
South Yorkshire Labs | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 10 | 130.4667 | 114.7 | 13.75% | G | 0.87915
Staffordshire Scientific Services | 50% TEA in Water | 2006 | K | AEA E&E Intercomparison |  | 12 | 100.8014 | 111.1667 | -9.324% | G | 1.1028
University of Essex | 50% TEA in Acetone | 2006 | K | AEA E&E Intercomparison |  | 12 | 74.1434 | 111.1667 | -33.3% | G | 1.4993
Tayside SS | 20% TEA in Water | 2006 | R | Dundee CC |  | 12 | 50.5028 | 34.8333 | 44.98% | G | 0.68973
Edinburgh Scientific Services | 50% TEA in Acetone | 2006 | R | Edinburgh CC |  | 12 | 42.2917 | 38.125 | 10.93% | G | 0.90148
Edinburgh Scientific Services | 50% TEA in Acetone | 2006 | R | Edinburgh CC |  | 12 | 41.3472 | 33.7667 | 22.45% | P | 0.81666
Gradko | 50% TEA in Acetone | 2006 | R | LB Croydon |  | 11 | 48.1852 | 53.7698 | -10.39% | G | 1.1159
Gradko | 50% TEA in Acetone | 2006 | K | LB Croydon |  | 12 | 60.8631 | 62.7003 | -2.93% | G | 1.0302
Gradko | 50% TEA in Acetone | 2006 | S | Slough BC |  | 9 | 28.237 | 32.4444 | -12.97% | G | 1.149
Gradko | 50% TEA in Acetone | 2006 | UC | Slough BC |  | 9 | 32.8889 | 36.6667 | -10.3% | G | 1.1149
Gradko | 50% TEA in Acetone | 2006 | S | Wokingham DC |  | 11 | 28.0606 | 20.8364 | 34.67% | G | 0.74255
Gradko | 50% TEA in Acetone | 2006 | S | Wokingham DC |  | 10 | 24.5167 | 21.2 | 15.64% | P | 0.86472
Gradko | 50% TEA in Acetone | 2006 | R | Wokingham DC |  | 10 | 41.7333 | 36.1 | 15.6% | G | 0.86502
Gradko | 50% TEA in Acetone | 2006 | R | Wokingham DC |  | 9 | 37.1111 | 36.6667 | 1.212% | P | 0.98802
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | R | Newtownabbey BC |  | 12 | 26.6944 | 31.75 | -15.92% | P | 1.1894
Lambeth Scientific Services | 50% TEA in Acetone | 2006 | R | Newtownabbey BC |  | 12 | 20.0556 | 28.6667 | -30.04% | P | 1.4294
South Yorkshire Labs | 50% TEA in Acetone | 2006 | UB | Bournemouth BC |  | 12 | 18.1667 | 16.8808 | 7.617% | S | 0.92922
South Yorkshire Labs | 50% TEA in Acetone | 2006 | UB | Chesterfield BC |  | 11 | 20.1212 | 18.261 | 10.19% | G | 0.90755
South Yorkshire Labs | 50% TEA in Acetone | 2006 | R | Chesterfield BC |  | 11 | 34.6667 | 28.3364 | 22.34% | G | 0.8174
South Yorkshire Labs | 50% TEA in Acetone | 2006 | UI | North Lincolnshire Council |  | 10 | 26.7333 | 20.313 | 31.61% | P | 0.75984
Bristol Scientific Services | 20% TEA in Water | 2007 | Rural | Pembrokeshire CC |  | 11 | 7.1182 | 5.2 | 36.89% | G | 0.73052
Bristol Scientific Services | 20% TEA in Water | 2007 | R | Brighton and Hove CC |  | 12 | 46.1861 | 33.4191 | 38.2% | G | 0.72358
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | Castlereagh BC |  | 12 | 30.5 | 32.75 | -6.87% | P | 1.0738
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | Dumfries and Galloway Council |  | 12 | 37.125 | 38.0723 | -1.949% | P | 1.0255
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | Lisburn CC |  | 11 | 28.6061 | 28.9091 | -1.048% | P | 1.0106
Environmental Scientific Groups | 10% TEA in Water | 2007 | UC | Manchester CC |  | 12 | 43.6667 | 44.5425 | -1.966% | P | 1.0201
Environmental Scientific Groups | 10% TEA in Water | 2007 | B | North Down BC |  | 11 | 46.8788 | 24.6364 | 90.28% | P | 0.52553
Environmental Scientific Groups | 10% TEA in Water | 2007 | B | Warrington BC |  | 12 | 25.4167 | 24.6333 | 3.18% | G | 0.96918
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | Wrexham CBC |  | 10 | 20.8667 | 19.694 | 5.954% | P | 0.9438
Environmental Scientific Groups | 10% TEA in Water | 2007 | S | Chichester DC |  | 10 | 35.2333 | 31.15 | 13.11% | P | 0.88411
Tayside SS | 20% TEA in Water | 2007 | K | Fife Council |  | 11 | 42.2909 | 51.8182 | -18.39% | G | 1.2253
Edinburgh Scientific Services | 50% TEA in Acetone | 2007 | R | West Lothian Council |  | 12 | 20.4167 | 23.1667 | -11.87% | P | 1.1347
Glasgow Scientific Services | 20% TEA in Water | 2007 | R | East Dunbartonshire Council |  | 12 | 29.2222 | 33.9 | -13.8% | P | 1.1601
Glasgow Scientific Services | 20% TEA in Water | 2007 | R | East Dunbartonshire Council |  | 12 | 38.0833 | 39.225 | -2.911% | P | 1.03
Glasgow Scientific Services | 20% TEA in Water | 2007 | R | North Lanarkshire Council |  | 12 | 20.4722 | 22.0519 | -7.163% | G | 1.0772
Gradko | 20% TEA in Water | 2007 | R | Dartford Council |  | 12 | 60.2886 | 55.25 | 9.12% | G | 0.91643
Gradko | 20% TEA in Water | 2007 | R | Macclesfield BC |  | 12 | 36.7874 | 33.1908 | 10.84% | G | 0.90223
Gradko | 20% TEA in Water | 2007 | R | Macclesfield BC |  | 12 | 56.4861 | 44.1133 | 28.05% | G | 0.78096
Gradko | 20% TEA in Water | 2007 | R | Macclesfield BC |  | 10 | 28.2665 | 26.75 | 5.669% | G | 0.94635
Gradko | 20% TEA in Water | 2007 | K | South Lakeland DC |  | 11 | 45.4088 | 41.1991 | 10.22% | G | 0.90729
Gradko | 20% TEA in Water | 2007 | UB | Coventry CC |  | 12 | 23.6528 | 18.5724 | 27.35% | G | 0.78521
Gradko | 20% TEA in Water | 2007 | R | Coventry CC |  | 10 | 52.4133 | 46.18 | 13.5% | P | 0.88107
Gradko | 20% TEA in Water | 2007 | R | Coventry CC |  | 12 | 40.4297 | 29.825 | 35.56% | P | 0.7377
Gradko | 20% TEA in Water | 2007 | R | Coventry CC |  | 12 | 40.5464 | 35.6 | 13.89% | G | 0.87801
Gradko | 20% TEA in Water | 2007 | R | Ellesmere Port & Neston BC |  | 12 | 46.3528 | 44.5 | 4.164% | G | 0.96003
Gradko | 50% TEA in Acetone | 2007 | R | Boston BC |  | 9 | 41.4185 | 32.7778 | 26.36% | G | 0.79138
Gradko | 50% TEA in Acetone | 2007 | R | LB Hammersmith and Fulham |  | 12 | 88.7842 | 79.4083 | 11.81% | S | 0.8944
Gradko | 50% TEA in Acetone | 2007 | UB | Reading BC |  | 12 | 25.2425 | 23.8392 | 5.887% | G | 0.94441
Gradko | 50% TEA in Acetone | 2007 | R | Stevenage BC |  | 12 | 38.989 | 32.3667 | 20.46% | G | 0.83015
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Adur DC |  | 12 | 47.3639 | 32.6667 | 44.99% | G | 0.6897
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Ashford BC |  | 12 | 45.4528 | 31.175 | 45.8% | G | 0.68588
Harwell Scientific Services | 50% TEA in Acetone | 2007 | BI | Gravesham BC |  | 12 | 34.9694 | 29.1667 | 19.9% | G | 0.83406
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Gravesham BC |  | 12 | 56.9222 | 51.0833 | 11.43% | G | 0.89742
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Hambleton DC |  | 11 | 28.1303 | 21.5894 | 30.3% | G | 0.76748
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Leeds CC |  | 12 | 43.7778 | 39.2 | 11.68% | P | 0.89543
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | UC | Newcastle upon Tyne CC |  | 9 | 34.3403 | 27.7889 | 23.58% | G | 0.80922
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Swale BC |  | 11 | 28.2621 | 26.6878 | 5.899% | G | 0.9443
Harwell Scientific Services | 50% TEA in Acetone | 2007 | B | Swale BC |  | 12 | 28.7611 | 26.9151 | 6.859% | G | 0.93582
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Falkirk Council |  | 10 | 34.02 | 25 | 36.08% | G | 0.73486
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Falkirk Council |  | 11 | 48.7545 | 31.6545 | 54.02% | G | 0.64926
Harwell Scientific Services | 50% TEA in Acetone | 2007 | UB | Falkirk Council |  | 10 | 26.0367 | 19.04 | 36.75% | G | 0.73128
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Cambridge CC |  | 12 | 43.4583 | 41.7833 | 4.009% | G | 0.96146
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | K | Gateshead Council |  | 12 | 44.7361 | 34.7833 | 28.61% | P | 0.77752
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | K | Gateshead Council |  | 11 | 46.5 | 33.0091 | 40.87% | P | 0.70987
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | K | Gateshead Council |  | 11 | 43.8758 | 36.3727 | 20.63% | P | 0.82899
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | R | Gateshead Council |  | 9 | 46.6481 | 30.3556 | 53.67% | P | 0.65073
Kirklees Council Scientific Services | 50% TEA in Acetone | 2007 | R | Kirklees MC |  | 12 | 33.6747 | 28.1803 | 19.5% | G | 0.83684
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | R | East Hertfordshire DC |  | 10 | 21.85 | 18.14 | 20.45% | P | 0.83021
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | R | Elmbridge BC |  | 11 | 17.7879 | 19.9091 | -10.65% | P | 1.1193
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | R | Elmbridge BC |  | 11 | 40.0455 | 35.3636 | 13.24% | P | 0.88309
Lancashire CC | 50% TEA in Acetone | 2007 | R | Lancaster CC |  | 11 | 28.1667 | 27.6387 | 1.91% | P | 0.98126
Northampton BC | 20% TEA in Water | 2007 | R | Northampton BC |  | 11 | 34.3636 | 33.8182 | 1.613% | G | 0.98413
Northampton BC | 20% TEA in Water | 2007 | B | Northampton BC |  | 12 | 20.1667 | 21.5 | -6.202% | P | 1.0661
South Yorkshire Labs | 50% TEA in Acetone | 2007 | R | Barnsley MBC |  | 12 | 41.8056 | 45.0833 | -7.27% | G | 1.0784
South Yorkshire Labs | 50% TEA in Acetone | 2007 | R | North Lincolnshire Council |  | 10 | 29.2 | 23.7156 | 23.13% | P | 0.81218
South Yorkshire Labs | 50% TEA in Acetone | 2007 | UR | North Lincolnshire Council |  | 12 | 34.8611 | 30.8821 | 12.88% | P | 0.88586
South Yorkshire Labs | 50% TEA in Acetone | 2007 | UI | North Lincolnshire Council |  | 12 | 27.25 | 17.865 | 52.53% | G | 0.6556
Staffordshire Scientific Services | 50% TEA in Water | 2007 | R | East Staffordshire BC |  | 12 | 33.6944 | 32.8115 | 2.691% | G | 0.9738
Staffordshire Scientific Services | 50% TEA in Water | 2007 | O | South Staffordshire DC |  | 12 | 31.6417 | 36.2326 | -12.67% | G | 1.1451
Staffordshire Scientific Services | 50% TEA in Water | 2007 | UC | Stoke-on-Trent City Council |  | 11 | 31.1879 | 26.6809 | 16.89% | G | 0.85549
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | R | Calderdale MBC |  | 12 | 61.8889 | 50.91 | 21.57% | G | 0.8226
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | K | Leeds CC |  | 11 | 56.2727 | 46.4455 | 21.16% | S | 0.82536
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | UC | Leeds CC |  | 9 | 46.4444 | 43.5444 | 6.66% | S | 0.93756
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | B | Wakefield MDC |  | 12 | 29.9722 | 29.2867 | 2.341% | G | 0.97713
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | UB | Wakefield MDC |  | 12 | 40.5417 | 37.4042 | 8.388% | P | 0.92261
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | UB | Leeds CC |  | 12 | 42.25 | 43.0083 | -1.763% | G | 1.0179
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | UB | Leeds CC |  | 11 | 32.4545 | 25.6273 | 26.64% | G | 0.78964
Bristol Scientific Services | 20% TEA in Water | 2007 | K | South Gloucestershire |  | 9 | 29.4074 | 24.3058 | 20.99% | G | 0.82652
Bristol Scientific Services | 20% TEA in Water | 2007 | R | West Wiltshire DC |  | 9 | 37.963 | 25.5556 | 48.55% | G | 0.67317
Cardiff Scientific Services | 50% TEA in Acetone | 2007 | UB | Cardiff CC |  | 11 | 36.697 | 32.5455 | 12.76% | G | 0.89
Environmental Scientific Groups | 10% TEA in Water | 2007 | B | Wigan MBC |  | 11 | 29.9697 | 27.0909 | 1.902% | G | 0.90394
Environmental Scientific Groups | 10% TEA in Water | 2007 | B | Wigan MBC |  | 11 | 32.8485 | 21.2727 | 54.42% | G | 0.6476
Gradko | 20% TEA in Water | 2007 | B | Dudley MBC |  | 11 | 29.9286 | 29.8144 | 0.3831% | G | 0.99618
Gradko | 20% TEA in Water | 2007 | R | Dudley MBC |  | 12 | 48.2367 | 43.8241 | 10.07% | G | 0.90852
Gradko | 20% TEA in Water | 2007 | Rural | Dudley MBC |  | 10 | 19.0059 | 18.7406 | 1.416% | G | 0.98604
Gradko | 20% TEA in Water | 2007 | R | Rushmoor BC |  | 12 | 42.7028 | 37.3833 | 14.23% | G | 0.87543
Gradko | 20% TEA in Water | 2007 | B | St Albans DC |  | 11 | 25.4848 | 22.8564 | 11.5% | G | 0.89686
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Swansea CC |  | 12 | 36.6944 | 30.4525 | 20.5% | G | 0.82989
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Swansea CC |  | 12 | 43.5444 | 35.8058 | 21.61% | G | 0.82228
Harwell Scientific Services | 50% TEA in Acetone | 2007 | R | Swansea CC |  | 12 | 41.4861 | 36.0917 | 14.95% | G | 0.86997
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | UC | Belfast CC |  | 12 | 32.8889 | 32.75 | 0.4241% | P | 0.99578
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | R | LB Haringey |  | 12 | 34.0833 | 42.2583 | -19.35% | S | 1.2399
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | O | Reigate & Banstead BC |  | 12 | 24.25 | 28.9358 | -16.19% | P | 1.1932
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | O | Reigate & Banstead BC |  | 11 | 33.0909 | 34.3518 | -3.671% | P | 1.0381
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | O | Reigate & Banstead BC |  | 12 | 17.9444 | 20.8492 | -13.93% | P | 1.1619
South Yorkshire Labs | 50% TEA in Acetone | 2007 | UB | Bournemouth BC |  | 12 | 17.4444 | 15.1225 | 15.35% | G | 0.86689
Aberdeen CC | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 11 | 117.5894 | 102 | 15.28% | G | 0.86743
Bristol Scientific Services | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 12 | 114.8611 | 102.5833 | 11.97% | G | 0.89311
Environmental Scientific Groups | 10% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 11 | 108.1049 | 104.5455 | 3.405% | G | 0.96707
Exova | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 11 | 100.1227 | 100.7273 | -0.6002% | G | 1.006
Tayside SS | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 11 | 115.9901 | 101.4545 | 14.33% | G | 0.87468
Edinburgh Scientific Services | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 12 | 104.7528 | 102.5833 | 2.115% | G | 0.97929
Glasgow Scientific Services | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 12 | 110.09 | 102.5833 | 7.318% | G | 0.93181
Gradko | 20% TEA in Water | 2007 | O | North Warwickshire BC |  | 11 | 53.5014 | 44.0682 | 21.41% | G | 0.82368
Gradko | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 12 | 108.9342 | 102.5833 | 6.191% | G | 0.9417
Gradko | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 12 | 97.1275 | 102.5833 | -5.318% | G | 1.0562
Gradko | 50% TEA in Acetone | 2007 | UC | Norwich CC |  | 12 | 20.8575 | 22.4667 | -7.162% | G | 1.0772
Gradko | 50% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 12 | 103.4896 | 102.5833 | 0.8834% | G | 0.99124
Harwell Scientific Services | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 9 | 129.1037 | 105.6667 | 22.18% | G | 0.81846
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 12 | 117.9389 | 102.5833 | 14.97% | G | 0.8698
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | R | Newcastle upon Tyne CC |  | 12 | 49.8284 | 41.8083 | 19.18% | G | 0.83905
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 | UC | Newcastle upon Tyne CC |  | 9 | 34.3403 | 30.0222 | 14.38% | G | 0.87426
Kent Scientific Services | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 12 | 106.6978 | 102.5833 | 4.011% | G | 0.96144
Kirklees Council Scientific Services | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 12 | 100.4917 | 102.5833 | -2.039% | G | 1.0208
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 10 | 101.1367 | 99.3 | 1.85% | G | 0.98184
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | R | LB Islington |  | 10 | 53.1167 | 65.34 | -18.71% | P | 1.2301
Lancashire CC | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 9 | 95.2407 | 101 | -5.702% | G | 1.0605
Milton Keynes Council | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 11 | 118.9485 | 104.5455 | 13.78% | G | 0.87891
Milton Keynes Council | 20% TEA in Water | 2007 | UC | Milton Keynes Council |  | 11 | 31.4333 | 21.0673 | 49.2% | G | 0.67022
Milton Keynes Council | 20% TEA in Water | 2007 | R | Milton Keynes Council |  | 10 | 32.815 | 23.391 | 40.29% | G | 0.71281
Milton Keynes Council | 20% TEA in Water | 2007 | R | Milton Keynes Council |  | 12 | 50.0278 | 31.7942 | 57.35% | G | 0.63553
Northampton BC | 20% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 11 | 93.3212 | 102 | -8.509% | P | 1.093
South Yorkshire Labs | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 11 | 118.7273 | 100.7273 | 17.87% | G | 0.84839
Staffordshire Scientific Services | 50% TEA in Water | 2007 | R | Stoke-on-Trent City Council |  | 11 | 42.4835 | 44.4183 | -4.356% | P | 1.0455
Staffordshire Scientific Services | 50% TEA in Water | 2007 | R | Stoke-on-Trent City Council |  | 11 | 52.2777 | 40.8944 | 27.84% | G | 0.78225
Staffordshire Scientific Services | 50% TEA in Water | 2007 | K | AEA Tech Intercomparison |  | 12 | 96.1667 | 102.5833 | -6.255% | G | 1.0667
University of Essex | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 11 | 86.6309 | 100.7273 | -13.99% | G | 1.1627
Walsall MBC | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 11 | 90.8091 | 102.3636 | -11.29% | P | 1.1272
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | K | AEA Tech Intercomparison |  | 11 | 114.0303 | 101.4545 | 12.4% | G | 0.88972
Harwell Scientific Services | 50% TEA in Acetone | 2006 | K | Suffolk Coastal DC |  | 12 | 47.2681 | 44 | 7.427% | G | 0.93086
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | Carlisle CC |  | 12 | 40.3472 | 31.2 | 29.32% | P | 0.77329
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | LB Croydon |  | 12 | 54.7572 | 57.84 | -5.33% | G | 1.0563
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | LB Croydon |  | 12 | 66.9438 | 65.2558 | 2.587% | G | 0.97479
Environmental Scientific Groups | 10% TEA in Water | 2007 | UC | Nottingham CC |  | 12 | 40.6667 | 33.2833 | 22.18% | P | 0.81844
Environmental Scientific Groups | 10% TEA in Water | 2007 | R | Nottingham CC |  | 11 | 50.5758 | 42.6455 | 18.6% | P | 0.8432
Environmental Scientific Groups | 10% TEA in Water | 2007 | UB | Stockport MBC |  | 10 | 39.0333 | 31.1 | 25.51% | P | 0.79675
Gradko | 20% TEA in Water | 2007 | UB | LB Ealing |  | 12 | 42.75 | 38.8333 | 10.09% | G | 0.90838
Gradko | 20% TEA in Water | 2007 | R | LB Ealing |  | 12 | 59.3208 | 58.8083 | 0.8715% | G | 0.99136
Gradko | 20% TEA in Water | 2007 | UB | LB Ealing |  | 10 | 32.947 | 29.93 | 10.08% | G | 0.90843
Gradko | 20% TEA in Water | 2007 | UB | North Tyneside Council |  | 11 | 21.7742 | 18.929 | 15.03% | G | 0.86933
Gradko | 20% TEA in Water | 2007 | R | Southampton CC |  | 9 | 38.6 | 34.6889 | 11.27% | G | 0.89868
Gradko | 50% TEA in Acetone | 2007 | R | LB Hounslow |  | 12 | 58.3801 | 63.35 | -7.845% | G | 1.0851
Gradko | 50% TEA in Acetone | 2007 | UB | LB Hounslow |  | 12 | 31.4751 | 33.7667 | -6.786% | G | 1.0728
Gradko | 50% TEA in Acetone | 2007 | K | LB Camden |  | 10 | 84.1993 | 80.356 | 4.783% | G | 0.95435
Gradko | 50% TEA in Acetone | 2007 | S | LB Greenwich |  | 12 | 27.4167 | 29.8272 | -8.082% | S | 1.0879
Gradko | 50% TEA in Acetone | 2007 | R | LB Greenwich |  | 12 | 55.75 | 49.0279 | 13.71% | S | 0.87942
Gradko | 50% TEA in Acetone | 2007 | R | LB Greenwich |  | 12 | 53.8333 | 49.5108 | 8.731% | S | 0.9197
Gradko | 50% TEA in Acetone | 2007 | R | LB Greenwich |  | 11 | 68.7273 | 72.0794 | -4.651% | S | 1.0488
Gradko | 50% TEA in Acetone | 2007 | R | LB Greenwich |  | 12 | 46.75 | 45.3859 | 3.006% | S | 0.97082
Gradko | 50% TEA in Acetone | 2007 | R | LB Greenwich |  | 10 | 48.4 | 56.5414 | -14.4% | S | 1.1682
Gradko | 50% TEA in Water | 2007 | K | New Forest DC |  | 12 | 51.8604 | 43.8 | 18.4% | G | 0.84457
Gradko | 50% TEA in Water | 2007 | R | New Forest DC |  | 12 | 32.9406 | 31.175 | 5.663% | G | 0.9464
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | UB | City of York Council |  | 12 | 19.4167 | 24.75 | -21.55% | P | 1.2747
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | R | City of York Council |  | 12 | 29.7222 | 34.25 | -13.22% | P | 1.1523
Lambeth Scientific Services | 50% TEA in Acetone | 2007 | R | City of York Council |  | 11 | 32.2121 | 30.9091 | 4.216% | P | 0.95955
Lancashire CC | 50% TEA in Acetone | 2007 | UB | Preston CC |  | 12 | 24.0694 | 23.25 | 3.524% | P | 0.96595
South Yorkshire Labs | 50% TEA in Acetone | 2007 | B | Chesterfield BC |  | 12 | 21.6111 | 21.99 | -1.723% | G | 1.0175
South Yorkshire Labs | 50% TEA in Acetone | 2007 | R | Chesterfield BC |  | 11 | 35.1212 | 27.7 | 26.79% | G | 0.7887
South Yorkshire Labs | 50% TEA in Acetone | 2007 | R | Oxford CC |  | 12 | 51.8333 | 57.7167 | -10.19% | G | 1.1135
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 | UC | Leeds CC |  | 12 | 43.7778 | 37.3083 | 17.34% | P | 0.85222
Cardiff Scientific Services | 50% TEA in Acetone | 2007 | R | Monmouthshire CC |  | 11 | 52.5152 | 36.9364 | 42.18% | G | 0.7
Cardiff Scientific Services | 50% TEA in Acetone | 2007 | UB | Monmouthshire CC |  | 11 | 26.4848 | 25.2909 | 4.721% | G | 0.95
Bristol Scientific Services | 20% TEA in Water | 2008 | UB | LB Waltham Forest |  | 12 | 41.075 | 35.975 | 14.18% | S | 0.87584
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | Castlereagh BC |  | 10 | 35.1 | 22.8 | 53.95% | P | 0.64957
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | Castlereagh BC |  | 10 | 43.1333 | 31 | 39.14% | G | 0.7187
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | Dumfries and Galloway Council |  | 11 | 40 | 37.5667 | 6.477% | G | 0.93917
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | Lisburn CC |  | 12 | 33.3056 | 25.75 | 29.34% | P | 0.77314
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | North Down BC |  | 10 | 46.1333 | 32.5 | 41.95% | G | 0.70448
Gradko | 20% TEA in Water | 2008 | UC | Nottingham CC |  | 9 | 36.0574 | 32.2333 | 11.86% | G | 0.89394
Gradko | 20% TEA in Water | 2008 | R | Nottingham CC |  | 9 | 44.3274 | 42.0222 | 5.486% | G | 0.948
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | Wrexham CBC |  | 12 | 21.8333 | 20.525 | 6.374% | G | 0.94008
Exova | 20% TEA in Water | 2008 | R | Stirling Council |  | 11 | 28.7712 | 30.4341 | -5.464% | P | 1.0578
Tayside SS | 20% TEA in Water | 2008 | K | Fife Council |  | 10 | 44.6233 | 45.3845 | -1.677% | G | 1.0171
Tayside SS | 20% TEA in Water | 2008 | R | Fife Council |  | 11 | 39.8758 | 29.1468 | 36.81% | G | 0.73094
Edinburgh Scientific Services | 50% TEA in Acetone | 2007 | R | Edinburgh CC |  | 10 | 44.9167 | 41.07 | 9.366% | G | 0.91436
Edinburgh Scientific Services | 50% TEA in Acetone | 2007 | R | Edinburgh CC |  | 11 | 44.7879 | 41.8182 | 7.101% | P | 0.93369
Edinburgh Scientific Services | 50% TEA in Acetone | 2007 | R | Edinburgh CC |  | 11 | 41.4242 | 35.3182 | 17.29% | G | 0.8526
Edinburgh Scientific Services | 50% TEA in Acetone | 2007 | K | Edinburgh CC |  | 10 | 100.2 | 93.67 | 6.971% | G | 0.93483
Glasgow Scientific Services | 20% TEA in Water | 2008 | R | East Dunbartonshire Council |  | 10 | 28.8333 | 30.89 | -6.658% | P | 1.0713
Glasgow Scientific Services | 20% TEA in Water | 2008 | R | East Dunbartonshire Council |  | 11 | 42.2424 | 44.6182 | -5.325% | G | 1.0562
Glasgow Scientific Services | 20% TEA in Water | 2008 | R | East Dunbartonshire Council |  | 11 | 39.9091 | 35.3909 | 12.77% | P | 0.88679
Gradko | 50% TEA in Acetone | 2008 | R | Boston BC |  | 12 | 45.1944 | 35.5 | 27.31% | G | 0.78549
Gradko | 20% TEA in Water | 2008 | UC | Belfast CC |  | 11 | 41.4055 | 32.6818 | 26.69% | G | 0.78931
Gradko | 20% TEA in Water | 2008 | R | Crewe and Nantwich BC |  | 10 | 32.8974 | 31.566 | 4.218% | G | 0.95953
Gradko | 20% TEA in Water | 2008 | R | Dudley MBC |  | 12 | 48.9101 | 41.575 | 17.64% | G | 0.85003
Gradko | 20% TEA in Water | 2008 | B | Dudley MBC |  | 12 | 30.2132 | 28.2554 | 6.929% | P | 0.9352
Gradko | 20% TEA in Water | 2008 | Rural | Dudley MBC |  | 12 | 17.7714 | 16.9794 | 4.665% | G | 0.95543
Gradko | 20% TEA in Water | 2008 | R | Ellesmere Port & Neston BC |  | 12 | 44.6306 | 41.5625 | 7.382% | G | 0.93126
Gradko | 50% TEA in Acetone | 2008 | R | LB Hammersmith and Fulham |  | 10 | 89.902 | 67.92 | 32.36% | S | 0.75549
Gradko | 50% TEA in Acetone | 2008 | UB | Reading BC |  | 12 | 26.2644 | 23.1583 | 13.41% | G | 0.88174
Gradko | 20% TEA in Water | 2008 | R | Rhondda Cynon Taf CBC |  | 12 | 35.4333 | 35.44 | -0.01881% | G | 1.0002
Gradko | 20% TEA in Water | 2008 | R | South Hams DC |  | 10 | 41.21 | 40.16 | 2.615% | G | 0.97452
Gradko | 50% TEA in Acetone | 2008 | UC | Uttlesford DC |  | 11 | 26.0076 | 27.9673 | -7.007% | G | 1.0754
Gradko | 50% TEA in Acetone | 2008 | R | Stevenage BC |  | 12 | 36.6861 | 29 | 26.5% | G | 0.79049
Harwell Scientific Services | 50% TEA in Acetone | 2008 | R | Cambridge CC |  | 12 | 51.65 | 42.7667 | 20.77% | G | 0.82801
Harwell Scientific Services | 50% TEA in Acetone | 2008 | UB | Falkirk Council |  | 9 | 29.5296 | 26.0372 | 13.41% | G | 0.88173
Harwell Scientific Services | 50% TEA in Acetone | 2008 | R | Falkirk Council |  | 11 | 42.997 | 31.7713 | 35.33% | G | 0.73892
Harwell Scientific Services | 50% TEA in Acetone | 2008 | R | Hambleton DC |  | 12 | 27.9139 | 20.9309 | 33.36% | G | 0.74984
Harwell Scientific Services | 50% TEA in Acetone | 2007 | I | Suffolk Coastal DC |  | 12 | 53.5694 | 41.9167 | 27.8% | G | 0.78247
Harwell Scientific Services | 50% TEA in Acetone | 2007 | K | Suffolk Coastal DC |  | 12 | 47.7694 | 46.1667 | 3.472% | G | 0.96645
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | B | East Hertfordshire DC |  | 11 | 24.3333 | 18.6364 | 30.57% | P | 0.76588
Northampton BC | 20% TEA in Water | 2008 | B | Northampton BC |  | 12 | 23.5889 | 21 | 12.33% | G | 0.89025
Northampton BC | 20% TEA in Water | 2008 | R | Northampton BC |  | 10 | 42.53 | 30.59 | 39.03% | G | 0.71926
South Yorkshire Labs | 50% TEA in Acetone | 2008 | UB | Bournemouth BC |  | 12 | 16.1667 | 14.2958 | 13.09% | G | 0.88428
South Yorkshire Labs | 50% TEA in Acetone | 2008 | R | Chesterfield BC |  | 10 | 28.4 | 24.65 | 15.21% | G | 0.86796
South Yorkshire Labs | 50% TEA in Acetone | 2008 | B | Chesterfield BC |  | 10 | 19.0333 | 19.52 | -2.493% | G | 1.0256
South Yorkshire Labs | 50% TEA in Acetone | 2008 | UB | Kingston upon Hull City Council |  | 11 | 28.1515 | 29.4545 | -4.424% | P | 1.0463
South Yorkshire Labs | 50% TEA in Acetone | 2008 | R | Barnsley MBC |  | 10 | 41.7 | 46.7 | -10.71% | G | 1.1199
South Yorkshire Labs | 50% TEA in Acetone | 2008 | R | Rotherham MBC |  | 10 | 36.5 | 38.27 | -4.625% | G | 1.0485
Staffordshire Scientific Services | 50% TEA in Water | 2008 | O | South Staffordshire DC |  | 12 | 32.3944 | 35.7459 | -9.376% | P | 1.1035
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | R | Calderdale MBC |  | 10 | 64.2667 | 52.11 | 23.33% | G | 0.81084
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | R | Leeds CC |  | 12 | 52.2083 | 48.0917 | 8.56% | G | 0.92115
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | R | Leeds CC |  | 11 | 37.8182 | 27.8575 | 35.76% | S | 0.73662
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | UC | Leeds CC |  | 12 | 42.1389 | 35.525 | 18.62% | G | 0.84305
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | R | Leeds CC |  | 12 | 40.6667 | 32 | 27.08% | P | 0.78689
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | UB | Wakefield MDC |  | 11 | 33.2727 | 29.0199 | 14.65% | G | 0.87218
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | R | Wakefield MDC |  | 11 | 27 | 20.083 | 34.44% | G | 0.74382
Staffordshire Scientific Services | 50% TEA in Water | 2008 | S | Cannock Chase Council |  | 12 | 22.6833 | 24.3833 | -6.972% | G | 1.0749
Staffordshire Scientific Services | 50% TEA in Water | 2008 | R | Cannock Chase Council |  | 11 | 43.9364 | 49.5909 | -11.4% | P | 1.1287
Cardiff Scientific Services | 50% TEA in Acetone | 2008 | UB | Cardiff CC |  | 12 | 34.25 | 28.6667 | 19.48% | G | 0.83698
Environmental Scientific Groups | 10% TEA in Water | 2008 | S | Chichester DC |  | 12 | 39.6389 | 34.3917 | 15.26% | G | 0.86762
Staffordshire Scientific Services | 50% TEA in Water | 2008 | R | East Staffordshire BC |  | 11 | 33.0015 | 33.4273 | -1.274% | G | 1.0129
Jesmond Dene Laboratory | 50% TEA in Acetone | 2008 | R | Gateshead Council |  | 11 | 43.9394 | 34.2818 | 28.17% | P | 0.78021
Jesmond Dene Laboratory | 50% TEA in Acetone | 2008 | R | Gateshead Council |  | 12 | 41.5528 | 31.625 | 31.39% | G | 0.76108
Jesmond Dene Laboratory | 50% TEA in Acetone | 2008 | R | Gateshead Council |  | 12 | 46.9361 | 34.4167 | 36.38% | P | 0.73327
Jesmond Dene Laboratory | 50% TEA in Acetone | 2008 | R | Gateshead Council |  | 11 | 40.0712 | 30.2182 | 32.61% | G | 0.75411
Gradko | 50% TEA in Acetone | 2008 | R | LB Hounslow |  | 12 | 58.4934 | 58.5917 | -0.1677% | G | 1.0017
Gradko | 50% TEA in Acetone | 2008 | R | LB Hounslow |  | 12 | 68.6682 | 71.9583 | -4.572% | G | 1.0479
Gradko | 50% TEA in Acetone | 2008 | UB | LB Hounslow |  | 10 | 37.5591 | 30.06 | 24.95% | G | 0.80034
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | R | Leeds CC |  | 10 | 43.1 | 35.4 | 21.75% | S | 0.82135
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | R | Leeds CC |  | 10 | 49.5 | 46.7 | 5.996% | S | 0.94343
Bristol Scientific Services | 20% TEA in Water | 2008 | R | Lewes DC |  | 11 | 39.7909 | 37.5455 | 5.981% | S | 0.94357
Milton Keynes Council | 20% TEA in Water | 2008 | UC | Milton Keynes Council |  | 12 | 33.1125 | 23.6542 | 39.99% | G | 0.71436
Cardiff Scientific Services | 50% TEA in Acetone | 2008 | R | Monmouthshire CC |  | 12 | 50.1944 | 41.8716 | 19.88% | G | 0.83419
Cardiff Scientific Services | 50% TEA in Acetone | 2008 | UB | Monmouthshire CC |  | 11 | 26.0303 | 22.5756 | 15.3% | P | 0.86728
Gradko | 50% TEA in Acetone | 2008 | R | LB Redbridge |  | 12 | 45.6489 | 45.75 | -0.221% | G | 1.0022
Gradko | 50% TEA in Acetone | 2008 | R | LB Redbridge |  | 11 | 49.593 | 55.1455 | -10.07% | G | 1.112
Gradko | 50% TEA in Acetone | 2008 | UB | LB Redbridge |  | 12 | 36.8869 | 32.3917 | 13.88% | G | 0.87813
Gradko | 50% TEA in Acetone | 2008 | K | LB Redbridge |  | 10 | 51.5603 | 52.85 | -2.44% | G | 1.025
Gradko | 20% TEA in Water | 2008 | R | Rushmoor BC |  | 12 | 41.5267 | 37.99 | 9.309% | G | 0.91483
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | UB | Spelthorne BC |  | 12 | 36.0556 | 35.9419 | 0.3161% | G | 0.99685
Environmental Scientific Groups | 10% TEA in Water | 2008 | UB | Stockport MBC |  | 10 | 39.5 | 27.6 | 43.12% | G | 0.69873
Gradko | 50% TEA in Water | 2008 | R | Winchester |  | 11 | 46.4904 | 47.53 | -2.187% | G | 1.0224
Gradko | 50% TEA in Water | 2008 | R | South Norfolk |  | 10 | 27.533 | 34.326 | -19.79% | S | 1.2467
Harwell Scientific Services | 50% TEA in Acetone | 2008 | R | Canterbury CC |  | 10 | 48.41 | 39.8 | 21.63% | G | 0.82214
Harwell Scientific Services | 50% TEA in Acetone | 2008 | B | Canterbury CC |  | 12 | 22.3278 | 16.5833 | 34.64% | G | 0.74272
Aberdeen CC | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 10 | 135.2733 | 118.4 | 14.25% | G | 0.87526
Bristol Scientific Services | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 12 | 122.0361 | 115.75 | 5.431% | G | 0.94849
Tayside SS | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 12 | 134.6217 | 115.75 | 16.3% | G | 0.85982
Environmental Scientific Groups | 10% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 11 | 136.7576 | 117.1818 | 16.71% | G | 0.85686
Edinburgh Scientific Services | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 11 | 127.0273 | 117.1818 | 8.402% | G | 0.92249
University of Essex | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 12 | 103.3068 | 115.75 | -10.75% | G | 1.1204
Glasgow Scientific Services | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 12 | 133.0406 | 115.75 | 14.94% | G | 0.87004
Gradko | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 12 | 117.1389 | 115.75 | 1.2% | G | 0.98814
Gradko | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 12 | 103.3283 | 115.75 | -10.73% | G | 1.1202
Gradko | 50% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 12 | 108.5814 | 115.75 | -6.193% | G | 1.066
Harwell Scientific Services | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 12 | 138.0611 | 115.75 | 19.28% | G | 0.8384
Kent Scientific Services | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 11 | 125.279 | 114.4545 | 9.457% | G | 0.9136
Kirklees Council Scientific Services | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 11 | 128.715 | 113.7273 | 13.18% | G | 0.88356
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 11 | 100.1788 | 113.7273 | -11.91% | G | 1.1352
Lancashire CC | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 12 | 110.4583 | 115.75 | -4.572% | G | 1.0479
Milton Keynes Council | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 10 | 141.4667 | 121.1 | 16.82% | G | 0.85603
Northampton BC | 20% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 11 | 120.7594 | 114.4545 | 5.509% | G | 0.94779
South Yorkshire Labs | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 12 | 132.4722 | 115.75 | 14.45% | G | 0.87377
Staffordshire Scientific Services | 50% TEA in Water | 2008 | K | AEA Tech Intercomparison |  | 12 | 107.4944 | 115.75 | -7.132% | G | 1.0768
Walsall MBC | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 11 | 73.9333 | 116.4545 | -36.51% | P | 1.5751
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | K | AEA Tech Intercomparison |  | 12 | 120.3611 | 115.75 | 3.984% | G | 0.96169
Bristol Scientific Services | 20% TEA in Water | 2008 | R | Brighton and Hove CC |  | 12 | 42.6792 | 30.7667 | 38.72% | G | 0.72088
Gradko | 20% TEA in Water | 2008 | R | Blackburn with Darwen BC |  | 12 | 31.0479 | 25.9885 | 19.47% | P | 0.83705
Gradko | 50% TEA in Acetone | 2008 | R | Horsham DC |  | 11 | 33.7879 | 29.9909 | 12.66% | G | 0.88762
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | R | LB Islington |  | 11 | 64.3333 | 62.8818 | 2.308% | G | 0.97744
Gradko | 50% TEA in Water | 2008 | R | Thurrock Council |  | 12 | 40.1333 | 35.7283 | 12.33% | G | 0.89024
Environmental Scientific Groups | 10% TEA in Water | 2008 | B | Wigan MBC |  | 12 | 27.9583 | 26.1667 | 6.847% | G | 0.93592
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | O | Reigate & Banstead BC |  | 12 | 24.7222 | 27.0898 | -8.74% | P | 1.0958
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | O | Reigate & Banstead BC |  | 11 | 32.7879 | 32.67 | 0.3609% | P | 0.9964
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | O | Reigate & Banstead BC |  | 12 | 19.6389 | 19.124 | 2.692% | P | 0.97378
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 | UB | Leeds CC |  | 10 | 31.95 | 23.74 | 34.58% | G | 0.74304
Staffordshire Scientific Services | 50% TEA in Water | 2008 | UC | Stoke-on-Trent City Council |  | 12 | 28.1194 | 27.5045 | 2.236% | G | 0.97813
Staffordshire Scientific Services | 50% TEA in Water | 2008 | R | Stoke-on-Trent City Council |  | 12 | 46.6694 | 39.3855 | 18.49% | G | 0.84392
Tayside SS | 20% TEA in Water | 2007 | K | Dundee CC |  | 11 | 58.3727 | 53.7525 | 8.595% | G | 0.92085
Tayside SS | 20% TEA in Water | 2007 | K | Dundee CC |  | 12 | 55.2583 | 49.1548 | 12.42% | G | 0.88955
Tayside SS | 20% TEA in Water | 2007 | R | Dundee CC |  | 12 | 50.4333 | 35.7486 | 41.08% | G | 0.70883
Tayside SS | 20% TEA in Water | 2007 | K | Dundee CC |  | 11 | 44.103 | 41.4354 | 6.438% | G | 0.93951
Harwell Scientific Services | 50% TEA in Acetone | 2008 | R | Adur DC |  | 11 | 43.9576 | 31.1273 | 41.22% | G | 0.70812
Gradko | 20% TEA in Water | 2008 | R | Central Bedfordshire |  | 11 | 43.9561 | 37.5091 | 17.19% | G | 0.85333
Gradko | 20% TEA in Water | 2008 | R | Cheshire East Council |  | 12 | 55.3278 | 45.3025 | 22.13% | G | 0.8188
Gradko | 20% TEA in Water | 2008 | R | Cheshire East Council |  | 12 | 28.0346 | 28.0167 | 0.06419% | G | 0.99936
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | LB Croydon |  | 12 | 66.5537 | 67.565 | -1.497% | G | 1.0152
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | LB Croydon |  | 12 | 49.5806 | 49.4458 | 0.2725% | P | 0.99728
Environmental Scientific Groups | 10% TEA in Water | 2008 | B | Derby CC |  | 12 | 37.5556 | 33.7417 | 11.3% | G | 0.89845
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | Derby CC |  | 12 | 53.9444 | 40.625 | 32.79% | P | 0.75309
Environmental Scientific Groups | 10% TEA in Water | 2008 | R | Derby CC |  | 12 | 53.9167 | 45.95 | 17.34% | G | 0.85224
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | UB | Elmbridge BC |  | 10 | 21.6333 | 20.7 | 4.509% | P | 0.95686
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | R | Elmbridge BC |  | 12 | 50.3333 | 44.6667 | 12.69% | G | 0.88742
Gradko | 20% TEA in Water | 2008 | R | Gedling BC |  | 9 | 43.7504 | 34.4444 | 27.02% | G | 0.78729
Gradko | 50% TEA in Acetone | 2008 | R | LB Lewisham |  | 10 | 68.5953 | 63.4782 | 8.061% | G | 0.9254
Harwell Scientific Services | 50% TEA in Acetone | 2008 | UB | Medway Council |  | 11 | 28.6364 | 23.3636 | 22.57% | G | 0.81587
Harwell Scientific Services | 50% TEA in Acetone | 2008 | R | Medway Council |  | 12 | 41.7639 | 32.5833 | 28.18% | G | 0.78018
Harwell Scientific Services | 50% TEA in Acetone | 2008 | Rural | Medway Council |  | 12 | 29.4722 | 18 | 63.73% | G | 0.61074
Gradko | 20% TEA in Water | 2008 | O | North Warwickshire BC |  | 10 | 51.628 | 46.297 | 11.51% | G | 0.89674
Gradko | 50% TEA in Acetone | 2008 | UC | Sheffield CC |  | 12 | 33.1708 | 29.9392 | 10.79% | G | 0.90258
South Yorkshire Labs | 50% TEA in Acetone | 2008 | UI | Sheffield CC |  | 12 | 41.5556 | 40.6242 | 2.293% | G | 0.97759
South Yorkshire Labs | 50% TEA in Acetone | 2008 | UC | Sheffield CC |  | 12 | 34.125 | 32.795 | 4.055% | G | 0.96103
South Yorkshire Labs | 50% TEA in Acetone | 2008 | UC | Sheffield CC |  | 12 | 36.0556 | 29.9392 | 20.43% | G | 0.83036
Cardiff Scientific Services | 50% TEA in Acetone | 2008 | R | Vale of Glamorgan |  | 11 | 28.303 | 23.6091 | 19.88% | G | 0.83415
Cardiff Scientific Services | 50% TEA in Acetone | 2008 | O | Vale of Glamorgan |  | 9 | 14.7407 | 11.9556 | 23.3% | P | 0.81106
Cardiff Scientific Services | 50% TEA in Acetone | 2008 | R | Vale of Glamorgan |  | 11 | 24.7576 | 25.2909 | -2.109% | G | 1.0215
Lambeth Scientific Services | 50% TEA in Acetone | 2008 | UB | City of York Council |  | 12 | 20.4264 | 22.1878 | -7.939% | P | 1.0862
Gradko | 20% TEA in Water | 2008 | R | Newtownabbey BC |  | 12 | 37.8611 | 35.6667 | 6.153% | P | 0.94204
Gradko | 50% TEA in Acetone | 2007 | R | LB Richmond |  | 12 | 43.41 | 43.25 | 0.3699% | G | 0.99631
Gradko | 50% TEA in Acetone | 2008 | R | LB Richmond |  | 11 | 27.3694 | 44.6455 | -38.7% | P | 1.6312
Gradko | 50% TEA in Acetone | 2007 | B | LB Richmond |  | 11 | 28.09 | 30.7545 | -8.664% | G | 1.0949
Gradko | 50% TEA in Acetone | 2008 | B | LB Richmond |  | 11 | 44.6203 | 29.2455 | 52.57% | G | 0.65543
Tayside SS | 20% TEA in Water | 2008 | R | Dundee CC |  | 12 | 51.3417 | 43.025 | 19.33% | G | 0.83801
Tayside SS | 20% TEA in Water | 2008 | K | Dundee CC |  | 12 | 59.0611 | 52.7417 | 11.98% | G | 0.893
Tayside SS | 20% TEA in Water | 2008 | K | Dundee CC |  | 10 | 58.5667 | 51.93 | 12.78% | G | 0.88668
Gradko | 20% TEA in Water | 2009 | R | Lancaster CC |  | 12 | 36.8889 | 31.0074 | 18.97% | G | 0.84056
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Falkirk Council |  | 12 | 38.7333 | 26.7823 | 44.62% | G | 0.69145
Harwell Scientific Services | 50% TEA in Acetone | 2009 | UB | Falkirk Council |  | 12 | 27.2667 | 22.5771 | 20.77% | G | 0.82801
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Hambleton DC |  | 11 | 27.8879 | 22.686 | 22.93% | G | 0.81347
Harwell Scientific Services | 50% TEA in Acetone | 2009 | UB | Medway Council |  | 9 | 23.0463 | 20.5556 | 12.12% | G | 0.89192
Harwell Scientific Services | 50% TEA in Acetone | 2009 | Rural | Medway Council |  | 10 | 26.1533 | 16.6 | 57.55% | G | 0.63472
Harwell Scientific Services | 50% TEA in Acetone | 2008 | I | Suffolk Coastal DC |  | 12 | 55.0694 | 42.3333 | 30.09% | G | 0.76873
Harwell Scientific Services | 50% TEA in Acetone | 2008 | K | Suffolk Coastal DC |  | 12 | 50.8 | 45.6667 | 11.24% | G | 0.89895
Aberdeen CC | 20% TEA in Water | 2008 | K | Aberdeen CC |  | 12 | 26.3889 | 26.0142 | 1.44% | G | 0.9858
Aberdeen CC | 20% TEA in Water | 2008 | UB | Aberdeen CC |  | 12 | 27.9167 | 24.5058 | 13.92% | G | 0.87782
Aberdeen CC | 20% TEA in Water | 2008 | K | Aberdeen CC |  | 10 | 76.0333 | 74.404 | 2.19% | G | 0.97857
Aberdeen CC | 20% TEA in Water | 2008 | K | Aberdeen CC |  | 11 | 57.1515 | 54.5164 | 4.834% | G | 0.95389
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Adur DC |  | 9 | 53.3444 | 45.7 | 16.73% | G | 0.8567
Harwell Scientific Services | 50% TEA in Acetone | 2009 | UB | Ashford BC |  | 12 | 22.8361 | 22.25 | 2.634% | P | 0.97433
Gradko | 50% TEA in Acetone | 2009 | R | Boston BC |  | 11 | 45.097 | 33.3636 | 35.17% | G | 0.73982
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | R | Calderdale MBC |  | 12 | 62.0417 | 48.8417 | 27.03% | G | 0.78724
Bristol Scientific Services | 20% TEA in Water | 2008 | UB | Cheltenham BC |  | 12 | 25.4056 | 21.5833 | 17.71% | G | 0.84955
Gradko | 20% TEA in Water | 2009 | R | Cheshire East Council |  | 12 | 33.9458 | 20.9417 | 62.1% | G | 0.61691
Gradko | 20% TEA in Water | 2008 | R | Dartford Council |  | 12 | 60.9889 | 57.9167 | 5.305% | P | 0.94963
Gradko | 20% TEA in Water | 2009 | R | Dartford Council |  | 12 | 60.4917 | 57.8333 | 4.597% | G | 0.95605
Glasgow Scientific Services | 20% TEA in Water | 2009 | R | East Dunbartonshire Council |  | 12 | 27.5333 | 33.4417 | -17.67% | P | 1.2146
Glasgow Scientific Services | 20% TEA in Water | 2009 | R | East Dunbartonshire Council |  | 12 | 34.7667 | 39.6417 | -12.3% | G | 1.1402
Glasgow Scientific Services | 20% TEA in Water | 2009 | R | East Dunbartonshire Council |  | 11 | 30.1788 | 42.6182 | -29.19% | P | 1.4122
Gradko | 50% TEA in Acetone | 2008 | R | East Hampshire DC |  | 12 | 28.6807 | 27.5917 | 3.947% | G | 0.96203
Gradko | 50% TEA in Acetone | 2009 | R | East Hampshire DC |  | 12 | 26.9319 | 24.8333 | 8.451% | G | 0.92208
Gradko | 20% TEA in Water | 2009 | R | Fareham BC |  | 11 | 43.6883 | 36.0909 | 21.05% | G | 0.8261
Gradko | 20% TEA in Water | 2009 | R | Gedling BC |  | 12 | 45.0464 | 36 | 25.13% | G | 0.79918
South Yorkshire Labs | 50% TEA in Acetone | 2009 | UC | Kingston upon Hull City Council |  | 11 | 31.8485 | 26.9091 | 18.36% | G | 0.84491
Gradko | 20% TEA in Water | 2009 | UB | East Hertfordshire DC |  | 9 | 29.3333 | 32.9444 | -10.96% | G | 1.1231
Gradko | 20% TEA in Water | 2009 | R | Exeter CC |  | 12 | 37.9104 | 39.4975 | -4.018% | G | 1.0419
Gradko | 50% TEA in Acetone | 2009 | B | LB Brent |  | 10 | 31.9879 | 31.16 | 2.657% | G | 0.97412
Lambeth Scientific Services | 50% TEA in Acetone | 2009 | R | LB Haringey |  | 10 | 46.3 | 41.52 | 11.51% | S | 0.89676
Gradko | 50% TEA in Acetone | 2009 | R | LB Richmond |  | 12 | 42.7023 | 42.8417 | -0.3253% | G | 1.0033
Gradko | 50% TEA in Acetone | 2009 | S | LB Richmond |  | 12 | 27.2517 | 27.9167 | -2.382% | G | 1.0244
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | B | Leeds CC |  | 11 | 30.3636 | 28.5091 | 6.505% | G | 0.93892
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | UC | Leeds CC |  | 10 | 68.9 | 64.45 | 6.905% | S | 0.93541
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | R | Leeds CC |  | 12 | 51.25 | 45.15 | 13.51% | G | 0.88098
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | R | Leeds CC |  | 10 | 57.8 | 43.2 | 33.8% | S | 0.7474
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | UC | Leeds CC |  | 12 | 42.3056 | 34 | 25.2% | G | 0.8
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | R | Leeds CC |  | 11 | 42.9091 | 37.6909 | 13.84% | G | 0.87839
Cardiff Scientific Services | 50% TEA in Acetone | 2009 | R | Monmouthshire CC |  | 12 | 48.6111 | 38.9602 | 24.77% | G | 0.80147
Cardiff Scientific Services | 50% TEA in Acetone | 2009 | UB | Monmouthshire CC |  | 12 | 25.1944 | 21.0037 | 19.95% | G | 0.83366
Gradko | 20% TEA in Water | 2009 | K | New Forest DC |  | 12 | 53.9803 | 44.9667 | 20.05% | G | 0.83302
Gradko | 20% TEA in Water | 2009 | R | New Forest DC |  | 12 | 37.2917 | 26.1833 | 42.43% | G | 0.70212
Environmental Scientific Groups | 20% TEA in Water | 2009 | S | Chichester DC |  | 11 | 43.7778 | 34 | 30.2% | G | 0.76504
Northampton BC | 20% TEA in Water | 2009 | B | Northampton BC |  | 12 | 26.8589 | 20.5 | 31.02% | G | 0.76325
Northampton BC | 20% TEA in Water | 2009 | R | Northampton BC |  | 9 | 49.9889 | 32.9667 | 51.63% | G | 0.65948
Bristol Scientific Services | 20% TEA in Water | 2009 | Rural | Pembrokeshire CC |  | 12 | 6.9861 | 5.7375 | 21.76% | P | 0.82127
South Yorkshire Labs | 50% TEA in Acetone | 2009 | R | Rotherham MBC |  | 10 | 35.2 | 31.69 | 11.08% | G | 0.90028
Gradko | 20% TEA in Water | 2009 | K | South Lakeland DC |  | 10 | 46.349 | 36.277 | 27.76% | P | 0.78269
Staffordshire Scientific Services | 20% TEA in Water | 2009 | O | South Staffordshire DC |  | 12 | 38.2722 | 34.1583 | 12.04% | G | 0.89251
Gradko | 50% TEA in Acetone | 2009 | R | Stevenage BC |  | 12 | 38.4589 | 29.125 | 32.05% | G | 0.7573
Staffordshire Scientific Services | 20% TEA in Water | 2009 | UB | Stoke-on-Trent City Council |  | 9 | 35.3037 | 30.3153 | 16.45% | G | 0.8587
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Vale of the White Horse DC |  | 9 | 34.1519 | 29.6667 | 15.12% | G | 0.86867
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | B | Wakefield MDC |  | 9 | 34.5185 | 38.2 | -9.637% | G | 1.1067
Edinburgh Scientific Services | 50% TEA in Acetone | 2008 | R | West Lothian Council |  | 12 | 22.3889 | 19.8167 | 12.98% | G | 0.88511
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | R | West Lothian Council |  | 12 | 22.8139 | 21.1167 | 8.037% | G | 0.92561
Harwell Scientific Services | 50% TEA in Acetone | 2008 | K | West Oxfordshire DC |  | 12 | 53.8 | 39.8 | 35.18% | G | 0.73978
Gradko | 20% TEA in Water | 2009 | R | Carlisle CC |  | 9 | 37.7341 | 31.0521 | 21.52% | G | 0.82292
Gradko | 20% TEA in Water | 2009 | R | Newtownabbey BC |  | 12 | 36.7222 | 34.3333 | 6.958% | G | 0.93495
Gradko | 20% TEA in Water | 2009 | UC | Nottingham CC |  | 12 | 36.4722 | 33.7017 | 8.221% | G | 0.92404
Gradko | 20% TEA in Water | 2009 | R | Nottingham CC |  | 12 | 45.437 | 40.6247 | 11.85% | G | 0.89409
Gradko | 20% TEA in Water | 2009 | R | Nottingham CC |  | 11 | 44.7407 | 40.8911 | 9.414% | G | 0.91396
Tayside SS | 20% TEA in Water | 2009 | R | Fife Council |  | 11 | 38.3545 | 27.9734 | 37.11% | G | 0.72934
Tayside SS | 20% TEA in Water | 2009 | R | Fife Council |  | 10 | 32.9617 | 26.494 | 24.41% | G | 0.80378
Staffordshire Scientific Services | 20% TEA in Water | 2009 | R | Cannock Chase Council |  | 12 | 49.2611 | 43.075 | 14.36% | G | 0.87442
Staffordshire Scientific Services | 20% TEA in Water | 2009 | S | Cannock Chase Council |  | 12 | 27.0667 | 21.75 | 24.44% | G | 0.80357
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Cambridge CC |  | 12 | 50.3472 | 41.9717 | 19.96% | G | 0.83364
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | UC | Leeds CC |  | 11 | 50.2727 | 41.7273 | 20.48% | S | 0.83002
Staffordshire Scientific Services | 20% TEA in Water | 2009 | R | Stoke-on-Trent City Council |  | 12 | 41.7139 | 30.1178 | 38.5% | G | 0.72201
Staffordshire Scientific Services | 20% TEA in Water | 2009 | R | Stoke-on-Trent City Council |  | 12 | 46.9806 | 38.2693 | 22.76% | G | 0.81458
Staffordshire Scientific Services | 20% TEA in Water | 2009 | R | Stoke-on-Trent City Council |  | 10 | 57.2233 | 41.0528 | 39.39% | G | 0.71741
Staffordshire Scientific Services | 20% TEA in Water | 2009 | R | East Staffordshire BC |  | 12 | 35.3125 | 24.775 | 42.53% | G | 0.70159
South Yorkshire Labs | 50% TEA in Acetone | 2009 | R | Barnsley MBC |  | 10 | 42.9 | 41.2 | 4.126% | G | 0.96037
Gradko | 20% TEA in Water | 2009 | UC | Belfast CC |  | 10 | 38.8987 | 34.002 | 14.4% | G | 0.87412
South Yorkshire Labs | 50% TEA in Acetone | 2009 | UB | Bournemouth BC |  | 12 | 15.2778 | 15.6425 | -2.332% | G | 1.0239
Gradko | 20% TEA in Water | 2009 | R | Bromsgrove DC |  | 11 | 54.5048 | 52.7623 | 3.303% | P | 0.96803
Harwell Scientific Services | 50% TEA in Acetone | 2009 | B | Canterbury CC |  | 11 | 19.8485 | 15.6727 | 26.64% | G | 0.78962
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Canterbury CC |  | 12 | 48.5167 | 37.3 | 30.07% | G | 0.76881
Environmental Scientific Groups | 20% TEA in Water | 2009 | R | Castlereagh BC |  | 12 | 48.7778 | 39.4183 | 23.74% | G | 0.80812
Environmental Scientific Groups | 20% TEA in Water | 2009 | R | Castlereagh BC |  | 12 | 31.2778 | 24.5533 | 27.39% | G | 0.78501
Gradko | 20% TEA in Water | 2009 | R | Chelmsford BC |  | 10 | 39.2397 | 35.849 | 9.458% | G | 0.91359
Gradko | 20% TEA in Water | 2009 | R | Coventry CC |  | 11 | 44.8889 | 43.6455 | 2.849% | P | 0.9723
Gradko | 20% TEA in Water | 2009 | R | Coventry CC |  | 11 | 37.5777 | 29.9273 | 25.56% | P | 0.79641
Gradko | 20% TEA in Water | 2009 | R | Coventry CC |  | 12 | 36.6303 | 35.875 | 2.105% | G | 0.97938
Gradko | 20% TEA in Water | 2009 | R | Coventry CC |  | 9 | 50.7656 | 65.0556 | -21.97% | G | 1.2815
Gradko | 20% TEA in Water | 2009 | R | Dudley MBC |  | 11 | 41.9421 | 37.0818 | 13.11% | G | 0.88412
Gradko | 20% TEA in Water | 2009 | B | Dudley MBC |  | 12 | 29.635 | 27.0917 | 9.388% | G | 0.91418
Gradko | 20% TEA in Water | 2009 | Rural | Dudley MBC |  | 12 | 18.8969 | 17 | 11.16% | G | 0.89962
Gradko | 20% TEA in Water | 2009 | R | Dudley MBC |  | 12 | 44.4836 | 39.95 | 11.35% | G | 0.89808
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Gateshead Council |  | 10 | 39.1467 | 33.8 | 15.82% | G | 0.86342
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Gateshead Council |  | 10 | 37.6733 | 27.57 | 36.65% | P | 0.73182
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Gateshead Council |  | 10 | 40.88 | 29.67 | 37.78% | G | 0.72578
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Gateshead Council |  | 9 | 34.7111 | 30.3333 | 14.43% | P | 0.87388
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | R | Leeds CC |  | 9 | 42 | 34.7111 | 21% | S | 0.82646
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | R | Leeds CC |  | 9 | 50.3333 | 40.8889 | 23.1% | S | 0.81236
Environmental Scientific Groups | 20% TEA in Water | 2009 | R | Lisburn CC |  | 12 | 32.0833 | 26.4817 | 21.15% | P | 0.8254
Milton Keynes Council | 20% TEA in Water | 2009 | UC | Milton Keynes Council |  | 12 | 31.008 | 22.26 | 39.3% | G | 0.71788
Milton Keynes Council | 20% TEA in Water | 2009 | R | Milton Keynes Council |  | 12 | 45.0917 | 34.665 | 30.08% | G | 0.76877
Environmental Scientific Groups | 20% TEA in Water | 2009 | R | North Down BC |  | 12 | 48.5556 | 35.6192 | 36.32% | G | 0.73358
Gradko | 20% TEA in Water | 2008 | R | Rhondda Cynon Taf CBC |  | 12 | 36.1486 | 39.9575 | -9.532% | G | 1.1054
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Rugby BC |  | 12 | 35.1361 | 33.8426 | 3.822% | G | 0.96319
Gradko | 50% TEA in Acetone | 2009 | R | Sandwell MBC |  | 12 | 45.2964 | 43.9766 | 3.001% | G | 0.97086
Gradko | 20% TEA in Water | 2009 | R | Sandwell MBC |  | 12 | 47.1192 | 43.9766 | 7.146% | S | 0.93331
Gradko | 20% TEA in Water | 2009 | UB | Sandwell MBC |  | 10 | 19.2377 | 16.1039 | 19.46% | S | 0.8371
Gradko | 50% TEA in Acetone | 2009 | UB | Sandwell MBC |  | 11 | 16.5107 | 16.7834 | -1.625% | S | 1.0165
Gradko | 20% TEA in Water | 2009 | UB | Sandwell MBC |  | 12 | 28.9838 | 27.374 | 5.881% | S | 0.94446
Gradko | 50% TEA in Acetone | 2009 | UB | Sandwell MBC |  | 11 | 26.5009 | 28.3824 | -6.629% | G | 1.071
Gradko | 20% TEA in Water | 2009 | R | Sandwell MBC |  | 11 | 41.8069 | 39.5229 | 5.779% | S | 0.94537
Gradko | 50% TEA in Acetone | 2009 | R | Sandwell MBC |  | 12 | 38.3861 | 39.821 | -3.603% | S | 1.0374
Harwell Scientific Services | 50% TEA in Acetone | 2009 | I | Swale BC |  | 12 | 26.4325 | 22.25 | 18.8% | G | 0.84177
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Swale BC |  | 12 | 44.4325 | 30.4167 | 46.08% | P | 0.68456
Gradko | 50% TEA in Acetone | 2009 | UB | Sheffield CC |  | 10 | 32.8333 | 37.7 | -12.91% | G | 1.1482
South Yorkshire Labs | 50% TEA in Acetone | 2009 | UB | Sheffield CC |  | 11 | 36 | 36.3636 | -1% | G | 1.0101
South Yorkshire Labs | 50% TEA in Acetone | 2009 | UI | Sheffield CC |  | 10 | 45.7667 | 34.7 | 31.89% | G | 0.75819
South Yorkshire Labs | 50% TEA in Acetone | 2009 | R | Sheffield CC |  | 10 | 36.6 | 35.4 | 3.39% | G | 0.96721
Gradko | 50% TEA in Acetone | 2009 | UC | Uttlesford DC |  | 9 | 24.2841 | 24.7111 | -1.728% | G | 1.0176
Cardiff Scientific Services | 50% TEA in Acetone | 2009 | R | Vale of Glamorgan |  | 12 | 28.0833 | 24.5833 | 14.24% | G | 0.87537
Cardiff Scientific Services | 50% TEA in Acetone | 2009 | R | Vale of Glamorgan |  | 12 | 15.5833 | 11 | 41.67% | G | 0.70588
Cardiff Scientific Services | 50% TEA in Acetone | 2009 | R | Vale of Glamorgan |  | 9 | 26.7407 | 26.8889 | -0.551% | G | 1.0055
Walsall MBC | 50% TEA in Acetone | 2009 | R | Walsall MBC |  | 9 | 59.1662 | 55.3724 | 6.852% | P | 0.93588
Walsall MBC | 50% TEA in Acetone | 2009 | O | Walsall MBC |  | 10 | 36.216 | 34.5412 | 4.849% | P | 0.95375
Walsall MBC | 50% TEA in Acetone | 2009 | I | Walsall MBC |  | 11 | 27.3394 | 25.7561 | 6.147% | P | 0.94209
Gradko | 50% TEA in Acetone | 2009 | R | West Berkshire Council |  | 12 | 45.02 | 53.55 | -15.93% | P | 1.1895
Environmental Scientific Groups | 20% TEA in Water | 2009 | R | Wrexham CBC |  | 11 | 26.3611 | 23.2436 | 13.41% | G | 0.88174
Gradko | 20% TEA in Water | 2009 | R | Rushmoor BC |  | 10 | 35.1767 | 33.1367 | 6.156% | G | 0.94201
Environmental Scientific Groups | 20% TEA in Water | 2009 | R | Horsham DC |  | 10 | 36.2833 | 32 | 13.39% | G | 0.88195
Aberdeen CC | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 9 | 125.1704 | 105.6667 | 18.46% | G | 0.84418
Bristol Scientific Services | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 11 | 125.4939 | 106.6364 | 17.68% | G | 0.84973
Environmental Scientific Groups | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 12 | 131.7639 | 107.25 | 22.86% | G | 0.81396
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 10 | 111.4373 | 107.1 | 4.05% | G | 0.96108
Glasgow Scientific Services | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 11 | 92.0182 | 107.6364 | -14.51% | G | 1.1697
Gradko | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 12 | 106.2881 | 107.25 | -0.8969% | G | 1.0091
Harwell Scientific Services | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 11 | 129.497 | 107.6364 | 20.31% | G | 0.83119
Kent Scientific Services | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 12 | 129.863 | 107.25 | 21.08% | G | 0.82587
Kirklees Council Scientific Services | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 12 | 126.7058 | 107.25 | 18.14% | G | 0.84645
Lambeth Scientific Services | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 11 | 91.0303 | 107.5455 | -15.36% | G | 1.1814
Lancashire CC | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 12 | 114.7917 | 107.5 | 6.783% | G | 0.93648
Milton Keynes Council | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 9 | 119.0556 | 108.4444 | 9.785% | G | 0.91087
Northampton BC | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 11 | 146.6061 | 106.6364 | 37.48% | G | 0.72737
Staffordshire Scientific Services | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 12 | 115.9381 | 107.25 | 8.101% | G | 0.92506
South Yorkshire Labs | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 12 | 129.7222 | 107.5 | 20.67% | G | 0.82869
Walsall MBC | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 9 | 71.1185 | 107 | -33.53% | P | 1.5045
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | K | AEA Tech Intercomparison |  | 10 | 135.9 | 107.5 | 26.42% | G | 0.79102
Tayside SS | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 11 | 138.6176 | 107.9091 | 28.46% | G | 0.77847
Gradko | 20% TEA in Water | 2009 | K | AEA Tech Intercomparison |  | 12 | 120.8019 | 107.25 | 12.64% | G | 0.88782
Walsall MBC | 50% TEA in Acetone | 2009 | UB | Warwick |  | 9 | 19.2037 | 30.3022 | -36.63% | P | 1.5779
Walsall MBC | 50% TEA in Acetone | 2009 | R | Warwick |  | 11 | 46.6364 | 63.35 | -26.38% | P | 1.3584
Gradko | 20% TEA in Water | 2009 | R | Cheshire West & Chester Council |  | 11 | 41.1182 | 37.397 | 9.951% | G | 0.9095
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 | R | Leeds CC |  | 10 | 37.5 | 30.47 | 23.07% | S | 0.81253
Lambeth Scientific Services | 50% TEA in Acetone | 2009 | O | Reigate & Banstead BC |  | 12 | 24.75 | 25.107 | -1.422% | P | 1.0144
Lambeth Scientific Services | 50% TEA in Acetone | 2009 | O | Reigate & Banstead BC |  | 12 | 18.2778 | 17.8913 | 2.16% | G | 0.97886
Lambeth Scientific Services | 50% TEA in Acetone | 2009 | O | Reigate & Banstead BC |  | 11 | 29.1212 | 30.4718 | -4.432% | G | 1.0464
Gradko | 50% TEA in Acetone | 2009 | R | LB Lewisham |  | 10 | 73.595 | 61.709 | 19.26% | G | 0.83849
Bristol Scientific Services | 20% TEA in Water | 2009 | R | Wiltshire Council |  | 11 | 48.8576 | 38.8718 | 25.69% | G | 0.79561
Bristol Scientific Services | 20% TEA in Water | 2009 | R | Wiltshire Council |  | 10 | 54.02 | 37.855 | 42.7% | G | 0.70076
Tayside SS | 20% TEA in Water | 2009 | K | Dundee CC |  | 10 | 64.77 | 56.14 | 15.37% | G | 0.86676
Tayside SS | 20% TEA in Water | 2009 | R | Dundee CC |  | 12 | 54.5333 | 44.875 | 21.52% | G | 0.82289
Tayside SS | 20% TEA in Water | 2009 | K | Dundee CC |  | 12 | 46.8972 | 37.775 | 24.15% | G | 0.80548
Harwell Scientific Services | 50% TEA in Acetone | 2009 | I | Suffolk Coastal DC |  | 10 | 51.6733 | 45.5 | 13.57% | G | 0.88053
Harwell Scientific Services | 50% TEA in Acetone | 2009 | K | Suffolk Coastal DC |  | 12 | 49.6111 | 44.5 | 11.49% | G | 0.89698
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | Dover DC |  | 12 | 46.4972 | 39.5833 | 17.47% | G | 0.85131
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | City of York Council |  | 12 | 37.5262 | 29.371 | 27.77% | G | 0.78268
Harwell Scientific Services | 50% TEA in Acetone | 2009 | UB | City of York Council |  | 12 | 27.1512 | 23.6967 | 14.58% | G | 0.87277
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | City of York Council |  | 12 | 41.8343 | 36.5144 | 14.57% | G | 0.87283
Harwell Scientific Services | 50% TEA in Acetone | 2009 | R | City of York Council |  | 12 | 43.7602 | 37.2323 | 17.53% | G | 0.85082
Gradko | 50% TEA in Acetone | 2009 | R | Bedford BC |  | 12 | 40.4122 | 39.525 | 2.245% | G | 0.97805
Environmental Scientific Groups | 20% TEA in Water | 2009 | R | Dumfries and Galloway Council |  | 10 | 40.9333 | 34.8869 | 17.33% | G | 0.85228
Harwell Scientific Services | 50% TEA in Acetone | 2009 | K | West Oxfordshire DC |  | 11 | 56.3758 | 39.7545 | 41.81% | G | 0.70517
Gradko | 20% TEA in Water | 2009 | R | Winchester | 9 | 9 | 41 | 47 | -13.2% | G | 1.15
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | K | Edinburgh CC |  | 12 | 76.0917 | 69.7417 | 9.105% | G | 0.91655
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | R | Edinburgh CC |  | 12 | 40.1125 | 33.325 | 20.37% | G | 0.83079
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | R | Edinburgh CC |  | 12 | 31.5181 | 25.6917 | 22.68% | G | 0.81514
South Yorkshire Labs | 50% TEA in Acetone | 2009 | UC | Oxford City Council |  | 10 | 51.25 | 48.56 | 5.54% | G | 0.95
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | R | Leeds CC |  | 12 | 50.7083 | 45.925 | 10.42% | G | 0.90567
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | R | Leeds CC |  | 10 | 58.6 | 49.36 | 18.72% | S | 0.84232
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | UR | Leeds CC |  | 11 | 31 | 33.9818 | -8.775% | S | 1.0962
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | UB | Leeds CC |  | 11 | 46.9394 | 37.0636 | 26.65% | G | 0.78961
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | UC | Leeds CC |  | 10 | 51.6 | 42.908 | 20.26% | S | 0.83155
Bristol Scientific Services | 20% TEA in Water | 2010 | R | Wiltshire Council |  | 12 | 40.2778 | 34.5833 | 16.47% | G | 0.85862
Gradko | 50% TEA in Acetone | 2010 | R | Reading BC |  | 12 | 40.0806 | 46.2 | -13.25% | G | 1.1527
South Yorkshire Labs | 50% TEA in Acetone | 2010 | R | Barnsley MBC |  | 10 | 44.9333 | 41.2 | 9.061% | G | 0.91691
Gradko | 50% TEA in Acetone | 2010 | R | East Hampshire DC |  | 11 | 26.6527 | 25.0182 | 6.533% | G | 0.93867
Bristol Scientific Services | 20% TEA in Water | 2010 | R | Wiltshire Council |  | 9 | 50.1111 | 40.1111 | 24.93% | G | 0.80044
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | R | Calderdale MBC |  | 11 | 65.697 | 54.5818 | 20.36% | G | 0.83081
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | R | Vale of Glamorgan |  | 12 | 29.5278 | 26.5242 | 11.32% | G | 0.89828
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | O | Vale of Glamorgan |  | 10 | 18.9667 | 12.406 | 52.88% | G | 0.65409
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | R | Vale of Glamorgan |  | 10 | 31.7333 | 31.4443 | 0.9193% | G | 0.99089
Gradko | 20% TEA in Water | 2010 | UC | Nottingham CC |  | 12 | 39.6481 | 40.05 | -1.004% | G | 1.0101
Gradko | 20% TEA in Water | 2010 | R | Nottingham CC |  | 12 | 45.4217 | 43.2833 | 4.94% | G | 0.95292
Gradko | 20% TEA in Water | 2010 | R | Nottingham CC |  | 12 | 43.7219 | 41.1167 | 6.336% | G | 0.94041
Gradko | 20% TEA in Water | 2010 | K | South Lakeland DC |  | 12 | 45.565 | 42.1458 | 8.113% | G | 0.92496
Bristol Scientific Services | 20% TEA in Water | 2010 | R | Wiltshire Council |  | 9 | 47.963 | 41.6667 | 15.11% | G | 0.86873
Gradko | 20% TEA in Water | 2010 | R | Gedling BC |  | 12 | 42.6061 | 39.0692 | 9.053% | G | 0.91699
Gradko | 20% TEA in Water | 2010 | B | East Hertfordshire DC |  | 12 | 20.9444 | 15.9583 | 31.24% | G | 0.76194
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Hambleton DC |  | 11 | 25.7394 | 17.5545 | 46.63% | G | 0.68201
Gradko | 50% TEA in Acetone | 2010 | R | Wolverhampton CC |  | 12 | 42.3264 | 40.6667 | 4.081% | G | 0.96079
Gradko | 50% TEA in Acetone | 2010 | R | Wolverhampton CC |  | 12 | 38.4647 | 38.1533 | 0.8162% | G | 0.9919
Gradko | 20% TEA in Water | 2010 | R | Dudley MBC |  | 12 | 42.1263 | 44.791 | -5.949% | G | 1.0633
Gradko | 20% TEA in Water | 2010 | B | Dudley MBC |  | 12 | 29.9369 | 29.9246 | 0.0411% | G | 0.99959
Gradko | 20% TEA in Water | 2010 | R | Dudley MBC |  | 10 | 46.6493 | 45.5502 | 2.413% | G | 0.97644
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | UC | Uttlesford DC |  | 9 | 33.3637 | 31.5067 | 5.894% | G | 0.94434
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | K | Caerphilly County BC |  | 12 | 54.6667 | 48.4167 | 12.91% | G | 0.88567
Gradko | 50% TEA in Acetone | 2010 | R | Exeter CC |  | 12 | 42.1042 | 39.875 | 5.59% | G | 0.94706
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Falkirk Council |  | 11 | 37.1545 | 31.3283 | 18.6% | P | 0.84319
Harwell Scientific Services | 50% TEA in Acetone | 2010 | UB | Falkirk Council |  | 10 | 26.6033 | 21.9737 | 21.07% | P | 0.82597
Staffordshire Scientific Services | 20% TEA in Water | 2010 | O | South Staffordshire DC |  | 11 | 41.1152 | 34.7109 | 18.45% | G | 0.84424
Gradko | 50% TEA in Acetone | 2010 | R | Lewisham Council |  | 10 | 74.2322 | 50.85 | 45.98% | G | 0.68501
Staffordshire Scientific Services | 20% TEA in Water | 2010 | UB | Stoke-on-Trent City Council |  | 12 | 37.5 | 36.9508 | 1.486% | G | 0.98535
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | UB | Wakefield MDC |  | 11 | 35.2121 | 33 | 6.703% | G | 0.93718
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | UB | Wakefield MDC |  | 9 | 43.9444 | 58.4321 | -24.79% | G | 1.3297
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | K | Wakefield MDC |  | 9 | 59.2222 | 57.7244 | 2.595% | G | 0.97471
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | R | Wakefield MDC |  | 12 | 44.75 | 35.1758 | 27.22% | G | 0.78605
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | UB | Wakefield MDC |  | 12 | 41.8056 | 37.5 | 11.48% | P | 0.89701
Bristol Scientific Services | 20% TEA in Water | 2010 | R | Wiltshire Council |  | 11 | 45.4848 | 36.1818 | 25.71% | G | 0.79547
Gradko | 50% TEA in Acetone | 2010 | B | LB Brent |  | 10 | 27.7901 | 28.199 | -1.45% | G | 1.0147
Bristol Scientific Services | 20% TEA in Water | 2010 | B | LB Waltham Forest |  | 12 | 40.375 | 37.8417 | 6.695% | S | 0.93725
Environmental Scientific Groups | 20% TEA in Water | 2010 | U | Chichester DC |  | 12 | 42.3588 | 27.9 | 51.82% | P | 0.65866
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Swale BC |  | 12 | 46.2639 | 39.1667 | 18.12% | G | 0.84659
Environmental Scientific Groups | 20% TEA in Water | 2010 | R | Wrexham CBC |  | 11 | 25.4803 | 23.7827 | 7.138% | P | 0.93338
Gradko | 20% TEA in Water | 2010 | R | Carlisle CC |  | 12 | 40.9046 | 32.5917 | 25.51% | G | 0.79677
Gradko | 50% TEA in Acetone | 2010 | R | Worthing BC |  | 10 | 44.4967 | 41.96 | 6.045% | G | 0.94299
Environmental Scientific Groups | 20% TEA in Water | 2010 | UB | Warrington BC |  | 12 | 28.2778 | 32.2583 | -12.34% | G | 1.1408
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | R | Boston BC |  | 10 | 57.2667 | 32.9 | 74.06% | G | 0.57451
Gradko | 50% TEA in Acetone | 2010 | B | LB Brent |  | 10 | 27.7901 | 28.199 | -1.45% | G | 1.0147
Northampton BC | 20% TEA in Water | 2010 | R | Northampton BC |  | 10 | 48.0867 | 33.31 | 44.36% | G | 0.69271
Northampton BC | 20% TEA in Water | 2010 | UB | Northampton BC |  | 12 | 26.2861 | 20.6083 | 27.55% | G | 0.784
Staffordshire Scientific Services | 20% TEA in Water | 2010 | R | Cannock Chase Council |  | 11 | 52.5879 | 43.5091 | 20.87% | G | 0.82736
Gradko | 20% TEA in Water | 2010 | UC | Nottingham CC |  | 12 | 39.6481 | 40.05 | -1.004% | G | 1.0101
Gradko | 20% TEA in Water | 2010 | R | Nottingham CC |  | 12 | 45.4217 | 43.2833 | 4.94% | G | 0.95292
Gradko | 20% TEA in Water | 2010 | R | Nottingham CC |  | 12 | 43.7219 | 41.1167 | 6.336% | G | 0.94041
South Yorkshire Labs | 50% TEA in Acetone | 2010 | UB | Bournemouth BC |  | 12 | 17.1111 | 17.0717 | 0.2311% | P | 0.99769
Gradko | 20% TEA in Water | 2010 | UC | Belfast CC |  | 11 | 38.447 | 35.4618 | 8.418% | P | 0.92236
Harwell Scientific Services | 50% TEA in Acetone | 2010 | UC | Dover DC |  | 12 | 44.2722 | 41 | 7.1% | G | 0.93
Environmental Scientific Groups | 20% TEA in Water | 2010 | R | Dumfries and Galloway Council |  | 12 | 43.345 | 39.839 | 8.8% | G | 0.91911
Gradko | 50% TEA in Acetone | 2010 | R | LB Richmond |  | 12 | 38.6214 | 40.95 | -5.686% | G | 1.0603
Gradko | 50% TEA in Acetone | 2010 | B | LB Richmond |  | 12 | 27.5308 | 26.2583 | 4.846% | G | 0.95378
Gradko | 20% TEA in Water | 2010 | UB | Luton Borough Council |  | 12 | 37.0218 | 33.7083 | 9.83% | G | 0.9105
South Yorkshire Labs | 50% TEA in Acetone | 2010 | UB | Kingston upon Hull City Council |  | 12 | 31.1944 | 27.3333 | 14.13% | G | 0.87622
Lambeth Scientific Services | 50% TEA in Acetone | 2010 | R | Waverley BC |  | 10 | 34.3 | 39.305 | -12.73% | P | 1.1459
Harwell Scientific Services | 50% TEA in Acetone | 2010 | B | Gravesham BC |  | 10 | 35.9767 | 27.3 | 31.78% | G | 0.75883
Staffordshire Scientific Services | 20% TEA in Water | 2010 | UC | Manchester CC |  | 12 | 48.4639 | 45.219 | 7.176% | G | 0.93304
Staffordshire Scientific Services | 20% TEA in Water | 2010 | S | Manchester CC |  | 12 | 31.5361 | 27.9114 | 12.99% | G | 0.88506
South Yorkshire Labs | 50% TEA in Acetone | 2010 | O | Kingston upon Hull City Council |  | 12 | 34.5139 | 29.1733 | 18.31% | P | 0.84526
Gradko | 50% TEA in Acetone | 2010 | UB | Reading BC |  | 9 | 20.337 | 25.5778 | -20.49% | G | 1.2577
Gradko | 20% TEA in Water | 2010 | K | Wandsworth Council |  | 11 | 135.7847 | 162.3364 | -16.36% | G | 1.1955
Harwell Scientific Services | 50% TEA in Acetone | 2010 | I | Swale BC |  | 10 | 25.0567 | 30.4 | -17.58% | G | 1.2132
Gradko | 20% Tea in Water | 2010 | UB | Wandsworth Council |  | 12 | 47.4306 | 52.6167 | -9.856% | P | 1.1093
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Tunbridge Wells BC |  | 12 | 66.6694 | 77.0833 | -13.51% | G | 1.1562
Harwell Scientific Services | 50% TEA in Acetone | 2010 | B | Canterbury CC |  | 12 | 20.7806 | 18 | 15.45% | G | 0.86619
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Canterbury CC |  | 12 | 47.8083 | 33.75 | 41.65% | G | 0.70594
Gradko | 20% TEA in Water | 2010 | R | Newtownabbey BC |  | 11 | 36.4242 | 34.9091 | 4.34% | G | 0.9584
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Gravesham BC |  | 11 | 41.8182 | 35.9091 | 16.46% | G | 0.8587
Environmental Scientific Groups | 20% TEA in Water | 2010 | R | North East Lincolnshire Council |  | 11 | 40.8121 | 34.6655 | 17.73% | G | 0.84939
Environmental Scientific Groups | 20% TEA in Water | 2010 | R | North East Lincolnshire Council |  | 10 | 47.9933 | 26.094 | 83.92% | P | 0.5437
Environmental Scientific Groups | 20% TEA in Water | 2010 | UB | North East Lincolnshire Council |  | 10 | 27.63 | 31.209 | -11.47% | P | 1.1295
Harwell Scientific Services | 50% TEA in Acetone | 2010 | UB | City of York Council |  | 12 | 26.1656 | 24.5358 | 6.642% | G | 0.93771
Gradko | 20% TEA in Water | 2010 | R | Cheshire East Council |  | 12 | 57.8444 | 44.7717 | 29.2% | G | 0.774
Environmental Scientific Groups | 20% TEA in Water | 2010 | R | Horsham DC |  | 12 | 36.3422 | 30.4817 | 19.23% | P | 0.83874
Gradko | 20% TEA in Water | 2010 | R | Cheshire East Council |  | 11 | 29.8485 | 30.3873 | -1.773% | G | 1.0181
Gradko | 20% TEA in Water | 2010 | R | Gateshead Council |  | 9 | 38.0952 | 34.4978 | 10.43% | G | 0.90557
Gradko | 20% TEA in Water | 2010 | R | Gateshead Council |  | 9 | 35.1767 | 33.1856 | 6% | G | 0.9434
Harwell Scientific Services | 50% TEA in Acetone | 2010 | B | Gravesham BC |  | 10 | 35.9767 | 27.3 | 31.78% | G | 0.75883
Gradko | 20% TEA in Water | 2010 | R | Gateshead Council |  | 9 | 34.0537 | 34.2667 | -0.6215% | G | 1.0063
Gradko | 20% TEA in Water | 2010 | R | Gateshead Council |  | 9 | 31.8363 | 35.3189 | -9.86% | G | 1.1094
Harwell Scientific Services | 50% TEA in Acetone | 2010 | B | Stockton on Tees |  | 12 | 30.0167 | 27.175 | 10.46% | G | 0.90533
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Stockton on Tees |  | 12 | 24.6833 | 21.0833 | 17.08% | G | 0.85415
Edinburgh Scientific Services | 50% TEA in Acetone | 2010 | R | West Lothian Council |  | 9 | 23.8778 | 26.2222 | -8.941% | G | 1.0982
Gradko | 20% TEA in Water | 2010 | R | Gosport BC |  | 10 | 31.386 | 23.2 | 35.28% | S | 0.73918
Lancashire CC | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 12 | 103.0919 | 93.2083 | 10.6% | P | 0.90413
Lambeth Scientific Services | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 11 | 90.7697 | 92.4091 | -1.774% | G | 1.0181
Harwell Scientific Services | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 11 | 119.9758 | 94.2273 | 27.33% | G | 0.78539
Milton Keynes Council | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 107.9814 | 93.2083 | 15.85% | G | 0.86319
South Yorkshire Labs | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 12 | 119.925 | 93.2083 | 28.66% | G | 0.77722
Edinburgh Scientific Services | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 12 | 98.5496 | 93.2083 | 5.73% | G | 0.9458
Bristol Scientific Services | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 118.5389 | 93.2083 | 27.18% | G | 0.78631
Northampton BC | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 127.9972 | 93.2083 | 37.32% | G | 0.72821
Glasgow Scientific Services | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 84.9917 | 93.2083 | -8.815% | G | 1.0967
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 12 | 129.8611 | 93.2083 | 39.32% | G | 0.71775
Aberdeen CC | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 114.3542 | 93.2083 | 22.69% | G | 0.81508
Staffordshire Scientific Services | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 118.7131 | 93.2083 | 27.36% | G | 0.78516
Kirklees Council Scientific Services | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 10 | 121.1136 | 94.85 | 27.69% | G | 0.78315
Tayside SS | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 127.4711 | 93.2083 | 36.76% | G | 0.73121
Kent Scientific Services | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 12 | 120 | 93.2083 | 28.9% | G | 0.78
Environmental Scientific Groups | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 11 | 117.2667 | 94.4091 | 24.21% | G | 0.80508
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | K | Marylebone Road Intercomparison |  | 11 | 118.8121 | 94.4091 | 25.85% | G | 0.79461
Harwell Scientific Services | 20% TEA in Water | 2010 | K | Marylebone Road Intercomparison |  | 9 | 117.7856 | 90.3889 | 30.31% | G | 0.7674
Gradko | 20% TEA in Water | 2010 | R | Rhondda Cynon Taf CBC |  | 10 | 35.0917 | 34.939 | 0.437% | G | 0.99565
Gradko | 20% TEA in Water | 2010 | O | North Warwickshire BC |  | 9 | 48.1241 | 42.3544 | 13.62% | P | 0.88011
Gradko | 20% TEA in Water | 2010 | UB | LB Ealing |  | 10 | 39.12 | 40.654 | -3.773% | G | 1.0392
Gradko | 20% TEA in Water | 2010 | R | South Norfolk Council |  | 9 | 28.3815 | 17.3333 | 63.74% | G | 0.61073
Lambeth Scientific Services | 50% TEA in Acetone | 2010 | R | Waverley BC |  | 9 | 38.4444 | 41.1111 | -6.486% | P | 1.0694
Gradko | 20% TEA in Water | 2010 | B | Chelmsford BC |  | 11 | 16.1502 | 17.0564 | -5.313% | G | 1.0561
Gradko | 20% TEA in Water | 2010 | R | Chelmsford BC |  | 12 | 33.1928 | 21.4108 | 55.03% | G | 0.64504
Gradko | 20% TEA in Water | 2010 | R | Chelmsford BC |  | 10 | 37.1602 | 32.423 | 14.61% | G | 0.87252
Gradko | 20% TEA in Water | 2010 | R | Wokingham BC |  | 10 | 37.279 | 35.8 | 4.131% | G | 0.96033
Gradko | 20% TEA in Water | 2010 | R | West Dunbartonshire Council |  | 9 | 22.0926 | 22.0667 | 0.1175% | G | 0.99883
Gradko | 20% TEA in Water | 2010 | R | Scarborough BC |  | 12 | 34.5808 | 29.25 | 18.23% | G | 0.84584
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Vale of White Horse DC |  | 12 | 39.0611 | 31.6083 | 23.58% | G | 0.8092
Environmental Scientific Groups | 20% TEA in Water | 2010 | R | North Down BC |  | 11 | 46.0318 | 34.16 | 34.75% | G | 0.7421
Bristol Scientific Services | 20% TEA in Water | 2010 | R | South Gloucestershire |  | 11 | 33.9499 | 31.1297 | 9.06% | G | 0.91693
Milton Keynes Council | 20% TEA in Water | 2010 | UC | Milton Keynes Council |  | 12 | 28.6699 | 22.9374 | 24.99% | G | 0.80005
Milton Keynes Council | 20% TEA in Water | 2010 | O | Milton Keynes Council |  | 9 | 36.5018 | 30.1227 | 21.18% | G | 0.82524
Milton Keynes Council | 20% TEA in Water | 2010 | UC | Milton Keynes Council |  | 12 | 28.6699 | 22.9374 | 24.99% | G | 0.80005
Milton Keynes Council | 20% TEA in Water | 2010 | R | Milton Keynes Council |  | 12 | 38.5572 | 34.1816 | 12.8% | G | 0.88652
Milton Keynes Council | 20% TEA in Water | 2010 | R | Milton Keynes Council |  | 12 | 44.8311 | 38.5511 | 16.29% | G | 0.85992
Harwell Scientific Services | 50% TEA in Acetone | 2010 | R | Thanet DC |  | 11 | 32 | 26 | 25.8% | G | 0.79
Gradko | 50% TEA in Acetone | 2010 | UB | Sandwell MBC |  | 12 | 26.8955 | 29.9499 | -10.2% | G | 1.1136
Gradko | 50% TEA in Acetone | 2010 | R | Sandwell MBC |  | 12 | 43.2571 | 46.6789 | -7.331% | G | 1.0791
Gradko | 50% TEA in Acetone | 2010 | R | Sandwell MBC |  | 12 | 32.4999 | 39.95 | -18.65% | S | 1.2292
Gradko | 50% TEA in Acetone | 2010 | UB | Sandwell MBC |  | 11 | 19.4658 | 23.1448 | -15.9% | S | 1.189
Gradko | 20% TEA in Water | 2010 | UB | Sandwell MBC |  | 11 | 30.9801 | 27.8221 | 11.35% | S | 0.89807
Gradko | 20% TEA in Water | 2010 | R | Sandwell MBC |  | 11 | 44.5934 | 45.017 | -0.9409% | S | 1.0095
Gradko | 20% TEA in Water | 2010 | R | Sandwell MBC |  | 11 | 37.1259 | 36.3891 | 2.025% | S | 0.98015
Gradko | 20% TEA in Water | 2010 | UB | Sandwell MBC |  | 10 | 22.4326 | 20.7509 | 8.104% | S | 0.92504
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | R | Rochford |  | 12 | 62.0231 | 46.5808 | 33.15% | G | 0.75102
Staffordshire Scientific Services | 20% TEA in Water | 2010 | R | Stoke-on-Trent City Council |  | 12 | 50.5819 | 41.433 | 22.08% | G | 0.81913
Staffordshire Scientific Services | 20% TEA in Water | 2010 | R | Stoke-on-Trent City Council |  | 12 | 46.2944 | 34.7692 | 33.15% | G | 0.75105
Glasgow Scientific Services | 20% TEA in Water | 2010 | UB | Glasgow City Council |  | 9 | 34.9833 | 38.6444 | -9.474% | P | 1.1047
Glasgow Scientific Services | 20% TEA in Water | 2010 | R | Glasgow City Council |  | 10 | 49.15 | 47.58 | 3.3% | P | 0.96806
Glasgow Scientific Services | 20% TEA in Water | 2010 | UC | Glasgow City Council |  | 12 | 38.1722 | 42.9 | -11.02% | P | 1.1239
Glasgow Scientific Services | 20% TEA in Water | 2010 | K | Glasgow City Council |  | 11 | 67.0258 | 80.6545 | -16.9% | G | 1.2033
Gradko | 20% TEA in Water | 2010 | R | Shropshire Council |  | 11 | 45.1736 | 38.1273 | 18.48% | G | 0.84402
Gradko | 20% TEA in Water | 2010 | R | Bromsgrove DC |  | 12 | 54.3875 | 52.606 | 3.386% | G | 0.96724
Harwell Scientific Services | 50% TEA in Acetone | 2010 | K | WODC |  | 12 | 53 | 45 | 17.9% | G | 0.85
Glasgow Scientific Services | 20% TEA in Water | 2010 | R | East Ayrshire Council |  | 11 | 33.603 | 41.8182 | -19.64% | P | 1.24
Lambeth Scientific Services | 50% TEA in Acetone | 2010 | R | London Borough of Haringey |  | 12 | 43.1667 | 44.4333 | -2.851% | S | 1.03
Tayside SS | 20% TEA in Water | 2010 | K | Dundee CC |  | 12 | 63.6417 | 54.1667 | 17.49% | G | 0.85112
Tayside SS | 20% TEA in Water | 2010 | R | Dundee CC |  | 10 | 51.1133 | 39.8 | 28.43% | G | 0.77866
Tayside SS | 20% TEA in Water | 2010 | K | Dundee CC |  | 11 | 45.1152 | 35.6727 | 26.47% | G | 0.7907
Staffordshire Scientific Services | 20% TEA in Water | 2010 | UB | Stockport MBC |  | 12 | 37.9083 | 32.1667 | 17.85% | G | 0.84854
Aberdeen CC | 20% TEA in Water | 2010 | R | Aberdeen CC |  | 10 | 31.3667 | 26.9092 | 16.56% | G | 0.85789
Harwell Scientific Services | 50% TEA in Acetone | 2010 | K | Suffolk Coastal DC |  | 9 | 50.563 | 46.1111 | 9.655% | G | 0.91195
Aberdeen CC | 20% TEA in Water | 2010 | R | Aberdeen CC |  | 12 | 35.5556 | 29.2282 | 21.65% | G | 0.82204
Aberdeen CC | 20% TEA in Water | 2010 | R | Aberdeen CC |  | 9 | 54.5185 | 56.2111 | -3.011% | P | 1.031
Aberdeen CC | 20% TEA in Water | 2010 | R | Aberdeen CC |  | 12 | 57.0139 | 52.065 | 9.505% | G | 0.9132
Gradko | 20% TEA in Water | 2010 | R | The Highland Council |  | 11 | 25.6403 | 24.2818 | 5.595% | G | 0.94702
Edinburgh Scientific Services | 50% TEA in acetone | 2010 | R | City of Edinburgh Council |  | 12 | 38.4125 | 30.2417 | 27.02% | G | 0.78729
Edinburgh Scientific Services | 50% TEA in acetone | 2010 | R | City of Edinburgh Council |  | 12 | 43.7556 | 36.925 | 18.5% | G | 0.84389
Edinburgh Scientific Services | 50% TEA in acetone | 2010 | R | City of Edinburgh Council |  | 12 | 35.05 | 29.8583 | 17.39% | G | 0.85188
Edinburgh Scientific Services | 50% TEA in acetone | 2010 | R | City of Edinburgh Council |  | 12 | 77.2889 | 71.2667 | 8.45% | G | 0.92208
Gradko | 20% TEA in water | 2011 | R | Scarborough Borough Council |  | 12 | 35.1892 | 36.9167 | -4.679% | G | 1.0491
Staffordshire Scientific Services | 20% TEA in water | 2011 | UB | Stoke-on-Trent City Council |  | 12 | 32.4028 | 32.2383 | 0.5101% | G | 0.99492
Bristol Scientific Services | 20% TEA in water | 2011 | R | Wiltshire Council |  | 10 | 42.8333 | 35.2 | 21.69% | G | 0.82179
Gradko | 20% TEA in Water | 2011 | R | Dudley MBC |  | 12 | 34.7834 | 28.2177 | 23.27% | G | 0.81124
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | R | Calderdale MBC |  | 12 | 62.5833 | 55.4083 | 12.95% | G | 0.88535
Gradko | 20% TEA in Water | 2011 | UB | Dudley MBC |  | 12 | 27.5671 | 25.0507 | 10.05% | G | 0.90872
ESG Didcot | 50% TEA in acetone | 2011 | B | Stockton on Tees |  | 12 | 27.1056 | 25.425 | 6.61% | G | 0.938
ESG Didcot | 50% TEA in acetone | 2011 | R | Stockton on Tees |  | 12 | 20.2111 | 18.4167 | 9.744% | G | 0.91121
Gradko | 20% TEA in Water | 2011 | R | Dudley MBC |  | 11 | 44.6933 | 39.981 | 11.79% | G | 0.89456
ESG Didcot | 50% TEA in Acetone | 2011 | R | Vale of White Horse DC |  | 12 | 34.8903 | 31.5 | 10.76% | G | 0.90283
ESG Didcot | 50% TEA in acetone | 2011 | R | Falkirk Council |  | 10 | 34.58 | 26.7882 | 29.09% | G | 0.77467
ESG Didcot | 50% TEA in acetone | 2011 | UB | Falkirk Council |  | 12 | 25.6417 | 21.7936 | 17.66% | G | 0.84993
ESG Didcot | 50% TEA in acetone | 2011 | R | Hambleton District Council |  | 12 | 25.6583 | 20.1333 | 27.44% | G | 0.78467
ESG Didcot | 50% TEA in acetone | 2011 | R | Swale Borough Council |  | 12 | 42.6403 | 28.5 | 49.62% | G | 0.66838
ESG Didcot | 50% TEA in acetone | 2011 | R | Swale Borough Council |  | 12 | 42.6403 | 28.5 | 49.62% | G | 0.66838
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | UC | Leeds City Council |  | 10 | 60.8 | 54.46 | 11.64% | S | 0.89572
Bristol Scientific Services | 20% TEA in water | 2011 | R | Wiltshire Council |  | 11 | 44.7424 | 35.4545 | 26.2% | G | 0.79241
Staffordshire Scientific Services | 20% TEA in water | 2011 | R | East Staffordshire Borough Council |  | 12 | 52.4722 | 39.0189 | 34.48% | G | 0.74361
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | R | Barnsley MBC |  | 12 | 45.9291 | 41.6667 | 10.23% | G | 0.9072
Gradko | 20% TEA in water | 2011 | K | South Lakeland District Council |  | 10 | 40.971 | 37.836 | 8.286% | G | 0.92348
Gradko | 50% TEA in acetone | 2011 | R | London Borough of Richmond upon Thames |  | 12 | 38.0894 | 39.7 | -4.057% | G | 1.0423
Gradko | 50% TEA in acetone | 2011 | B | London Borough of Richmond upon Thames |  | 12 | 27.8417 | 26.2583 | 6.03% | P | 0.94313
Gradko | 50% TEA in acetone | 2011 | UB | Reading Borough Council |  | 12 | 22.2439 | 23.5142 | -5.402% | G | 1.0571
Gradko | 50% TEA in Acetone | 2011 | R | Worthing Borough Council |  | 11 | 43.3939 | 38.3727 | 13.09% | G | 0.88429
Gradko | 20% TEA in water | 2011 | R | Gedling Borough Council |  | 11 | 43.0667 | 34.5782 | 24.55% | G | 0.8029
Gradko | 50% TEA in acetone | 2011 | R | Reading Borough Council |  | 10 | 38.2647 | 42.881 | -10.77% | G | 1.1206
Cardiff Scientific Services | 50% TEA in acetone | 2011 | R | Vale of Glamorgan Council |  | 10 | 28.5 | 24.7 | 15.38% | P | 0.86667
ESG Didcot | 50% TEA in acetone | 2011 | R | Dover District Council |  | 12 | 42.0806 | 36.9167 | 13.99% | G | 0.87729
Bristol Scientific Services | 20% TEA in water | 2011 | R | Wiltshire Council |  | 11 | 44.0303 | 35.8182 | 22.93% | G | 0.81349
Cardiff Scientific Services | 50% TEA in acetone | 2011 | O | Vale of Glamorgan Council |  | 10 | 16.7 | 13.3 | 25.56% | G | 0.79641
Gradko | 20% TEA in water | 2011 | R | Gateshead |  | 12 | 39.1917 | 37.3583 | 4.907% | G | 0.95322
Gradko | 20% TEA in water | 2011 | R | Gateshead |  | 12 | 37.1056 | 36.4333 | 1.845% | P | 0.98188
Gradko | 20% TEA in water | 2011 | R | Gateshead |  | 10 | 32.85 | 31.27 | 5.053% | G | 0.9519
Gradko | 20% TEA in water | 2011 | R | Gosport Borough Council |  | 10 | 27.615 | 24.8456 | 11.15% | S | 0.89971
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | R | Leeds City Council |  | 10 | 52.95 | 44.85 | 18.06% | G | 0.84703
Lambeth Scientific Services | 50% TEA in acetone | 2011 | R | LB Islington |  | 11 | 58.1667 | 57.4455 | 1.255% | G | 0.9876
ESG Didcot | 50% TEA in acetone | 2011 | UB | Medway Council |  | 12 | 21.9389 | 25 | -12.24% | G | 1.1395
Staffordshire Scientific Services | 20% TEA in water | 2011 | R | Stoke-on-Trent City Council |  | 12 | 40.3194 | 38.6731 | 4.257% | G | 0.95917
Bristol Scientific Services | 20% TEA in water | 2011 | B | Pembrokeshire Council |  | 12 | 6.5222 | 4.6417 | 40.51% | G | 0.71167
Cardiff Scientific Services | 50% TEA in acetone | 2011 | R | Vale of Glamorgan Council |  | 10 | 22.2667 | 20.8 | 7.051% | G | 0.93413
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | UC | Leeds City Council |  | 11 | 48.3485 | 37.8182 | 27.84% | G | 0.7822
ESG Didcot | 50% TEA in acetone | 2011 | R | North East Lincolnshire Council |  | 10 | 52.09 | 47.818 | 8.934% | G | 0.91799
ESG Didcot | 50% TEA in acetone | 2011 | R | North East Lincolnshire Council |  | 9 | 37.6815 | 35.0589 | 7.481% | G | 0.9304
ESG Didcot | 50% TEA in acetone | 2011 | R | North East Lincolnshire Council |  | 12 | 40.9167 | 30.8183 | 32.77% | G | 0.7532
ESG Didcot | 50% TEA in acetone | 2011 | UB | North East Lincolnshire Council |  | 12 | 22.0694 | 20.5375 | 7.459% | G | 0.93059
ESG Didcot | 50% TEA in acetone | 2011 | B | Medway Council |  | 9 | 31.5852 | 20.6667 | 52.83% | P | 0.65432
Staffordshire Scientific Services | 20% TEA in water | 2011 | O | South Staffordshire Council |  | 11 | 36.3364 | 33.9536 | 7.018% | G | 0.93443
Gradko | 20% TEA in water | 2011 | UC | Southampton City Council |  | 12 | 30.9356 | 34.675 | -10.78% | G | 1.1209
Gradko | 20% TEA in Water | 2011 | R | Dudley MBC |  | 9 | 49.7397 | 50.5096 | -1.524% | G | 1.0155
ESG Didcot | 50% TEA in acetone | 2011 | R | Wrexham County Borough Council |  | 12 | 21.5194 | 19.2567 | 11.75% | G | 0.89485
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | R | Leeds City Council |  | 9 | 37.8333 | 30.1778 | 25.37% | G | 0.79765
ESG Didcot | 50% TEA in acetone | 2011 | R | Medway Council |  | 9 | 36.0852 | 30.4444 | 18.53% | G | 0.84368
Gradko | 50% TEA in acetone | 2011 | R | West Berkshire Council |  | 11 | 41.1064 | 44.0455 | -6.673% | G | 1.0715
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | KS | Wakefield MDC |  | 12 | 61.6667 | 57.9367 | 6.438% | G | 0.93951
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | UB | Wakefield MDC |  | 12 | 40.0833 | 29.9233 | 33.95% | G | 0.74653
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | UB | Wakefield MDC |  | 12 | 42.0556 | 41.239 | 1.98% | G | 0.98058
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | UB | Wakefield MDC |  | 11 | 30.7576 | 29.8655 | 2.987% | G | 0.971
Lambeth Scientific Services | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 11 | 89.3818 | 98.8048 | -9.537% | G | 1.1054
ESG Glasgow | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 11 | 118.6091 | 99.2866 | 19.46% | G | 0.83709
ESG Glasgow | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 11 | 120.5939 | 99.2866 | 21.46% | G | 0.82331
ESG Didcot | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 11 | 125.4 | 99.2866 | 26.3% | G | 0.79176
ESG Didcot | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 11 | 128.2364 | 99.2866 | 29.16% | G | 0.77425
Milton Keynes Council | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 121.525 | 99.8794 | 21.67% | G | 0.82188
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 11 | 129.0424 | 99.0966 | 30.22% | G | 0.76794
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 12 | 103.0798 | 99.8794 | 3.204% | G | 0.96895
Bristol Scientific Services | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 118.8556 | 99.8794 | 19% | G | 0.84034
Northampton BC | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 139.1 | 99.8794 | 39.27% | G | 0.71804
Glasgow Scientific services | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 11 | 115.2 | 98.8048 | 16.59% | G | 0.85768
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 12 | 137.0833 | 99.8794 | 37.25% | G | 0.7286
Gradko | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 111.2733 | 99.8794 | 11.41% | P | 0.8976
Gradko | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 12 | 97.8767 | 99.8794 | -2.005% | G | 1.0205
Aberdeen Scientific Services | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 11 | 117.87 | 100.6694 | 17.09% | G | 0.85407
Staffordshire Scientific Services | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 118.6417 | 99.8794 | 18.78% | G | 0.84186
Kirklees Council | 50% TEA in acetone | 2011 | K | Marylebone Road Intercomparison |  | 12 | 135.7818 | 99.8794 | 35.95% | G | 0.73559
Tayside Scientific Services | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 127.4722 | 99.8794 | 27.63% | G | 0.78354
Exova | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 127.6361 | 99.8794 | 27.79% | P | 0.78253
Kent Scientific Services | 20% TEA in water | 2011 | K | Marylebone Road Intercomparison |  | 12 | 129.5822 | 99.8794 | 29.74% | G | 0.77078
Staffordshire Scientific Services | 20% TEA in water | 2011 | R | Cannock Chase Council |  | 9 | 43.5407 | 43.4667 | 0.1704% | G | 0.9983
ESG Didcot | 50% TEA in acetone | 2011 | R | Castlereagh Borough Council |  | 11 | 48.1212 | 39.8182 | 20.85% | G | 0.82746
ESG Didcot | 50% TEA in acetone | 2011 | R | Down District Council |  | 12 | 50.5139 | 36.3333 | 39.03% | G | 0.71927
ESG Didcot | 50% TEA in acetone | 2011 | R | Lisburn City Council |  | 12 | 30.1667 | 20.1667 | 49.59% | G | 0.66851
ESG Didcot | 50% TEA in acetone | 2011 | R | North Down Borough Council |  | 11 | 45 | 27 | 66.67% | G | 0.6
Gradko | 50% TEA in acetone | 2011 | R | Lewisham Council |  | 12 | 80.1936 | 48.1921 | 66.4% | P | 0.60095
Gradko | 50% TEA in acetone | 2011 | R | Wolverhampton City |  | 12 | 38.09 | 35.8125 | 6.36% | G | 0.94021
Gradko | 50% TEA in acetone | 2011 | R | Wolverhampton City |  | 12 | 39.1964 | 33.2658 | 17.83% | G | 0.8487
Gradko | 50% TEA in acetone | 2011 | R | Wolverhampton City |  | 9 | 34.1448 | 38.8367 | -12.08% | G | 1.1374
ESG Didcot | 50% TEA in acetone | 2011 | R | Maidstone Borough Council |  | 12 | 48.2736 | 53.25 | -9.345% | P | 1.1031
ESG Didcot | 50% TEA in acetone | 2011 | B | Maidstone Borough Council |  | 12 | 17.2583 | 12.5833 | 37.15% | G | 0.72912
Gradko | 20% TEA in water | 2011 | UB | Luton Borough Council |  | 11 | 38.7921 | 34.9091 | 11.12% | G | 0.8999
ESG Didcot | 50% TEA in Acetone | 2011 | K | Suffolk Coastal District Council |  | 12 | 50.5028 | 42.5833 | 18.6% | G | 0.84319
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | UB | Chesterfield BC |  | 10 | 20.6 | 15.999 | 28.76% | G | 0.77665
ESG Didcot | 50% TEA in acetone | 2011 | UB | CITY OF YORK COUNCIL |  | 12 | 26.8673 | 24.865 | 8.053% | G | 0.92547
ESG Didcot | 50% TEA in acetone | 2011 | R | CITY OF YORK COUNCIL |  | 11 | 38.0958 | 30.4245 | 25.21% | G | 0.79863
ESG Didcot | 50% TEA in acetone | 2011 | R | CITY OF YORK COUNCIL |  | 11 | 35.3176 | 28.5845 | 23.56% | G | 0.80936
ESG Didcot | 50% TEA in acetone | 2011 | R | CITY OF YORK COUNCIL |  | 12 | 44.7444 | 50.0058 | -10.52% | G | 1.1176
Glasgow Scientific Services | 20% TEA in water | 2011 | R | West Dunbartonshire Council |  | 12 | 23.7278 | 18.1833 | 30.49% | G | 0.76633
ESG Didcot | 50% TEA in acetone | 2011 | R | Dumfries and Galloway Council |  | 12 | 37.9306 | 31.5961 | 20.05% | G | 0.833
Gradko | 50% TEA in acetone | 2011 | R | East Hampshire District Council |  | 11 | 24.3824 | 22.3545 | 9.071% | G | 0.91683
Gradko | 20% TEA in water | 2011 | R | Exeter City Council |  | 11 | 37.455 | 32.5364 | 15.12% | G | 0.86868
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | UB | Leeds City Council |  | 9 | 25.8889 | 20.6111 | 25.61% | S | 0.79614
Northampton BC | 20% TEA in water | 2011 | R | Northampton BC |  | 10 | 55.4433 | 36.54 | 51.73% | G | 0.65905
Northampton BC | 20% TEA in water | 2011 | UB | Northampton BC |  | 12 | 24.6861 | 18.4667 | 33.68% | G | 0.74806
Gradko | 50% TEA in Acetone | 2011 | R | Stevenage Borough Council |  | 13 | 35.3333 | 30.3846 | 16.29% | G | 0.85994
Glasgow Scientific Services | 20% TEA in water | 2011 | R | West Dunbartonshire Council |  | 11 | 25.3773 | 20.7273 | 22.43% | P | 0.81677
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | R | West Lothian Council |  | 11 | 40.9333 | 42.7 | -4.137% | G | 1.0432
Staffordshire Scientific Services | 20% TEA in water | 2011 | UB | WIGAN |  | 11 | 30.4697 | 23.4545 | 29.91% | G | 0.76977
Staffordshire Scientific Services | 20% TEA in water | 2011 | UB | WIGAN |  | 11 | 29.697 | 25.4545 | 16.67% | G | 0.85714
Gradko | 20% TEA in water | 2011 | UB | Belfast City Council |  | 12 | 35.6126 | 28.8333 | 23.51% | G | 0.80964
Gradko | 20% TEA in water | 2011 | R | Bromsgrove District Council (Worcester Regulatory Services) |  | 10 | 56.4517 | 52.9 | 6.714% | G | 0.93708
ESG Didcot | 50% TEA in acetone | 2011 | R | Cambridge City Council |  | 11 | 47.1121 | 40.1818 | 17.25% | G | 0.8529
Gradko | 20% TEA in water | 2011 | R | Monmouthshire County Council |  | 11 | 46.7047 | 39.6 | 17.94% | G | 0.84788
Bristol Scientific Services | 20% TEA in water | 2011 | UB | LB Waltham Forest |  | 12 | 38.4333 | 37.5 | 2.489% | S | 0.97572
Kirklees Council | 50% TEA in acetone | 2011 | R | Kirklees Council |  | 12 | 53.05 | 41.6667 | 27.32% | S | 0.78542
ESG Didcot | 50% TEA in acetone | 2011 | UB | Gravesham Borough Council |  | 12 | 30.4389 | 25.3333 | 20.15% | G | 0.83227
ESG Didcot | 50% TEA in acetone | 2011 | R | Gravesham Borough Council |  | 12 | 43.125 | 33.8333 | 27.46% | G | 0.78454
Kirklees Council | 50% TEA in acetone | 2011 | UC | Kirklees Council |  | 11 | 48.2473 | 36.4364 | 32.42% | S | 0.7552
Kirklees Council | 50% TEA in acetone | 2011 | R | Kirklees Council |  | 9 | 65.0637 | 43.0922 | 50.99% | G | 0.66231
Kirklees Council | 50% TEA in acetone | 2011 | UB | Kirklees Council |  | 13 | 30.5015 | 24.9454 | 22.27% | P | 0.81784
Gradko | 20% TEA in water | 2011 | K | New Forest District Council |  | 10 | 48.961 | 41.94 | 16.74% | G | 0.8566
Gradko | 20% TEA in water | 2011 | R | New Forest District Council |  | 12 | 34.3747 | 26.4583 | 29.92% | G | 0.7697
Gradko | 50% TEA in Acetone | 2011 | UB | London Borough of Bexley |  | 11 | 25.7826 | 23.0791 | 11.71% | G | 0.89514
Gradko | 50% TEA in Acetone | 2011 | SU | London Borough of Bexley |  | 10 | 27.675 | 28.689 | -3.534% | G | 1.0366
Gradko | 50% TEA in Acetone | 2011 | SU | London Borough of Bexley |  | 9 | 27.327 | 26.3678 | 3.638% | G | 0.9649
Gradko | 50% TEA in Acetone | 2011 | R | London Borough of Bexley |  | 11 | 53.3462 | 43.8627 | 21.62% | G | 0.82223
Bristol Scientific Services | 20% TEA in water | 2011 | R | Bath & North East Somerset |  | 13 | 63.3228 | 56.9549 | 11.18% | G | 0.89944
Tayside Scientific Services | 20% TEA in water | 2011 | UB | Dundee City Council |  | 9 | 15.6037 | 10.1111 | 54.32% | G | 0.64799
Tayside Scientific Services | 20% TEA in water | 2011 | R | Dundee City Council |  | 12 | 46.3556 | 36.1667 | 28.17% | G | 0.7802
Tayside Scientific Services | 20% TEA in water | 2011 | R | Dundee City Council |  | 9 | 46.3111 | 35.6667 | 29.84% | G | 0.77015
Lambeth Scientific Services | 50% TEA in acetone | 2011 | R | Epsom & Ewell Borough Council |  | 10 | 39.6667 | 44.43 | -10.72% | P | 1.1201
Tayside Scientific Services | 20% TEA in water | 2011 | K | Fife Council |  | 9 | 40.987 | 30.1111 | 36.12% | G | 0.73465
Tayside Scientific Services | 20% TEA in water | 2011 | R | Fife Council |  | 9 | 38.7074 | 31.4444 | 23.1% | G | 0.81236
Tayside Scientific Services | 20% TEA in water | 2011 | R | Fife Council |  | 9 | 23.663 | 19.2222 | 23.1% | G | 0.81233
Tayside Scientific Services | 20% TEA in water | 2011 | R | Fife Council |  | 9 | 31.8148 | 29.2222 | 8.872% | G | 0.91851
Glasgow Scientific Services | 20% TEA in water | 2011 | UB | Glasgow City Council |  | 9 | 33.5481 | 37.2222 | -9.871% | P | 1.1095
Glasgow Scientific Services | 20% TEA in water | 2011 | UC | Glasgow City Council |  | 12 | 36.3333 | 34.475 | 5.39% | G | 0.94885
Glasgow Scientific Services | 20% TEA in water | 2011 | K | Glasgow City Council |  | 12 | 71.6667 | 72.475 | -1.115% | P | 1.0113
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | R | Leeds City Council 8 |  | 12 | 53.25 | 41.6667 | 27.8% | S | 0.78247
Staffordshire Scientific Services | 20% TEA in water | 2011 | SU | Manchester City Council |  | 12 | 30.3861 | 22.9482 | 32.41% | G | 0.75522
Staffordshire Scientific Services | 20% TEA in water | 2011 | K | Manchester City Council |  | 12 | 77.3444 | 66.2111 | 16.81% | G | 0.85605
Staffordshire Scientific Services | 20% TEA in water | 2011 | UB | Manchester City Council |  | 12 | 47.575 | 43.9336 | 8.288% | G | 0.92346
Gradko | 20% TEA in water | 2011 | R | Fareham Borough Council |  | 12 | 38.8408 | 33.0717 | 17.44% | G | 0.85147
ESG Didcot | 50% TEA in acetone | 2011 | R | Rugby Borough Council |  | 10 | 34.2833 | 34.3962 | -0.3281% | G | 1.0033
Gradko | 20% TEA in water | 2011 | R | Rushcliffe BC |  | 11 | 34.9606 | 38.6275 | -9.493% | G | 1.1049
ESG Didcot | 50% TEA in acetone | 2011 | R | Wycombe District Council |  | 10 | 43.3333 | 38.85 | 11.54% | G | 0.89654
Bristol Scientific Services | 20% TEA in water | 2011 | R | South Gloucestershire Council |  | 12 | 31.1087 | 26.6833 | 16.58% | G | 0.85775
Gradko | 20% TEA in Water | 2011 | R | Carlisle City Council |  | 12 | 34.6681 | 27.7827 | 24.78% | G | 0.80139
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | O | Blaby District Council |  | 12 | 36.0139 | 25.7 | 40.13% | G | 0.71361
Lambeth Scientific Services | 50% TEA in acetone | 2011 | R | Waverley BC |  | 12 | 32.7222 | 35.381 | -7.515% | P | 1.0813
Lambeth Scientific Services | 50% TEA in acetone | 2011 | R | Waverley BC |  | 9 | 25.3333 | 24.1552 | 4.878% | G | 0.95349
Lambeth Scientific Services | 50% TEA in acetone | 2011 | R | Waverley BC |  | 12 | 34.1667 | 38.8756 | -12.11% | G | 1.1378
ESG Didcot | 50% TEA in acetone | 2011 | R | Horsham Distric Council |  | 10 | 32.8033 | 26.474 | 23.91% | G | 0.80705
Gradko | 20% TEA in Water | 2011 | O | North Warwickshire Borough Council |  | 12 | 47.9139 | 38.945 | 23.03% | G | 0.81281
Gradko | 50% TEA in acetone | 2011 | UB | Norwich City Council |  | 10 | 14.389 | 13.182 | 9.156% | G | 0.91612
ESG Didcot | 50% TEA in acetone | 2011 | R | Horsham Distric Council |  | 12 | 31.4306 | 25.3117 | 24.17% | G | 0.80532
Gradko | 50% TEA in acetone | 2011 | UB | Sandwell Metropolitan Borough Council |  | 12 | 24.0539 | 23.34 | 3.059% | G | 0.97032
Gradko | 50% TEA in acetone | 2011 | R | Sandwell Metropolitan Borough Council |  | 10 | 41.7821 | 41.3774 | 0.9781% | G | 0.99031
ESG Didcot | 50% TEA in acetone | 2011 | R | Horsham Distric Council |  | 12 | 29.9028 | 23.6825 | 26.27% | G | 0.79198
ESG Didcot | 50% TEA in acetone | 2011 | UB | Canterbury City Council |  | 11 | 17.4485 | 14.8182 | 17.75% | G | 0.84925
ESG Didcot | 50% TEA in acetone | 2011 | R | Canterbury City Council |  | 12 | 39.1667 | 33.9167 | 15.48% | G | 0.86596
Glasgow Scientific Services | 20% TEA in water | 2011 | R | East Ayrshire Council |  | 10 | 31.4167 | 35.6 | -11.75% | P | 1.1332
Gradko | 20% TEA in water | 2011 | R | Wokingham Borough Council |  | 11 | 41.265 | 38 | 8.592% | G | 0.92088
Staffordshire Scientific Services | 20% TEA in water | 2011 | R | Stockport MBC |  | 11 | 30.9561 | 29.0909 | 6.411% | G | 0.93975
ESG Didcot | 50% TEA in acetone | 2011 | R | Tunbridge Wells Borough Council |  | 12 | 58.8556 | 42.5 | 38.48% | P | 0.72211
Gradko | 50% TEA in acetone | 2011 | R | LB Newham |  | 12 | 39.7535 | 47.0333 | -15.48% | G | 1.1831
Bristol Scientific Services | 20% TEA in water | 2011 | R | Brighton and Hove City Council |  | 11 | 48.5515 | 34.3909 | 41.18% | G | 0.70834
Aberdeen Scientific Services | 20% TEA in water | 2011 | R | ABERDEEN CITY COUNCIL |  | 12 | 25.9722 | 23.5475 | 10.3% | G | 0.90664
Aberdeen Scientific Services | 20% TEA in water | 2011 | UB | ABERDEEN CITY COUNCIL |  | 13 | 34.0769 | 22.9392 | 48.55% | G | 0.67316
Aberdeen Scientific Services | 20% TEA in water | 2011 | R | ABERDEEN CITY COUNCIL |  | 12 | 36.6667 | 31.9283 | 14.84% | G | 0.87077
Aberdeen Scientific Services | 20% TEA in water | 2011 | R | ABERDEEN CITY COUNCIL |  | 13 | 51.359 | 40.4392 | 27% | G | 0.78738
Aberdeen Scientific Services | 20% TEA in water | 2011 | R | ABERDEEN CITY COUNCIL |  | 13 | 51.8205 | 50.5246 | 2.565% | G | 0.97499
Gradko | 20% TEA in water | 2011 | R | Bedford Borough Council |  | 11 | 46.8182 | 49.4545 | -5.331% | G | 1.0563
Gradko | 50% TEA in acetone | 2011 | R | London Borough of Hounslow |  | 12 | 52.5158 | 50.6232 | 3.739% | G | 0.96396
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | R | City of Edinburgh Council |  | 11 | 43.8515 | 38.1364 | 14.99% | G | 0.86967
Gradko | 20% TEA in water | 2011 | R | Liverpool City Council |  | 9 | 43.8889 | 35 | 25.4% | G | 0.79747
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | R | Leeds City Council 7 |  | 11 | 46.8182 | 36.3364 | 28.85% | S | 0.77612
Gradko | 20% TEA in water | 2011 | R | Newtownabbey BC |  | 11 | 32.9012 | 30.1182 | 9.24% | G | 0.91541
Gradko | 20% TEA in water | 2011 | KS | Liverpool City Council |  | 12 | 54.1944 | 46 | 17.81% | G | 0.8488
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | R | Oxford City Council |  | 10 | 54.0727 | 63.48 | -14.82% | G | 1.174
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | UB | Oxford City Council |  | 10 | 19.39 | 18.03 | 7.543% | G | 0.92986
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | R | City of Edinburgh Council |  | 11 | 61.5515 | 40.3636 | 52.49% | G | 0.65577
Gradko | 20% TEA in water | 2011 | R | Liverpool City Council |  | 12 | 41.75 | 34 | 22.79% | G | 0.81437
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | R | City of Edinburgh Council |  | 12 | 41.7389 | 28.9 | 44.43% | G | 0.6924
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | R | City of Edinburgh Council |  | 12 | 37.6889 | 29.0917 | 29.55% | G | 0.77189
Gradko | 20% TEA in water | 2011 | UB | Liverpool City Council |  | 9 | 23.2037 | 23.7778 | -2.414% | S | 1.0247
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | KS | City of Edinburgh Council |  | 11 | 82.5242 | 66.5818 | 23.94% | G | 0.80682
Gradko | 20% TEA in water | 2011 | R | NOTTINGHAM CITY COUNCIL |  | 10 | 43.9223 | 41.71 | 5.304% | G | 0.94963
Gradko | 20% TEA in water | 2011 | R | NOTTINGHAM CITY COUNCIL |  | 11 | 44.7433 | 38.3818 | 16.57% | G | 0.85782
Gradko | 20% TEA in water | 2011 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 45.0461 | 41.0333 | 9.779% | G | 0.91092
Gradko | 20% TEA in water | 2011 | R | Bedford Borough Council |  | 9 | 47.6667 | 48.3778 | -1.47% | G | 1.0149
Gradko | 20% TEA in water | 2011 | R | Blackburn with Darwen Borough Council |  | 11 | 29.2812 | 25.0455 | 16.91% | G | 0.85534
Gradko | 20% TEA in Water | 2011 | R | Central Bedfordshire Council |  | 9 | 37.1741 | 35 | 6.212% | G | 0.94152
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | R | Stirling Council |  | 11 | 27.3333 | 27.9773 | -2.302% | G | 1.0236
ESG Didcot | 50% TEA in acetone | 2011 | UB | Kingston upon Hull City Council |  | 12 | 34.5833 | 28.6633 | 20.65% | P | 0.82882
ESG Didcot | 50% TEA in acetone | 2011 | UB | Kingston upon Hull City Council |  | 12 | 30.0944 | 32.15 | -6.394% | G | 1.0683
ESG Didcot | 50% TEA in acetone | 2011 | KS | Kingston upon Hull City Council |  | 12 | 36.7194 | 30.4833 | 20.46% | G | 0.83017
Gradko | 50% TEA in acetone | 2011 | R | London Borough of Croydon |  | 9 | 46.4233 | 41.4778 | 11.92% | P | 0.89347
Gradko | 50% TEA in acetone | 2011 | KS | London Borough of Croydon |  | 12 | 60.6254 | 60.5833 | 0.0695% | G | 0.99931
Gradko | 20% TEA in water | 2011 | UB | North Tyneside |  | 12 | 21.2167 | 21.8342 | -2.828% | G | 1.0291
Bristol Scientific Services | 20% TEA in water | 2011 | R | Wiltshire Council |  | 11 | 44.4091 | 35.0909 | 26.55% | G | 0.79017
Staffordshire Scientific Services | 20% TEA in water | 2011 | R | Bury Council |  | 10 | 73.4 | 69.3 | 5.916% | G | 0.94414
Gradko | 20% TEA in water | 2011 | Rural | Chelmsford Borough Council |  | 10 | 16.9757 | 14.4904 | 17.15% | G | 0.8536
Gradko | 20% TEA in water | 2011 | R | Chelmsford Borough Council |  | 10 | 28.291 | 20.4315 | 38.47% | G | 0.72219
Staffordshire Scientific Services | 20% TEA in water | 2011 | UI | Salford City Council |  | 10 | 36.1467 | 32.7 | 10.54% | G | 0.90465
Staffordshire Scientific Services | 20% TEA in water | 2011 | R | Salford City Council |  | 10 | 65.8567 | 66.3 | -0.6687% | G | 1.0067
ESG Glasgow | 50% TEA in acetone | 2011 | UB | Slough BC |  | 9 | 47.8815 | 42.5556 | 12.52% | G | 0.88877
ESG Glasgow | 50% TEA in acetone | 2011 | UB | Slough BC |  | 13 | 34.5051 | 31.0769 | 11.03% | G | 0.90065
ESG Glasgow | 50% TEA in acetone | 2011 | R | Slough BC |  | 13 | 41.3718 | 35.7692 | 15.66% | P | 0.86458
Gradko | 20% TEA in Water | 2011 | UB | East Herts Council |  | 12 | 21.3889 | 16.4667 | 29.89% | G | 0.76987
Gradko | 50% TEA in acetone | 2011 | R | London Borough of Hounslow |  | 10 | 58.9936 | 56.22 | 4.934% | P | 0.95298
Milton Keynes Council | 20% TEA in water | 2011 | UC | Milton Keynes Council |  | 12 | 23.3878 | 20.9033 | 11.89% | G | 0.89377
Milton Keynes Council | 20% TEA in water | 2011 | R | Milton Keynes Council |  | 12 | 41.068 | 34.6256 | 18.61% | G | 0.84313
Milton Keynes Council | 20% TEA in water | 2011 | R | Milton Keynes Council |  | 12 | 33.066 | 26.8351 | 23.22% | G | 0.81156
ESG Didcot | 50% TEA in acetone | 2011 | R | Northumberland County Council |  | 9 | 37.6333 | 27.1111 | 38.81% | S | 0.7204
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | R | Wakefield MDC |  | 9 | 39.5 | 33.0642 | 19.46% | G | 0.83707
Gradko | 20% TEA in water | 2012 | R | Wiltshire Council |  | 11 | 38.1515 | 39.7273 | -3.966% | G | 1.0413
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | KS | Wakefield MDC |  | 11 | 59.0152 | 57.6841 | 2.308% | G | 0.97745
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | UB | Wakefield MDC |  | 11 | 41.9091 | 34.2601 | 22.33% | G | 0.81749
Gradko | 20% TEA in Water | 2012 | R | Dudley MBC |  | 11 | 38.5702 | 33.2391 | 16.04% | G | 0.86178
Gradko | 20% TEA in Water | 2012 | UB | Dudley MBC |  | 10 | 27.014 | 26.19 | 3.146% | G | 0.9695
South Yorkshire Air Quality Samplers | 50% TEA in Acetone | 2012 | R | Chesterfield BC |  | 12 | 27.7778 | 20.2325 | 37.29% | G | 0.72837
Gradko | 20% TEA in Water | 2012 | R | Dudley MBC |  | 11 | 43.1033 | 42.2678 | 1.977% | G | 0.98062
Gradko | 20% TEA in water | 2012 | R | Monmouthshire County Council |  | 9 | 47.1596 | 41.8532 | 12.68% | G | 0.88748
ESG Didcot | 50% TEA in Acetone | 2012 | R | Swale Borough Council |  | 9 | 44.2519 | 38.5556 | 14.77% | G | 0.87128
ESG Didcot | 50% TEA in Acetone | 2012 | R | Swale Borough Council |  | 11 | 45.5424 | 35.9091 | 26.83% | G | 0.78848
ESG Didcot | 50% TEA in acetone | 2012 | R | Vale Of White Horse District Council |  | 12 | 36.8833 | 29.625 | 24.5% | G | 0.80321
ESG Didcot | 20% TEA in water | 2012 | KS | SOUTH LAKELAND DISTRICT COUNCIL |  | 9 | 45.1593 | 29.1789 | 54.77% | G | 0.64613
Gradko | 20% TEA in water | 2012 | R | Cheshire West & Chester |  | 11 | 40.2843 | 44.7909 | -10.06% | G | 1.1119
Staffordshire Scientific Services | 20% TEA in water | 2012 | R | Cannock Chase Coujncil |  | 12 | 49.4472 | 38.6333 | 27.99% | G | 0.7813
Gradko | 20% TEA in Water | 2012 | UB | East Herts Council |  | 11 | 20.7273 | 12.0545 | 71.95% | G | 0.58158
Staffordshire Scientific Services | 20% TEA in Water | 2012 | R | East Staffordshire Borough Council |  | 9 | 48.8333 | 31.4778 | 55.14% | G | 0.6446
Gradko | 20% TEA in Water | 2012 | R | Gateshead Council |  | 11 | 33.8152 | 34.2909 | -1.387% | G | 1.0141
Gradko | 20% TEA in Water | 2012 | R | Gateshead Council |  | 11 | 35.5515 | 36.7091 | -3.153% | G | 1.0326
Gradko | 20% TEA in Water | 2012 | R | Gateshead Council |  | 11 | 32.3758 | 33.2273 | -2.563% | G | 1.0263
Gradko | 50% TEA in acetone | 2012 | UB | Reading Borough Council |  | 12 | 21.3194 | 25.9667 | -17.9% | G | 1.218
Gradko | 50% TEA in acetone | 2012 | R | Reading Borough Council |  | 12 | 39.1111 | 47.0083 | -16.8% | G | 1.2019
Gradko | 50% TEA in Acetone | 2012 | R | Stevenage Borough Council |  | 11 | 32.8358 | 28.1545 | 16.63% | G | 0.85744
Gradko | 20% TEA in Water | 2012 | R | Dudley MBC |  | 9 | 55.4588 | 59.955 | -7.499% | G | 1.0811
ESG Didcot | 50% TEA in acetone | 2012 | R | Wrexham County Borough Council |  | 12 | 23.5306 | 21.4104 | 9.903% | P | 0.9099
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | R | Calderdale MBC |  | 12 | 62.1389 | 46.2417 | 34.38% | G | 0.74417
Glasgow Scientific Services | 20% TEA in water | 2012 | UB | Glasgow City Council |  | 11 | 29.4879 | 33.6636 | -12.4% | P | 1.1416
Glasgow Scientific Services | 20% TEA in water | 2012 | R | Glasgow City Council |  | 10 | 44.6 | 38.72 | 15.19% | P | 0.86816
Glasgow Scientific Services | 20% TEA in water | 2012 | KS | Glasgow City Council |  | 11 | 69.9621 | 71.0455 | -1.525% | P | 1.0155
ESG Didcot | 50% TEA in acetone | 2012 | UB | Gravesham Borough Council |  | 12 | 32.1611 | 27.1667 | 18.38% | G | 0.84471
ESG Didcot | 50% TEA in acetone | 2012 | UB | Gravesham Borough Council |  | 12 | 44.3944 | 35.4167 | 25.35% | G | 0.79777
ESG Didcot | 50% TEA in acetone | 2012 | R | Hambleton District Council |  | 12 | 25.325 | 19.3 | 31.22% | G | 0.76209
Lambeth Scientific Services | 50% TEA in acetone | 2012 | R | London Borough of Islington |  | 12 | 63.7083 | 53.0667 | 20.05% | G | 0.83296
Gradko | 50% TEA in acetone | 2012 | R | BC King's Lynn & West Norfolk |  | 10 | 28.0037 | 24.73 | 13.24% | G | 0.8831
Gradko | 50% TEA in acetone | 2012 | R | London Borough of Richmond upon Thames |  | 12 | 36.0776 | 39.3 | -8.199% | G | 1.0893
Gradko | 50% TEA in acetone | 2012 | KS | London Borough of Richmond upon Thames |  | 9 | 45.6367 | 40.8222 | 11.79% | G | 0.8945
Gradko | 50% TEA in acetone | 2012 | B | London Borough of Richmond upon Thames |  | 12 | 23.3222 | 24.2333 | -3.76% | G | 1.0391
Gradko | 20% TEA in Water | 2012 | UB | Luton Borough Council |  | 11 | 38.3939 | 29.6636 | 29.43% | G | 0.77261
Gradko | 50% TEA in acetone | 2012 | UB | Norwich City Council |  | 12 | 13.8275 | 17.8392 | -22.49% | G | 1.2901
ESG Didcot | 50% TEA in acetone | 2012 | R | North East Lincolnshire Council |  | 12 | 38.4194 | 30.4467 | 26.19% | G | 0.79248
ESG Didcot | 50% TEA in acetone | 2012 | R | North East Lincolnshire Council |  | 12 | 42.4889 | 32.1217 | 32.27% | G | 0.756
ESG Glasgow | 50% TEA in acetone | 2012 | UB | Slough Borough Council |  | 12 | 47.3306 | 39.9167 | 18.57% | G | 0.84336
ESG Glasgow | 50% TEA in Acetone | 2012 | UB | Slough Borough Council |  | 12 | 34.0417 | 29.9167 | 13.79% | G | 0.87882
ESG Glasgow | 50% TEA in acetone | 2012 | R | Slough Borough Council |  | 12 | 42.9611 | 37.3333 | 15.07% | G | 0.869
gradko | 20% TEA in water | 2012 | UC | Southampton City Council |  | 11 | 30.3121 | 33.0636 | -8.322% | G | 1.0908
Staffordshire Scientific Services | 20% TEA in water | 2012 | UB | Stoke-on-Trent City Council |  | 12 | 33.9306 | 31.7694 | 6.803% | G | 0.93631
Staffordshire Scientific Services | 20% TEA in water | 2012 | R | Stoke-on-Trent City Council |  | 11 | 58.1848 | 55.3192 | 5.18% | G | 0.95075
ESG Didcot | 50% TEA in Acetone | 2012 | R | Falkirk Council |  | 12 | 38.4972 | 33.2168 | 15.9% | G | 0.86284
ESG Didcot | 50% TEA in Acetone | 2012 | UB | Falkirk Council |  | 12 | 26.7861 | 24.1769 | 10.79% | G | 0.90259
Glasgow Scientific Services | 20% TEA in water | 2012 | R | East Dunbartonshire Council |  | 11 | 30.2515 | 29.7273 | 1.764% | G | 0.98267
Glasgow Scientific Services | 20% TEA in water | 2012 | R | East Dunbartonshire Council |  | 12 | 40.8917 | 41.5 | -1.466% | P | 1.0149
Glasgow Scientific Services | 20% TEA in water | 2012 | R | East Dunbartonshire Council |  | 12 | 40.2 | 33.8333 | 18.82% | P | 0.84163
Glasgow Scientific Services | 20% TEA in water | 2012 | R | East Dunbartonshire Council |  | 12 | 27.2917 | 25.4167 | 7.377% | P | 0.9313
Gradko | 20% TEA in water | 2012 | R | Exeter City Council |  | 11 | 33.8757 | 33.9622 | -0.2546% | G | 1.0026
Gradko | 50% TEA in acetone | 2012 | R | London Borough of Croydon |  | 12 | 44.0702 | 43.9083 | 0.3685% | P | 0.99633
Gradko | 50% TEA in acetone | 2012 | KS | London Borough of Croydon |  | 12 | 59.681 | 66.0583 | -9.654% | G | 1.1069
Staffordshire Scientific Services | 20% TEA in water | 2012 | R | Stoke-on-Trent City Council |  | 12 | 42.9 | 34.0745 | 25.9% | G | 0.79428
ESG Didcot | 50% TEA in acetone | 2012 | R | Thanet District Council |  | 12 | 32.225 | 25.3333 | 27.2% | G | 0.78614
Gradko | 20% TEA in water | 2012 | R | Scarborough B C |  | 11 | 32.4861 | 36.6364 | -11.33% | G | 1.1278
Lambeth Scientific services | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 95.4444 | 94.5553 | 0.9403% | G | 0.99068
ESG Glasgow | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 11 | 132.5121 | 94.7163 | 39.9% | G | 0.71477
ESG Glasgow | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 11 | 129.9667 | 94.7163 | 37.22% | G | 0.72877
ESG Didcot | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 11 | 126.9848 | 94.7163 | 34.07% | G | 0.74589
ESG Didcot | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 11 | 129.0091 | 94.7163 | 36.21% | G | 0.73418
Milton Keynes Council | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 116.2472 | 94.5553 | 22.94% | G | 0.8134
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 122.4056 | 94.5553 | 29.45% | G | 0.77248
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 109.6945 | 94.5553 | 16.01% | G | 0.86199
Northampton BC | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 121.4861 | 94.5553 | 28.48% | G | 0.77832
Glasgow Scientific services | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 106.5611 | 94.5553 | 12.7% | G | 0.88733
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 129.5278 | 94.5553 | 36.99% | G | 0.73
Gradko | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 11 | 105.587 | 94.1812 | 12.11% | G | 0.89198
Gradko | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 95.7453 | 94.5553 | 1.258% | G | 0.98757
Aberdeen Scientific Services | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 11 | 114.0882 | 94.8083 | 20.34% | G | 0.83101
Staffordshire Scientific Services | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 105.8056 | 94.5553 | 11.9% | G | 0.89367
Kirklees Council | 50% TEA in acetone | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 119.6586 | 94.5553 | 26.55% | G | 0.79021
Tayside Scientific Services | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 105.2236 | 94.5553 | 11.28% | G | 0.89861
Exova | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 105.8517 | 94.5553 | 11.95% | P | 0.89328
Kent Scientific Services | 20% TEA in Water | 2012 | KS | Marylebone Road Intercomparison |  | 12 | 115.94 | 94.5553 | 22.62% | G | 0.81555
ESG Didcot | 50% TEA in acetone | 2012 | B | Stockton on Tees |  | 12 | 28.1944 | 21.2167 | 32.89% | G | 0.75251
ESG Didcot | 50% TEA in acetone | 2012 | R | Stockton on Tees |  | 11 | 22.3924 | 17.1545 | 30.53% | G | 0.76609
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2012 | R | Barnsley MBC |  | 12 | 43.9236 | 40 | 9.809% | G | 0.91067
Somerset County Council | 20% TEA in water | 2012 | R | Bath & North East Somerset |  | 12 | 58.7768 | 55.5528 | 5.803% | P | 0.94515
Staffordshire Scientific Services | 20% TEA in water | 2012 | UC | Manchester City Council |  | 10 | 46.4833 | 41.3196 | 12.5% | G | 0.88891
Staffordshire Scientific Services | 20% TEA in water | 2012 | SU | Manchester City Council |  | 12 | 30.0417 | 24.4835 | 22.7% | G | 0.81498
Staffordshire Scientific Services | 20% TEA in water | 2012 | KS | Manchester City Council |  | 12 | 76.7556 | 62.1858 | 23.43% | G | 0.81018
Glasgow Scientific Services | 20% TEA in water | 2012 | R | West Dunbartonshire Council |  | 12 | 26.7 | 24.475 | 9.091% | P | 0.91667
Glasgow Scientific Services | 20% TEA in water | 2012 | R | West Dunbartonshire Council |  | 12 | 22.2861 | 21.2767 | 4.744% | P | 0.95471
Gradko | 20% TEA in water | 2012 | KS | New Forest DC |  | 10 | 45.6737 | 40.26 | 13.45% | G | 0.88147
Staffordshire Scientific Services | 20% TEA in water | 2012 | O | South Staffordshire Council |  | 12 | 37.7639 | 38.8258 | -2.735% | G | 1.0281
Staffordshire Scientific Services | 20% TEA in water | 2012 | R | Stockport MBC |  | 11 | 33.8939 | 29.2727 | 15.79% | G | 0.86366
ESG Didcot | 50% TEA in acetone | 2012 | SU | Thanet District Council |  | 12 | 21.1806 | 18.1667 | 16.59% | G | 0.8577
Gradko | 20% TEA in water | 2012 | R | New Forest DC |  | 10 | 32.886 | 29.41 | 11.82% | G | 0.8943
Staffordshire Scientific Services | 20% TEA in water | 2012 | UB | WIGAN |  | 12 | 29 | 24.3333 | 19.18% | G | 0.83908
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | UC | Leeds City Council |  | 11 | 48.3636 | 36.8455 | 31.26% | G | 0.76184
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | R | Leeds City Council |  | 9 | 56.2222 | 43.3444 | 29.71% | S | 0.77095
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | R | Leeds City Council |  | 11 | 60.6364 | 47.0182 | 28.96% | S | 0.77541
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | R | Leeds City Council |  | 10 | 42.1 | 31.08 | 35.46% | S | 0.73824
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | R | Leeds City Council |  | 10 | 43.6 | 32.132 | 35.69% | G | 0.73697
Northampton BC | 20% TEA in water | 2012 | R | Northampton BC |  | 11 | 53.8697 | 36.4545 | 47.77% | G | 0.67672
Northampton BC | 20% TEA in water | 2012 | R | Northampton BC |  | 11 | 45.3091 | 36.6 | 23.8% | G | 0.80778
ESG Didcot | 50% TEA in acetone | 2012 | UB | CITY OF YORK COUNCIL |  | 12 | 27.5153 | 23.8729 | 15.26% | P | 0.86762
ESG Didcot | 50% TEA in acetone | 2012 | R | CITY OF YORK COUNCIL |  | 12 | 41.4829 | 31.7758 | 30.55% | P | 0.766
ESG Didcot | 50% TEA in acetone | 2012 | R | CITY OF YORK COUNCIL |  | 12 | 36.6605 | 27.8954 | 31.42% | G | 0.76091
ESG Didcot | 50% TEA in acetone | 2012 | R | CITY OF YORK COUNCIL |  | 12 | 40.638 | 30.2287 | 34.44% | G | 0.74385
ESG Didcot | 50% TEA in Acetone | 2012 | KS | Suffolk Coastal District Council |  | 12 | 49.8806 | 43.8333 | 13.8% | G | 0.87877
Gradko | 20% TEA in water | 2012 | R | Brighton & Hove City Council |  | 11 | 40.7667 | 36.9 | 10.48% | G | 0.90515
Gradko | 20% TEA in water | 2012 | R | City of Lincoln Council |  | 11 | 52.5829 | 44.4164 | 18.39% | G | 0.84469
Gradko | 20% TEA in water | 2012 | R | Fareham Borough Council |  | 9 | 37.5065 | 39.1 | -4.075% | G | 1.0425
Gradko | 20% TEA in water | 2012 | R | NOTTINGHAM CITY COUNCIL |  | 10 | 44.257 | 44.35 | -0.2097% | G | 1.0021
Gradko | 20% TEA in water | 2012 | R | NOTTINGHAM CITY COUNCIL |  | 11 | 43.1824 | 41.1636 | 4.904% | G | 0.95325
Gradko | 20% TEA in water | 2012 | R | NOTTINGHAM CITY COUNCIL |  | 10 | 46.3857 | 46.51 | -0.2673% | G | 1.0027
Gradko | 20% TEA in water | 2012 | R | The Highland Council |  | 9 | 24.2526 | 31.9333 | -24.05% | G | 1.3167
Gradko | 50% TEA in acetone | 2012 | R | Leeds City Council |  | 9 | 49.8889 | 38.4444 | 29.77% | S | 0.7706
Gradko | 50% TEA in Acetone | 2012 | UB | London Borough of Bexley |  | 11 | 24.2998 | 24.1655 | 0.5561% | G | 0.99447
Gradko | 50% TEA in Acetone | 2012 | SU | London Borough of Bexley |  | 10 | 26.1993 | 28.144 | -6.91% | G | 1.0742
Gradko | 50% TEA in Acetone | 2012 | SU | London Borough of Bexley |  | 11 | 26.8221 | 26.1209 | 2.684% | P | 0.97386
Gradko | 50% TEA in Acetone | 2012 | R | London Borough of Bexley |  | 11 | 48.7624 | 45.8891 | 6.261% | G | 0.94107
ESG Didcot | 50% TEA in Acetone | 2012 | R | Maidstone Borough Council |  | 12 | 48.3889 | 43.5 | 11.24% | P | 0.89897
ESG Didcot | 50% TEA in Acetone | 2012 | B | Maidstone Borough Council |  | 12 | 20.2194 | 13.9167 | 45.29% | G | 0.68828
Kirklees Council | 50% TEA in acetone | 2012 | UB | Kirklees Council |  | 12 | 27.7225 | 24.8533 | 11.54% | G | 0.8965
Kirklees Council | 50% TEA in acetone | 2012 | R | Kirklees Council |  | 10 | 45.997 | 37.8 | 21.69% | S | 0.82179
Kirklees Council | 50% TEA in acetone | 2012 | R | Kirklees Council |  | 11 | 56.9873 | 43.7573 | 30.23% | G | 0.76784
Kirklees Council | 50% TEA in acetone | 2012 | UC | Kirklees Council |  | 12 | 50.9108 | 36.6833 | 38.78% | S | 0.72054
ESG Didcot | 50% TEA in acetone | 2012 | R | Armagh City and District Council |  | 12 | 39.8333 | 27.4167 | 45.29% | G | 0.68828
ESG Didcot | 50% TEA in acetone | 2012 | R | Dumfries and Galloway Council |  | 12 | 37.7083 | 33.034 | 14.15% | G | 0.87604
Gradko | 20% TEA in water | 2012 | R | Wiltshire Council |  | 10 | 36.0667 | 34.7 | 3.939% | G | 0.96211
Gradko | 50% TEA in acetone | 2012 | R | West Berkshire Council |  | 11 | 42.3152 | 45.5073 | -7.015% | G | 1.0754
Gradko | 50% TEA in acetone | 2012 | R | West Berkshire Council |  | 12 | 25.2621 | 27.01 | -6.471% | G | 1.0692
Staffordshire Scientific Services | 20% TEA in Water | 2012 | R | Salford City Council |  | 12 | 59.2111 | 61.1719 | -3.205% | G | 1.0331
ESG Didcot | 50% TEA in acetone | 2012 | R | Cambridge City Council |  | 12 | 45.9111 | 34.9167 | 31.49% | G | 0.76053
ESG Didcot | 50% TEA in Acetone | 2012 | R | Swale Borough Council |  | 11 | 43.8712 | 31.6364 | 38.67% | G | 0.72112
Gradko | 20% TEA in Water | 2012 | UB | LB Waltham Forest |  | 11 | 33.2009 | 37.6364 | -11.79% | S | 1.1336
Gradko | 50% TEA in acetone | 2012 | R | Wolverhampton CC |  | 12 | 39.6131 | 49.1667 | -19.43% | P | 1.2412
Gradko | 50% TEA in acetone | 2012 | R | Wolverhampton CC |  | 12 | 37.3711 | 33.6369 | 11.1% | G | 0.90008
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | R | Stirling Council |  | 11 | 30.1091 | 28.8182 | 4.479% | G | 0.95713
Somerset County Council | 20% TEA in water | 2012 | R | South Gloucestershire Council |  | 10 | 28.5717 | 27.03 | 5.704% | P | 0.94604
Gradko | 20% TEA in water | 2012 | R | Pendle |  | 10 | 38.616 | 32.01 | 20.64% | G | 0.82893
ESG Didcot | 50% TEA in acetone | 2012 | R | Northumberland County Council |  | 12 | 36.1667 | 27.5833 | 31.12% | S | 0.76267
ESG Didcot | 50% TEA in acetone | 2012 | R | North Down Borough Council |  | 12 | 44.9722 | 32.9167 | 36.62% | G | 0.73193
ESG Didcot | 50% TEA in acetone | 2012 | R | Medway Council |  | 11 | 38.103 | 32.0909 | 18.73% | G | 0.84221
ESG Didcot | 50% TEA in acetone | 2012 | UB | Medway Council |  | 11 | 24.9318 | 24.3636 | 2.332% | G | 0.97721
ESG Didcot | 50% TEA in acetone | 2012 | B | Medway Council |  | 10 | 28.59 | 18.9 | 51.27% | P | 0.66107
ESG Didcot | 50% TEA in acetone | 2012 | R | Lisburn City Council |  | 9 | 27.963 | 24.6667 | 13.36% | P | 0.88212
Gradko | 20% TEA in water | 2012 | R | Lancaster City Council |  | 10 | 46.412 | 41.699 | 11.3% | G | 0.89845
Gradko | 20% TEA in water | 2012 | R | Lancaster City Council |  | 11 | 36.7542 | 35.9844 | 2.139% | G | 0.97906
Lambeth Scientific Services | 50% TEA in acetone | 2012 | R | Elmbridge BC |  | 10 | 50.5 | 40.99 | 23.2% | P | 0.81168
Glasgow Scientific Services | 20% TEA in Water | 2012 | R | East Ayrshire Council |  | 10 | 29.0233 | 29.8 | -2.606% | G | 1.0268
Tayside Scientific Services | 20% TEA in water | 2012 | R | Dundee City Council |  | 11 | 55.3848 | 54 | 2.565% | G | 0.975
Tayside Scientific Services | 20% TEA in water | 2012 | UB | Dundee City Council |  | 11 | 17.0364 | 10.4545 | 62.96% | G | 0.61366
Tayside Scientific Services | 20% TEA in water | 2012 | R | Dundee City Council |  | 11 | 64.6061 | 55.1818 | 17.08% | G | 0.85413
Tayside Scientific Services | 20% TEA in water | 2012 | R | Dundee City Council |  | 10 | 55.1267 | 49.2 | 12.05% | G | 0.89249
Tayside Scientific Services | 20% TEA in water | 2012 | R | Dundee City Council |  | 9 | 40.8444 | 33.3333 | 22.53% | G | 0.8161
Tayside Scientific Services | 20% TEA in water | 2012 | R | Dundee City Council |  | 10 | 46.9633 | 44.3 | 6.012% | G | 0.94329
ESG Didcot | 50% TEA in acetone | 2012 | R | Down District Council |  | 11 | 49.8485 | 37.7273 | 32.13% | G | 0.75684
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | R | City Of Edinburgh Council |  | 11 | 45.9348 | 39.3636 | 16.69% | G | 0.85694
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | R | City Of Edinburgh Council |  | 10 | 73.03 | 51.7 | 41.26% | G | 0.70793
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | R | City Of Edinburgh Council |  | 12 | 43.2944 | 28.3333 | 52.8% | G | 0.65443
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | R | City Of Edinburgh Council |  | 12 | 37.7792 | 30.25 | 24.89% | G | 0.80071
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | KS | City Of Edinburgh Council |  | 11 | 75.7076 | 56.2727 | 34.54% | G | 0.74329
ESG Didcot | 50% TEA in acetone | 2012 | R | Castlereagh Borough Council |  | 12 | 47.8889 | 29.5833 | 61.88% | G | 0.61775
ESG Didcot | 50% TEA in acetone | 2012 | R | Bridgend County Borough Council |  | 12 | 28.0278 | 26.8333 | 4.451% | G | 0.95738
Gradko | 20% TEA in water | 2012 | R | Wokingham Borough Council |  | 9 | 31.513 | 34.0444 | -7.436% | G | 1.0803
ESG Didcot | 50% TEA in acetone | 2012 | R | Tunbridge Wells BC |  | 12 | 61.5819 | 48.5 | 26.97% | G | 0.78757
Gradko | 20% TEA in water | 2012 | R | London Borough of Ealing |  | 10 | 54.6635 | 53.7111 | 1.773% | P | 0.98258
Gradko | 20% TEA in water | 2012 | R | London Borough of Ealing |  | 10 | 83.3208 | 83.5338 | -0.2549% | P | 1.0026
Gradko | 20% TEA in water | 2012 | UB | London Borough of Ealing |  | 9 | 32.2882 | 36.2132 | -10.84% | G | 1.1216
ESG Didcot | 50% TEA in Acetone | 2012 | R | West Oxfordshire District Council (WODC) |  | 11 | 48.1364 | 36.3909 | 32.28% | G | 0.756
Gradko | 50% TEA in acetone | 2012 | R | Sandwell Metropolitan Borough Council |  | 11 | 45.6436 | 45.3545 | 0.6374% | G | 0.99367
Gradko | 20% TEA in water | 2012 | B | Chelmsford City Council |  | 11 | 18.8689 | 14.4121 | 30.92% | G | 0.7638
ESG Didcot | 20% TEA in water | 2013 | KS | South Lakeland District Council |  | 9 | 44.4074 | 34.3333 | 29.34% | G | 0.77314
Glasgow Scientific Services | 20% TEA in water | 2013 | R | East Dunbartonshire Council |  | 12 | 29.4847 | 30.75 | -4.115% | P | 1.0429
Glasgow Scientific Services | 20% TEA in water | 2013 | R | East Dunbartonshire Council |  | 12 | 39.7306 | 36.4167 | 9.1% | G | 0.91659
Glasgow Scientific Services | 20% TEA in water | 2013 | R | East Dunbartonshire Council |  | 10 | 32.9933 | 36.2 | -8.858% | P | 1.0972
Glasgow Scientific Services | 20% TEA in water | 2013 | R | East Dunbartonshire Council |  | 11 | 25.4788 | 26.7273 | -4.671% | G | 1.049
Gradko | 20% TEA in water | 2013 | R | Cheshire West and Chester |  | 12 | 38.9961 | 40.77 | -4.351% | G | 1.0455
Gradko | 20% TEA in Water | 2013 | R | Dudley MBC |  | 12 | 38.3657 | 31.1546 | 23.15% | G | 0.81204
Gradko | 20% TEA in Water | 2013 | UB | Dudley MBC |  | 10 | 24.9342 | 25.3739 | -1.733% | G | 1.0176
Gradko | 20% TEA in Water | 2013 | R | Dudley MBC |  | 11 | 41.1767 | 39.0854 | 5.351% | G | 0.94921
Somerset County Council | 20% TEA in water | 2013 | R | South Gloucestershire Council |  | 11 | 28.7803 | 26.3364 | 9.28% | G | 0.91508
Staffordshire Scientific Services | 20% TEA in water | 2013 | UC | Manchester City Council |  | 11 | 42 | 39.3448 | 6.749% | G | 0.93678
Staffordshire Scientific Services | 20% TEA in water | 2013 | SU | Manchester City Council |  | 11 | 27.1212 | 22.4859 | 20.61% | G | 0.82909
Staffordshire Scientific Services | 20% TEA in water | 2013 | KS | Manchester City Council |  | 10 | 68.4667 | 54.3629 | 25.94% | G | 0.79401
Staffordshire Scientific Services | 20% TEA in water | 2013 | R | Cannock Chase Council |  | 12 | 47.8389 | 36.2 | 32.15% | G | 0.75671
Gradko | 20% TEA in water | 2013 | R | East Herts Council |  | 10 | 35.2333 | 29.5 | 19.44% | G | 0.83728
Gradko | 20% TEA in water | 2013 | R | Fareham Borough Council |  | 9 | 34.3833 | 33.7222 | 1.96% | G | 0.98077
Gradko | 20% TEA in water | 2013 | R | Fareham Borough Council |  | 12 | 42.2764 | 45.0583 | -6.174% | G | 1.0658
Gradko | 20% TEA in water | 2013 | R | Gateshead Council |  | 11 | 33.797 | 37.0273 | -8.724% | G | 1.0956
Gradko | 20% TEA in water | 2013 | R | Gateshead Council |  | 11 | 34.7121 | 32.6545 | 6.301% | G | 0.94072
Gradko | 20% TEA in water | 2013 | R | Gateshead Council |  | 10 | 32.7167 | 32.03 | 2.144% | G | 0.97901
Gradko | 20% TEA in water | 2013 | R | Borough Council of King's Lynn & West Norfolk |  | 12 | 28.7153 | 25.5167 | 12.54% | G | 0.88861
Gradko | 20% TEA in water | 2013 | R | Gedling Borough Council |  | 10 | 37.0167 | 34.544 | 7.158% | G | 0.9332
Gradko | 20% TEA in water | 2013 | R | The Highland Council |  | 12 | 23.8664 | 20.9092 | 14.14% | G | 0.87609
ESG Didcot | 50% TEA in acetone | 2013 | B | Gravesham Borough Council |  | 11 | 33.3303 | 31.8182 | 4.752% | G | 0.95463
ESG Didcot | 50% TEA in acetone | 2013 | B | Gravesham Borough Council |  | 12 | 44.3944 | 31.75 | 39.83% | G | 0.71518
Gradko | 50% TEA in acetone | 2013 | R | West Berkshire Council |  | 12 | 39.3725 | 42.225 | -6.755% | G | 1.0724
Gradko | 50% TEA in acetone | 2013 | R | West Berkshire Council |  | 10 | 21.4707 | 23.717 | -9.471% | G | 1.1046
ESG Didcot | 50% TEA in acetone | 2013 | R | Falkirk Council |  | 12 | 34.9194 | 30.5444 | 14.32% | G | 0.87471
ESG Didcot | 50% TEA in acetone | 2013 | UB | Falkirk Council |  | 12 | 25.0972 | 20.4762 | 22.57% | G | 0.81588
ESG Didcot | 50% TEA in acetone | 2013 | B | Pembrokeshire Council |  | 12 | 6.9417 | 5.9167 | 17.32% | P | 0.85234
ESG Didcot | 50% TEA in acetone | 2013 | UB | Medway |  | 12 | 24.2778 | 25.1667 | -3.532% | G | 1.0366
ESG Didcot | 50% TEA in acetone | 2013 | R | Medway Council |  | 10 | 36.4433 | 26.7 | 36.49% | G | 0.73264
ESG Didcot | 50% TEA in acetone | 2013 | B | Medway |  | 11 | 25.8864 | 14 | 84.9% | P | 0.54083
Gradko | 50% TEA in acetone | 2013 | R | London Borough of Richmond upon Thames |  | 11 | 41.4788 | 39.7273 | 4.409% | G | 0.95777
Gradko | 50% TEA in acetone | 2013 | R | London Borough of Richmond upon Thames |  | 9 | 50.058 | 43.2222 | 15.82% | G | 0.86344
Gradko | 50% TEA in acetone | 2013 | B | London Borough of Richmond upon Thames |  | 11 | 25.9009 | 24.7273 | 4.746% | G | 0.95469
Gradko | 50% TEA in Acetone | 2013 | SU | London Borough of Bexley |  | 12 | 27.8861 | 27.2908 | 2.181% | G | 0.97865
Gradko | 50% TEA in Acetone | 2013 | R | London Borough of Bexley |  | 11 | 59.9897 | 51.34 | 16.85% | G | 0.85581
Gradko | 50% TEA in acetone | 2013 | R | East Hampshire District Council |  | 12 | 24.4112 | 23.5333 | 3.731% | G | 0.96404
Gradko | 50% TEA in acetone | 2013 | R | Royal Borough Windsor and Maidenhead |  | 10 | 35.4139 | 46.21 | -23.36% | G | 1.3049
Gradko | 50% TEA in acetone | 2013 | R | Royal Borough Windsor and Maidenhead |  | 12 | 41.0538 | 45.75 | -10.26% | G | 1.1144
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2013 | R | Barnsley MBC |  | 12 | 42.5926 | 36.8333 | 15.64% | G | 0.86478
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | R | Calderdale MBC |  | 12 | 61.5278 | 42.975 | 43.17% | G | 0.69847
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | UB | Wakefield MDC |  | 12 | 40.2083 | 31.0544 | 29.48% | G | 0.77234
ESG Didcot | 50% TEA in acetone | 2013 | R | Wrexham County Borough Council |  | 12 | 23.4361 | 21.6417 | 8.292% | G | 0.92343
ESG Didcot | 50% TEA in Acetone | 2013 | UI | Stockton on Tees |  | 12 | 27.3222 | 20.1583 | 35.54% | G | 0.7378
ESG Didcot | 50% TEA in Acetone | 2013 | R | Stockton on Tees |  | 12 | 21.1222 | 16.1083 | 31.13% | G | 0.76262
ESG Didcot | 50% TEA in acetone | 2013 | SU | Thanet District Council |  | 11 | 20.6 | 15.9091 | 29.49% | P | 0.77229
ESG Didcot | 50% TEA in acetone | 2013 | R | Thanet District Council |  | 11 | 28.6091 | 24.2727 | 17.87% | P | 0.84843
Gradko | 20% TEA in Water | 2013 | R | Dudley MBC |  | 12 | 51.8436 | 58.8973 | -11.98% | P | 1.1361
Staffordshire Scientific Services | 20% TEA in water | 2013 | UC | Stoke-on-Trent City Council |  | 11 | 30.6697 | 28.7991 | 6.495% | G | 0.93901
Staffordshire Scientific Services | 20% TEA in water | 2013 | R | Stoke-on-Trent City Council |  | 11 | 56.6182 | 48.9 | 15.78% | G | 0.86368
Staffordshire Scientific Services | 20% TEA in water | 2013 | R | Stoke-on-Trent City Council |  | 11 | 38.8515 | 32.4664 | 19.67% | G | 0.83565
Somerset County Council | 20% TEA in water | 2013 | R | Bath & North East Somerset |  | 12 | 56.0231 | 56.6789 | -1.157% | G | 1.0117
ESG Didcot | 50% TEA in acetone | 2013 | R | Cambridge City Council |  | 12 | 46.4278 | 34.8333 | 33.29% | G | 0.75027
ESG Didcot | 50% TEA in acetone | 2013 | R | Swale Borough Council |  | 10 | 45.27 | 41.4 | 9.348% | G | 0.91451
ESG Didcot | 50% TEA in acetone | 2013 | R | Swale Borough Council |  | 12 | 39.9278 | 34.4167 | 16.01% | P | 0.86197
ESG Didcot | 50% TEA in acetone | 2013 | R | Swale Borough Council |  | 12 | 41.2528 | 39.6667 | 3.999% | G | 0.96155
ESG Didcot | 50% TEA in acetone | 2013 | R | Swale Borough Council |  | 11 | 53.0455 | 34.2727 | 54.77% | G | 0.6461
Gradko | 50% TEA in acetone | 2013 | R | Worthing Borough Council |  | 12 | 41.2125 | 40.1458 | 2.657% | G | 0.97412
Gradko | 50% TEA in acetone | 2013 | R | London Borough of Croydon |  | 9 | 47.673 | 49.8222 | -4.314% | P | 1.0451
Gradko | 50% TEA in acetone | 2013 | KS | London Borough of Croydon |  | 10 | 69.0662 | 70.83 | -2.49% | G | 1.0255
ESG Didcot | 50% TEA in acetone | 2013 | R | North East Lincolnshire Council |  | 11 | 58.9894 | 49.3573 | 19.52% | G | 0.83671
ESG Didcot | 50% TEA in acetone | 2013 | R | North East Lincolnshire Council |  | 11 | 33.9685 | 30.2373 | 12.34% | G | 0.89016
ESG Didcot | 50% TEA in acetone | 2013 | R | North East Lincolnshire Council |  | 11 | 39.5394 | 31.1482 | 26.94% | G | 0.78778
ESG Didcot | 50% TEA in acetone | 2013 | R | Rugby Borough Council |  | 12 | 25.5583 | 21.9122 | 16.64% | P | 0.85734
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | UC | Leeds City Council 1 |  | 12 | 45.3611 | 34.3667 | 31.99% | G | 0.75762
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | KS | Leeds City Council 2 |  | 10 | 54.5 | 44.87 | 21.46% | S | 0.8233
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2013 | UC | Barnsley MBC |  | 12 | 36.4043 | 34.325 | 6.058% | S | 0.94288
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | R | Leeds City Council 8 |  | 11 | 49.2727 | 36.9091 | 33.5% | S | 0.74908
Staffordshire Scientific Services | 20% TEA in water | 2013 | R | Stockport |  | 9 | 31.8426 | 28.2222 | 12.83% | G | 0.8863
Gradko | 20% TEA in water | 2013 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 43.0067 | 43.9583 | -2.165% | G | 1.0221
Gradko | 20% TEA in water | 2013 | R | NOTTINGHAM CITY COUNCIL |  | 10 | 41.179 | 38.71 | 6.378% | G | 0.94004
Gradko | 20% TEA in water | 2013 | R | NOTTINGHAM CITY COUNCIL |  | 11 | 42.8061 | 41.9909 | 1.941% | G | 0.98096
Gradko | 20% TEA in water | 2013 | R | Brighton & Hove City Council |  | 11 | 61.5091 | 60.3718 | 1.884% | G | 0.98151
Gradko | 20% TEA in water | 2013 | R | Brighton & Hove City Council |  | 11 | 40.7652 | 29.641 | 37.53% | G | 0.72712
Kirklees Council | 50% TEA in acetone | 2013 | UC | Kirklees Council |  | 10 | 46.41 | 34.84 | 33.21% | S | 0.7507
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | R | Leeds City Council 9 |  | 11 | 41.3182 | 43.9273 | -5.94% | P | 1.0631
Gradko | 50% TEA in acetone | 2013 | UB | Norwich City Council |  | 11 | 13.1082 | 14.5455 | -9.881% | G | 1.1096
Staffordshire Scientific Services | 20% TEA in water | 2013 | UI | Salford City Council |  | 12 | 32.8521 | 30.0544 | 9.309% | G | 0.91484
Staffordshire Scientific Services | 20% TEA in water | 2013 | R | Salford City Council |  | 10 | 58.3332 | 62.0008 | -5.915% | G | 1.0629
Lambeth Scientific Services | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 97.1528 | 80.6083 | 20.52% | G | 0.82971
ESG Glasgow | 20% TEA in water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 111.9157 | 80.6083 | 38.84% | G | 0.72026
ESG Glasgow | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 110.6617 | 80.6083 | 37.28% | G | 0.72842
ESG Didcot | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 108.6885 | 80.6083 | 34.84% | G | 0.74165
ESG Didcot | 20% TEA in water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 107.8419 | 80.6083 | 33.79% | G | 0.74747
Milton Keynes Council | 20% TEA in water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 96.3833 | 80.6083 | 19.57% | G | 0.83633
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 110.9827 | 80.6083 | 37.68% | G | 0.72631
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 102.2944 | 80.6083 | 26.9% | G | 0.788
Northampton BC | 20% TEA in water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 113.1401 | 80.6083 | 40.36% | G | 0.71246
Glasgow Scientific Services | 20% TEA in water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 94.8 | 80.6083 | 17.61% | G | 0.8503
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 114.4444 | 80.6083 | 41.98% | G | 0.70434
Gradko | 20% TEA in water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 101.3688 | 80.6083 | 25.75% | G | 0.7952
Gradko | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 11 | 89.7595 | 81.1909 | 10.55% | G | 0.90454
Aberdeen Scientific Services | 20% TEA in Water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 97.3336 | 80.6083 | 20.75% | G | 0.82817
Staffordshire Scientific Services | 20% TEA in Water | 2013 | KS | Marylebone Road Intercomparison |  | 11 | 96.8025 | 80.3455 | 20.48% | G | 0.82999
Kirklees Council | 50% TEA in acetone | 2013 | KS | Marylebone Road Intercomparison |  | 11 | 112.5028 | 81.1909 | 38.57% | G | 0.72168
Tayside Scientific Services | 20% TEA in Water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 102.85 | 80.6083 | 27.59% | G | 0.78375
Exova | 20% TEA in water | 2013 | KS | Marylebone Road Intercomparison |  | 9 | 88.9667 | 80.8222 | 10.08% | G | 0.90846
Kent Scientific Services | 20% TEA in Water | 2013 | KS | Marylebone Road Intercomparison |  | 12 | 105.0686 | 80.6083 | 30.34% | G | 0.7672
Somerset County Council | 20% TEA in Water | 2013 | KS | Marylebone Road Intercomparison |  | 11 | 98.7674 | 80.2909 | 23.01% | G | 0.81293
Gradko | 20% TEA in Water | 2013 | R | Brighton & Hove City Council |  | 9 | 53.8741 | 45.0413 | 19.61% | G | 0.83605
ESG Didcot | 50% TEA in acetone | 2013 | UB | City of York Council |  | 11 | 24.9791 | 19.2658 | 29.66% | G | 0.77128
ESG Didcot | 50% TEA in acetone | 2013 | R | City of York Council |  | 12 | 40.0375 | 28.3636 | 41.16% | G | 0.70843
ESG Didcot | 50% TEA in acetone | 2013 | R | City of York Council |  | 12 | 33.6289 | 24.3776 | 37.95% | G | 0.7249
ESG Didcot | 50% TEA in acetone | 2013 | R | City of York Council |  | 10 | 39.6629 | 30.9182 | 28.28% | G | 0.77953
Gradko | 20% TEA in water | 2013 | R | Wiltshire Council |  | 12 | 40.1111 | 36.4167 | 10.14% | G | 0.90789
ESG Didcot | 50% TEA in acetone | 2013 | KS | Suffolk Coastal District Council |  | 11 | 45.8364 | 41 | 11.8% | G | 0.89449
Gradko | 20% TEA in water | 2013 | R | Wiltshire Council |  | 11 | 41.4067 | 37.0909 | 11.64% | G | 0.89577
Gradko | 50% TEA in acetone | 2013 | R | Wolverhampton CC |  | 10 | 43.4667 | 40.9 | 6.275% | G | 0.94095
Gradko | 50% TEA in acetone | 2013 | R | Wolverhampton CC |  | 12 | 35.0278 | 32.25 | 8.613% | G | 0.9207
Gradko | 20% TEA in water | 2013 | R | Wiltshire Council |  | 12 | 39.0556 | 48.8333 | -20.02% | G | 1.2504
Northampton BC | 20% TEA in water | 2013 | R | Northampton BC |  | 12 | 49.6417 | 32.7667 | 51.5% | G | 0.66006
Northampton BC | 20% TEA in water | 2013 | UB | Northampton BC |  | 12 | 19.2 | 14.0583 | 36.57% | G | 0.7322
Northampton BC | 20% TEA in water | 2013 | R | Northampton BC |  | 10 | 41.7033 | 34.18 | 22.01% | G | 0.8196
Lambeth Scientific Services | 50% TEA in acetone | 2013 | R | Elmbridge BC |  | 12 | 42.8603 | 46.5333 | -7.893% | G | 1.0857
Lambeth Scientific Services | 50% TEA in acetone | 2013 | R | Elmbridge BC |  | 12 | 49.9824 | 43.0583 | 16.08% | G | 0.86147
Gradko | 20% TEA in water | 2013 | R | Breckland Council |  | 12 | 32.0667 | 33.175 | -3.341% | G | 1.0346
ESG Didcot | 20% TEA in water | 2013 | R | Rhondda Cynon Taf CBC |  | 10 | 32.9067 | 35.051 | -6.118% | G | 1.0652
Lambeth Scientific Services | 50% TEA in acetone | 2013 | SU | Reigate and Banstead BC |  | 12 | 28.75 | 21.5902 | 33.16% | G | 0.75096
Lambeth Scientific Services | 50% TEA in acetone | 2013 | SU | Reigate and Banstead BC |  | 12 | 34.0833 | 28.324 | 20.33% | G | 0.83102
Lambeth Scientific Services | 50% TEA in acetone | 2013 | B | Reigate and Banstead BC |  | 12 | 22.7778 | 19.4069 | 17.37% | G | 0.85201
ESG Didcot | 50% TEA in acetone | 2013 | R | Maidstone Borough Council |  | 12 | 47.9375 | 47.25 | 1.455% | P | 0.98566
Gradko | 50% TEA in acetone | 2013 | UB | Reading Borough Council |  | 12 | 23.3106 | 26.5833 | -12.31% | G | 1.1404
Gradko | 50% TEA in acetone | 2013 | R | Reading Borough Council |  | 12 | 40.8106 | 43.6667 | -6.541% | G | 1.07
Gradko | 20% TEA in water | 2013 | R | City of Lincoln Council |  | 12 | 43.4125 | 43.1833 | 0.5307% | G | 0.99472
Gradko | 20% TEA in water | 2013 | R | Monmouthshire County Council |  | 12 | 40.9488 | 34.4118 | 19% | G | 0.84036
ESG Didcot | 50% TEA in acetone | 2013 | R | Dacorum Borough Council |  | 11 | 33.7515 | 27 | 25.01% | G | 0.79996
Gradko | 20% TEA in water | 2013 | R | Lancaster City Council |  | 12 | 44.0569 | 40.081 | 9.92% | G | 0.90976
Gradko | 20% TEA in water | 2013 | R | Lancaster City Council |  | 12 | 36.4392 | 34.332 | 6.138% | G | 0.94217
Lambeth Scientific Services | 50% TEA in acetone | 2013 | R | LB Haringey |  | 11 | 46.7273 | 42.9091 | 8.898% | S | 0.91829
Gradko | 20% TEA in water | 2013 | UB | Luton Borough Council |  | 12 | 35.6111 | 33.25 | 7.101% | G | 0.9337
ESG Didcot | 50% TEA in acetone | 2013 | R | Vale Of White Horse District Council |  | 12 | 37.7222 | 31.1 | 21.29% | G | 0.82445
ESG Didcot | 50% TEA in acetone | 2013 | R | Castlereagh Borough Council |  | 11 | 49.3636 | 32.1818 | 53.39% | G | 0.65193
ESG Didcot | 50% TEA in acetone | 2013 | R | North Down Borough Council |  | 11 | 40.4848 | 29.6364 | 36.61% | G | 0.73204
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | R | City Of Edinburgh Council |  | 12 | 43.1833 | 27.5 | 57.03% | G | 0.63682
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | R | City Of Edinburgh Council |  | 9 | 43.8741 | 38.3333 | 14.45% | G | 0.87371
Gradko | 20% TEA in water | 2013 | R | Pendle |  | 12 | 35.0106 | 38.3417 | -8.688% | P | 1.0951
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | R | City Of Edinburgh Council |  | 10 | 65.5633 | 46.6 | 40.69% | G | 0.71076
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | R | City Of Edinburgh Council |  | 11 | 40.9091 | 28.7273 | 42.41% | G | 0.70222
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | R | City Of Edinburgh Council |  | 11 | 36.8545 | 28.7273 | 28.29% | G | 0.77948
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | KS | City Of Edinburgh Council |  | 12 | 75.3917 | 58.1667 | 29.61% | G | 0.77153
ESG Didcot | 50% TEA in acetone | 2013 | R | West Lothian Council |  | 10 | 40.66 | 41.29 | -1.526% | G | 1.0155
ESG Didcot | 50% TEA in acetone | 2013 | R | Watford Borough Council |  | 12 | 42.4583 | 40.7667 | 4.15% | S | 0.96016
ESG Didcot | 50% TEA in acetone | 2013 | R | Dumfries and Galloway Council |  | 12 | 35.5583 | 30.3028 | 17.34% | G | 0.8522
Gradko | 20% TEA in water | 2013 | R | North Ayrshire Council |  | 12 | 32.2108 | 32.25 | -0.1214% | G | 1.0012
Gradko | 20% TEA in water | 2013 | KS | New Forest DC |  | 11 | 45.6339 | 40.2364 | 13.41% | G | 0.88172
Gradko | 20% TEA in water | 2013 | R | New Forest District Council |  | 11 | 33.5473 | 28.0909 | 19.42% | G | 0.83735
Gradko | 20% TEA in water | 2013 | UB | southampton city council |  | 12 | 30.1397 | 30.2833 | -0.4742% | G | 1.0048
ESG Didcot | 50% TEA in acetone | 2013 | R | Tunbridge Wells BC |  | 12 | 62.4181 | 47.0833 | 32.57% | G | 0.75432
ESG Didcot | 50% TEA in acetone | 2013 | R | Horsham District Council |  | 11 | 33.8515 | 23.8818 | 41.75% | G | 0.70549
ESG Didcot | 50% TEA in acetone | 2013 | R | Horsham District Council |  | 12 | 32.3778 | 29.8333 | 8.529% | G | 0.92141
ESG Didcot | 50% TEA in acetone | 2013 | R | Horsham District Council |  | 12 | 32.2833 | 26.425 | 22.17% | G | 0.81853
Gradko | 20% TEA in water | 2013 | UC | Belfast City Council |  | 11 | 33.1821 | 31.2114 | 6.314% | G | 0.94061
ESG Didcot | 50% TEA in acetone | 2013 | UI | Slough Borough Council |  | 12 | 42.2722 | 37.5 | 12.73% | G | 0.88711
ESG Didcot | 50% TEA in acetone | 2013 | UB | Slough Borough Council |  | 11 | 33.4561 | 29.2727 | 14.29% | G | 0.87496
ESG Didcot | 50% TEA in acetone | 2013 | R | Slough Borough Council |  | 11 | 37.2758 | 35.4545 | 5.137% | G | 0.95114
Gradko | 50% TEA in acetone | 2013 | SU | London Borough of Bexley |  | 12 | 27.8861 | 27.9767 | -0.3237% | G | 1.0032
ESG Didcot | 50% TEA in acetone | 2013 | R | West Oxfordshire District Council (WODC) |  | 10 | 51.3533 | 39.62 | 29.61% | G | 0.77152
Somerset County Council | 20% TEA in water | 2014 | R | Bath & North East Somerset |  | 12 | 52.8343 | 56.7509 | -6.901% | G | 1.0741
Gradko | 20% TEA in water | 2014 | UC | Belfast City Council |  | 11 | 33.3955 | 31.6273 | 5.591% | G | 0.94705
Gradko | 20% TEA in water | 2014 | R | Borough Council of King's Lynn & West Norfolk |  | 12 | 29.154 | 21.1667 | 37.74% | G | 0.72603
Gradko | 20% TEA in water | 2014 | R | Brighton & Hove City Council |  | 12 | 55.3511 | 48.0676 | 15.15% | G | 0.86841
Gradko | 20% TEA in water | 2014 | R | Brighton & Hove City Council |  | 11 | 60.0388 | 56.5496 | 6.17% | G | 0.94188
Somerset County Council | 20% TEA in water | 2014 | UB | Bristol City Council |  | 12 | 33.5417 | 31.4167 | 6.764% | G | 0.93664
Somerset County Council | 20% TEA in water | 2014 | R | Bristol City Council |  | 12 | 54.7951 | 52.25 | 4.871% | G | 0.95355
Somerset County Council | 20% TEA in water | 2014 | R | Bristol City Council |  | 12 | 45.6177 | 45.4167 | 0.4426% | G | 0.99559
Somerset County Council | 20% TEA in water | 2014 | UB | Bristol City Council |  | 9 | 29.9168 | 26.2222 | 14.09% | G | 0.8765
Somerset County Council | 20% TEA in water | 2014 | KS | Bristol City Council |  | 11 | 48.4836 | 40.5455 | 19.58% | G | 0.83627
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | R | Calderdale MBC |  | 10 | 62.8333 | 43.6 | 44.11% | P | 0.6939
ESG Didcot | 50% TEA in acetone | 2014 | R | Cambridge City Council |  | 12 | 46.75 | 37.25 | 25.5% | G | 0.79679
Staffordshire Scientific Services | 20% TEA in water | 2014 | R | Cannock Chase Council |  | 10 | 48.5067 | 34.38 | 41.09% | G | 0.70877
Gradko | 20% TEA in water | 2014 | R | Cheshire West and Chester |  | 11 | 40.0018 | 40.4182 | -1.03% | G | 1.0104
Gradko | 20% TEA in water | 2014 | R | Dudley MBC |  | 12 | 36.1926 | 30.6573 | 18.06% | G | 0.84706
Gradko | 20% TEA in water | 2014 | UB | Dudley MBC |  | 12 | 25.7562 | 23.1687 | 11.17% | G | 0.89954
Gradko | 20% TEA in water | 2014 | R | Dudley MBC |  | 12 | 40.7578 | 35.3738 | 15.22% | G | 0.8679
Gradko | 20% TEA in water | 2014 | R | Dudley MBC |  | 12 | 52.1122 | 59.6347 | -12.61% | G | 1.1444
Gradko | 50% TEA in acetone | 2014 | R | East Hampshire District Council |  | 12 | 25.1186 | 22.6417 | 10.94% | G | 0.90139
Staffordshire Scientific Services | 20% TEA in water | 2014 | R | East Staffordshire Borough Council |  | 11 | 49.2303 | 36.9091 | 33.38% | G | 0.74972
ESG Didcot | 50% TEA in acetone | 2014 | UB | Falkirk |  | 12 | 23.45 | 19.4585 | 20.51% | G | 0.82979
Tayside Scientific Services | 20% TEA in water | 2014 | R | Fife Council |  | 9 | 35.2741 | 27.8011 | 26.88% | G | 0.78815
Tayside Scientific Services | 20% TEA in water | 2014 | R | Fife Council |  | 12 | 23.5806 | 18.7817 | 25.55% | G | 0.79649
Tayside Scientific Services | 20% TEA in water | 2014 | R | Fife Council |  | 10 | 30.4067 | 25.403 | 19.7% | G | 0.83544
Tayside Scientific Services | 20% TEA in water | 2014 | R | Fife Council |  | 11 | 34.2061 | 24.4827 | 39.72% | G | 0.71574
Gradko | 20% TEA in water | 2014 | R | Gateshead Council |  | 10 | 35.1567 | 31.73 | 10.8% | G | 0.90253
Gradko | 20% TEA in water | 2014 | R | Gateshead Council |  | 12 | 35.8389 | 35.8917 | -0.147% | G | 1.0015
Gradko | 20% TEA in water | 2014 | R | Gateshead Council |  | 12 | 34.2361 | 32.1667 | 6.434% | G | 0.93955
Glasgow Scientific Services | 20% TEA in water | 2014 | KS | Glasgow City Council |  | 10 | 74.5633 | 65.09 | 14.55% | P | 0.87295
ESG Didcot | 50% TEA in acetone | 2014 | UB | Gravesham Borough Council |  | 12 | 27.3306 | 24.5 | 11.55% | P | 0.89643
ESG Didcot | 50% TEA in acetone | 2014 | UB | Gravesham Borough Council |  | 12 | 40.3933 | 31.1667 | 29.6% | G | 0.77158
ESG Didcot | 50% TEA in acetone | 2014 | UB | Kingston upon Hull City Council |  | 12 | 31.5611 | 25.75 | 22.57% | G | 0.81588
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | UC | Leeds City Council |  | 12 | 46.5556 | 38.5667 | 20.71% | G | 0.8284
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | KS | Leeds City Council |  | 10 | 52.3 | 42.94 | 21.8% | S | 0.82103
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | R | Leeds City Council |  | 10 | 40.35 | 26.67 | 51.29% | G | 0.66097
Gradko | 50% TEA in acetone | 2014 | R | London Borough of Croydon |  | 11 | 48.2612 | 46.1818 | 4.503% | P | 0.95691
Gradko | 50% TEA in acetone | 2014 | R | London Borough of Richmond upon Thames |  | 10 | 38.6464 | 36.4 | 6.171% | G | 0.94187
Gradko | 50% TEA in acetone | 2014 | R | London Borough of Richmond upon Thames |  | 12 | 48.0001 | 41.6667 | 15.2% | G | 0.86805
Gradko | 50% TEA in acetone | 2014 | B | London Borough of Richmond upon Thames |  | 11 | 23.6612 | 24.6364 | -3.958% | G | 1.0412
Gradko | 20% TEA in water | 2014 | UB | Luton Borough Council |  | 9 | 35.6585 | 37.1333 | -3.972% | G | 1.0414
Staffordshire Scientific Services | 20% TEA in water | 2014 | KS | Manchester City Council |  | 11 | 76.2455 | 68.5973 | 11.15% | G | 0.89969
Staffordshire Scientific Services | 20% TEA in water | 2014 | SU | Manchester City Council |  | 12 | 29.1611 | 21.573 | 35.17% | G | 0.73979
Staffordshire Scientific Services | 20% TEA in water | 2014 | UC | Manchester City Council |  | 12 | 43.3083 | 40.332 | 7.38% | G | 0.93127
Lambeth Scientific Services | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 99.9028 | 80.2505 | 24.49% | G | 0.80329
ESG Glasgow | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 115.1883 | 80.2505 | 43.54% | G | 0.69669
ESG Glasgow | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 105.9239 | 80.2505 | 31.99% | P | 0.75762
ESG Didcot | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 108.5076 | 80.2505 | 35.21% | P | 0.73958
ESG Didcot | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 114.4506 | 80.2505 | 42.62% | G | 0.70118
Milton Keynes Council | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 107.1458 | 80.2505 | 33.51% | G | 0.74898
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 119.7508 | 80.2505 | 49.22% | G | 0.67015
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 103.175 | 80.2505 | 28.57% | G | 0.77781
Northampton BC | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 118.677 | 80.2505 | 47.88% | G | 0.67621
Glasgow Scientific Services | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 101.4528 | 80.2505 | 26.42% | G | 0.79101
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 126.7778 | 80.2505 | 57.98% | G | 0.633
Gradko | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 114.5717 | 80.2505 | 42.77% | G | 0.70044
Gradko | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 92.1308 | 80.2505 | 14.8% | G | 0.87105
Aberdeen Scientific Services | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 106.6828 | 80.2505 | 32.94% | G | 0.75223
Staffordshire Scientific Services | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 105.9397 | 80.2505 | 32.01% | G | 0.75751
Kirklees Council | 50% TEA in acetone | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 108.0775 | 80.2505 | 34.68% | G | 0.74253
Tayside Scientific Services | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 112.6694 | 80.2505 | 40.4% | G | 0.71226
Somerset County Council | 20% TEA in water | 2014 | KS | Marylebone Road Intercomparison |  | 12 | 116.7322 | 80.2505 | 45.46% | G | 0.68747
Gradko | 20% TEA in water | 2014 | R | Monmouthshire County Council |  | 10 | 42.2787 | 38.3959 | 10.11% | G | 0.90816
ESG Didcot | 50% TEA in acetone | 2014 | R | North East Lincolnshire Council |  | 11 | 33.903 | 32.6636 | 3.794% | G | 0.96344
ESG Didcot | 50% TEA in acetone | 2014 | R | North East Lincolnshire Council |  | 12 | 36.8472 | 33.6 | 9.664% | P | 0.91187
ESG Didcot | 50% TEA in acetone | 2014 | R | North East Lincolnshire Council |  | 12 | 57.5889 | 47.1667 | 22.1% | G | 0.81902
Northampton BC | 20% TEA in water | 2014 | UB | Northampton BC |  | 12 | 17.2662 | 14.1 | 22.45% | G | 0.81663
Northampton BC | 20% TEA in water | 2014 | R | Northampton BC |  | 12 | 41.5263 | 33.2667 | 24.83% | G | 0.8011
Gradko | 50% TEA in acetone | 2014 | UB | Norwich City Council |  | 12 | 12.9083 | 13.8533 | -6.821% | G | 1.0732
Gradko | 20% TEA in water | 2014 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 44.3531 | 38.5917 | 14.93% | G | 0.8701
ESG Didcot | 50% TEA in acetone | 2014 | B | Pembrokeshire Council |  | 11 | 6.9 | 3.2727 | 110.8% | P | 0.47431
Gradko | 50% TEA in acetone | 2014 | R | Reading Borough Council |  | 11 | 41.9978 | 40.5455 | 3.582% | G | 0.96542
ESG Didcot | 20% TEA in water | 2014 | R | Rhondda Cynon Taf CBC |  | 11 | 33.6636 | 30.4618 | 10.51% | G | 0.90489
Somerset County Council | 20% TEA in water | 2014 | R | South Gloucestershire Council |  | 12 | 27.47 | 23.525 | 16.77% | G | 0.85639
ESG Didcot | 20% TEA in water | 2014 | KS | South Lakeland District Council |  | 9 | 41.4185 | 32.0478 | 29.24% | G | 0.77375
ESG Didcot | 50% TEA in acetone | 2014 | KS | South Northamptonshire Council |  | 11 | 42.9879 | 31.4964 | 36.49% | G | 0.73268
Staffordshire Scientific Services | 20% TEA in water | 2014 | O | South Staffordshire District Council |  | 10 | 37.1633 | 42.3 | -12.14% | G | 1.1382
ESG Didcot | 50% TEA in acetone | 2014 | UI | Stockton on Tees |  | 11 | 25.3515 | 21.5455 | 17.67% | P | 0.84987
ESG Didcot | 50% TEA in acetone | 2014 | R | Stockton on Tees |  | 12 | 21.075 | 15.5917 | 35.17% | G | 0.73982
Staffordshire Scientific Services | 20% TEA in water | 2014 | R | Stoke-on-Trent City Council |  | 10 | 58.6633 | 46.9115 | 25.05% | G | 0.79967
Staffordshire Scientific Services | 20% TEA in water | 2014 | UC | Stoke-on-Trent City Council |  | 10 | 31.1517 | 26.652 | 16.88% | G | 0.85556
ESG Didcot | 50% TEA in acetone | 2014 | R | Swale Borough Council |  | 9 | 41.95 | 32.6667 | 28.42% | P | 0.7787
ESG Didcot | 50% TEA in acetone | 2014 | R | Swale Borough Council |  | 12 | 50.2722 | 38.1667 | 31.72% | P | 0.7592
ESG Didcot | 50% TEA in acetone | 2014 | SU | Thanet District Council |  | 12 | 18.5292 | 17 | 8.995% | P | 0.91747
ESG Didcot | 50% TEA in acetone | 2014 | R | Thanet District Council |  | 12 | 28.4333 | 26.8333 | 5.963% | P | 0.94373
Staffordshire Scientific Services | 20% TEA in water | 2014 | R | Trafford Council |  | 12 | 38.3167 | 31.6667 | 21% | S | 0.82645
Staffordshire Scientific Services | 20% TEA in water | 2014 | UB | Trafford Council |  | 9 | 28.1333 | 21.5556 | 30.52% | S | 0.76619
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | R | Wakefield MDC |  | 12 | 37.4583 | 30.4008 | 23.21% | G | 0.81159
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | KS | Wakefield MDC |  | 12 | 53.1667 | 45.9673 | 15.66% | G | 0.86459
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | UB | Wakefield MDC |  | 12 | 39.2917 | 28.2486 | 39.09% | G | 0.71895
ESG Didcot | 20% TEA in water | 2014 | UB | Wigan Council |  | 13 | 28.0154 | 21.9685 | 27.53% | P | 0.78416
Gradko | 50% TEA in acetone | 2014 | R | Worthing Borough Council |  | 12 | 43.2222 | 50.9713 | -15.2% | G | 1.1793
ESG Didcot | 50% TEA in acetone | 2014 | R | Wrexham County Borough Council |  | 10 | 23.0067 | 21.78 | 5.632% | G | 0.94668
ESG Didcot | 50% TEA in acetone | 2014 | UB | City of York Council |  | 11 | 24.2107 | 18.8627 | 28.35% | P | 0.77911
ESG Didcot | 50% TEA in acetone | 2014 | R | City of York Council |  | 10 | 37.1049 | 27.1392 | 36.72% | G | 0.73142
ESG Didcot | 50% TEA in acetone | 2014 | R | City of York Council |  | 11 | 31.8102 | 28.305 | 12.38% | G | 0.88981
ESG Didcot | 50% TEA in acetone | 2014 | R | City of York Council |  | 11 | 40.411 | 35.8678 | 12.67% | G | 0.88758
Staffordshire Scientific Services | 20% TEA in water | 2014 | UB | Tameside MBC |  | 12 | 20.7375 | 16.5 | 25.68% | G | 0.79566
Staffordshire Scientific Services | 20% TEA in water | 2014 | UI | Salford City Council |  | 12 | 34.0028 | 28.5 | 19.31% | G | 0.83817
Staffordshire Scientific Services | 20% TEA in water | 2014 | R | Salford City Council |  | 12 | 57.7556 | 59.5 | -2.932% | G | 1.0302
Aberdeen Scientific Services | 20% TEA in water | 2014 | UB | Aberdeen City Council |  | 12 | 27.7778 | 21.32 | 30.29% | G | 0.76752
Aberdeen Scientific Services | 20% TEA in water | 2014 | R | Aberdeen City Council |  | 13 | 51.2821 | 39.7469 | 29.02% | G | 0.77507
Aberdeen Scientific Services | 20% TEA in water | 2014 | R | Aberdeen City Council |  | 10 | 53.1667 | 47.402 | 12.16% | G | 0.89157
Aberdeen Scientific Services | 20% TEA in water | 2014 | R | Aberdeen City Council |  | 11 | 18.4242 | 25.4 | -27.46% | G | 1.3786
Aberdeen Scientific Services | 20% TEA in water | 2014 | R | Aberdeen City Council |  | 13 | 32.8205 | 26.81 | 22.42% | G | 0.81687
Aberdeen Scientific Services | 20% TEA in water | 2014 | R | Aberdeen City Council |  | 12 | 54.5833 | 45.4533 | 20.09% | G | 0.83273
Gradko | 20% TEA in water | 2014 | R | Bedford Borough Council |  | 12 | 38.4444 | 39.4917 | -2.652% | G | 1.0272
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | R | City of Edinburgh Council |  | 12 | 39.425 | 26.4167 | 49.24% | G | 0.67005
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | R | City Of Edinburgh Council |  | 12 | 39.4528 | 33.8333 | 16.61% | G | 0.85757
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | R | City Of Edinburgh Council |  | 11 | 39.2333 | 25.0909 | 56.36% | G | 0.63953
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | R | City of Edinburgh Council |  | 11 | 66.0879 | 45.9091 | 43.95% | G | 0.69467
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | R | City Of Edinburgh Council |  | 12 | 34.5403 | 26.5833 | 29.93% | G | 0.76963
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | KS | City Of Edinburgh Council |  | 11 | 71.4848 | 58.3636 | 22.48% | G | 0.81645
Gradko | 20% TEA in water | 2014 | R | City of Lincoln Council |  | 12 | 44.5346 | 38.1142 | 16.85% | G | 0.85583
Tayside Scientific Services | 20% TEA in water | 2014 | R | Dundee City Council |  | 12 | 52.5556 | 45.5417 | 15.4% | G | 0.86654
Tayside Scientific Services | 20% TEA in water | 2014 | R | Dundee City Council |  | 12 | 53.2972 | 39.7083 | 34.22% | G | 0.74504
Tayside Scientific Services | 20% TEA in water | 2014 | R | Dundee City Council |  | 12 | 55.4417 | 54.5167 | 1.697% | G | 0.98332
Tayside Scientific Services | 20% TEA in water | 2014 | R | Dundee City Council |  | 12 | 39.0306 | 28.35 | 37.67% | G | 0.72635
Tayside Scientific Services | 20% TEA in water | 2014 | R | Dundee City Council |  | 12 | 51.4167 | 42.4833 | 21.03% | G | 0.82626
Gradko | 20% TEA in water | 2014 | R | East Herts Council |  | 11 | 37.4242 | 32.6818 | 14.51% | G | 0.87328
ESG Didcot | 50% TEA in acetone | 2014 | R | Horsham District Council |  | 12 | 29.6583 | 26.5583 | 11.67% | G | 0.89548
Gradko | 20% TEA in water | 2014 | R | Lancaster City Council |  | 11 | 36.1027 | 37.5886 | -3.953% | G | 1.0412
Lambeth Scientific Services | 50% TEA in acetone | 2014 | R | LB Haringey |  | 11 | 45.5455 | 44.7182 | 1.85% | S | 0.98184
ESG Didcot | 50% TEA in acetone | 2014 | R | Medway Council |  | 12 | 32.2944 | 25.6229 | 26.04% | G | 0.79342
ESG Didcot | 50% TEA in acetone | 2014 | B | Medway Council |  | 12 | 20.8611 | 14.555 | 43.33% | P | 0.69771
ESG Didcot | 50% TEA in acetone | 2014 | R | Sandwell MBC |  | 10 | 54.8467 | 46.694 | 17.46% | G | 0.85136
Staffordshire Scientific Services | 20% TEA in water | 2014 | R | Stockport |  | 12 | 33.8861 | 26.8333 | 26.28% | G | 0.79187
ESG Didcot | 50% TEA in acetone | 2014 | KS | Suffolk Coastal District Council |  | 10 | 45.5633 | 38.7 | 17.73% | G | 0.84937
ESG Didcot | 50% TEA in acetone | 2014 | R | Vale of White Horse District Council |  | 12 | 34.1667 | 29.9833 | 13.95% | G | 0.87756
Gradko | 50% TEA in acetone | 2014 | R | West Berkshire Council |  | 12 | 42.4025 | 46.0292 | -7.879% | G | 1.0855
ESG Didcot | 50% TEA in acetone | 2014 | R | West Oxfordshire District Council (WODC) |  | 12 | 48.2222 | 36.4167 | 32.42% | G | 0.75518
Gradko | 20% TEA in water | 2014 | R | Wokingham Borough Council |  | 12 | 40.4897 | 37.0467 | 9.294% | G | 0.91496
Gradko | 20% TEA in water | 2014 | UC | Southampton City Council |  | 11 | 32.1424 | 31.0636 | 3.473% | G | 0.96644
ESG Didcot | 50% TEA in acetone | 2014 | R | Dumfries and Galloway Council |  | 12 | 35.3861 | 30.3141 | 16.73% | G | 0.85667
Lambeth Scientific Services | 50% TEA in acetone | 2014 | SU | Reigate and Banstead |  | 10 | 27.0167 | 20.36 | 32.69% | P | 0.75361
Lambeth Scientific Services | 50% TEA in acetone | 2014 | R | Reigate and Banstead |  | 12 | 31.4722 | 28.6917 | 9.691% | G | 0.91165
Lambeth Scientific Services | 50% TEA in acetone | 2014 | B | Reigate and Banstead |  | 12 | 18.3611 | 17.4917 | 4.971% | P | 0.95265
ESG Didcot | 50% TEA in acetone | 2014 | R | Hambleton District Council |  | 12 | 23.7583 | 19.7333 | 20.4% | G | 0.83059
Gradko | 20% TEA in water | 2015 | R | Ards and North Down Borough Council |  | 12 | 38.3817 | 25.825 | 48.62% | G | 0.67285
Gradko | 50% TEA in acetone | 2015 | R | Bedford Borough Council |  | 12 | 35.4167 | 33.275 | 6.436% | G | 0.93953
Gradko | 20% TEA in water | 2015 | UC | Breckland Council |  | 12 | 29.6406 | 29.2167 | 1.451% | G | 0.9857
Gradko | 20% TEA in water | 2015 | R | Cheltenham Borough Council |  | 12 | 35.4389 | 34.5167 | 2.672% | G | 0.97398
ESG Didcot | 50% TEA in acetone | 2015 | R | Dumfries and Galloway Council |  | 12 | 34.7028 | 30.2695 | 14.65% | G | 0.87225
Glasgow Scientific Services | 20% TEA in water | 2015 | R | East Dunbartonshire Council |  | 12 | 33.5306 | 34.1667 | -1.862% | G | 1.019
Glasgow Scientific Services | 20% TEA in water | 2015 | R | East Dunbartonshire Council |  | 12 | 28.7389 | 27.5833 | 4.189% | P | 0.95979
Glasgow Scientific Services | 20% TEA in water | 2015 | R | East Dunbartonshire Council |  | 11 | 29.7 | 28.6364 | 3.714% | P | 0.96419
Glasgow Scientific Services | 20% TEA in water | 2015 | R | East Dunbartonshire Council |  | 11 | 23.0636 | 23.0909 | -0.1181% | P | 1.0012
ESG Didcot | 50% TEA in acetone | 2015 | B | Gravesham Borough Council |  | 12 | 40.1319 | 29.9167 | 34.15% | G | 0.74546
ESG Didcot | 50% TEA in acetone | 2015 | B | Gravesham Borough Council |  | 12 | 30.3861 | 23.4167 | 29.76% | P | 0.77064
Gradko | 20% TEA in water | 2015 | R | Lisburn & Castlereagh City Council |  | 10 | 35.6869 | 28.6 | 24.78% | G | 0.80141
Gradko | 20% TEA in water | 2015 | R | Luton Borough Council |  | 12 | 46.3411 | 43.7167 | 6.003% | G | 0.94337
Staffordshire Scientific Services | 20% TEA in water | 2015 | UC | Manchester City Council |  | 12 | 42.2778 | 39.5404 | 6.923% | G | 0.93525
Staffordshire Scientific Services | 20% TEA in water | 2015 | SU | Manchester City Council |  | 12 | 24.55 | 19.717 | 24.51% | G | 0.80314
Staffordshire Scientific Services | 20% TEA in water | 2015 | KS | Manchester City Council |  | 11 | 72.4848 | 66.2962 | 9.335% | G | 0.91462
Gradko | 20% TEA in water | 2015 | R | Monmouthshire County Council |  | 12 | 40.6778 | 36.6455 | 11% | G | 0.90087
ESG Didcot | 50% TEA in acetone | 2015 | UI | North Lincolnshire |  | 11 | 23.9576 | 17.5455 | 36.55% | P | 0.73236
Gradko | 50% TEA in acetone | 2015 | UB | Norwich City Council |  | 9 | 11.5846 | 11.9778 | -3.282% | G | 1.0339
Gradko | 20% TEA in water | 2015 | B | Pembrokeshire Council |  | 10 | 3.6917 | 2.7 | 36.73% | G | 0.73138
ESG Didcot | 50% TEA in acetone | 2015 | R | Swale Borough Council |  | 12 | 40.5361 | 32.6667 | 24.09% | P | 0.80587
ESG Didcot | 50% TEA in acetone | 2015 | R | Swale Borough Council |  | 10 | 47.5633 | 38.3 | 24.19% | G | 0.80524
ESG Didcot | 50% TEA in acetone | 2015 | R | Swale Borough Council |  | 11 | 38.2848 | 29.8182 | 28.39% | P | 0.77885
Staffordshire Scientific Services | 20% TEA in water | 2015 | R | Trafford |  | 9 | 35.1389 | 29.8889 | 17.57% | G | 0.85059
Staffordshire Scientific Services | 20% TEA in water | 2015 | UB | Trafford |  | 12 | 22.9917 | 20.5 | 12.15% | G | 0.89163
Gradko | 50% TEA in acetone | 2015 | R | West Berkshire Council |  | 11 | 38.443 | 34.7318 | 10.69% | G | 0.90346
Staffordshire Scientific Services | 20% TEA in water | 2015 | UB | Wigan Council |  | 11 | 26.5364 | 20 | 32.68% | G | 0.75368
ESG Didcot | 50% TEA in acetone | 2015 | R | Wrexham County Borough Council |  | 12 | 19.4806 | 19.3667 | 0.5881% | G | 0.99415
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | R | West Lothian Council |  | 12 | 36.7361 | 28.3333 | 29.66% | G | 0.77127
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | R | West Lothian Council |  | 11 | 38.6364 | 32.6364 | 18.38% | G | 0.84471
Gradko | 20% TEA in water | 2015 | R | City of Lincoln Council |  | 12 | 39.1578 | 33.2017 | 17.94% | G | 0.84789
Somerset County Council | 20% TEA in water | 2015 | R | Bath & North East Somerset |  | 12 | 50.9553 | 53.8418 | -5.361% | G | 1.0566
Gradko | 20% TEA in water | 2015 | R | Borough Council of King's Lynn and West Norfolk |  | 12 | 28.5278 | 21.5333 | 32.48% | G | 0.75482
Gradko | 20% TEA in water | 2015 | R | Cheshire West and Chester |  | 10 | 37.6531 | 39.73 | -5.228% | G | 1.0552
Gradko | 20% TEA in water | 2015 | R | Dudley MBC |  | 12 | 47.1936 | 50.1458 | -5.887% | G | 1.0626
Gradko | 20% TEA in water | 2015 | R | Dudley MBC |  | 12 | 40.0269 | 35.1225 | 13.96% | G | 0.87747
Gradko | 20% TEA in water | 2015 | R | Dudley MBC |  | 12 | 34.0128 | 30.9167 | 10.01% | G | 0.90897
Gradko | 20% TEA in water | 2015 | UB | Dudley MBC |  | 11 | 23.387 | 19.34 | 20.93% | G | 0.82696
Gradko | 50% TEA in acetone | 2015 | R | East Hampshire District Council |  | 11 | 22.0092 | 20.1 | 9.499% | G | 0.91325
Glasgow Scientific Services | 20% TEA in water | 2015 | KS | Glasgow City Council |  | 12 | 60.2528 | 60.7932 | -0.889% | P | 1.009
Glasgow Scientific Services | 20% TEA in water | 2015 | UB | Glasgow City Council |  | 10 | 25.3533 | 24.55 | 3.272% | P | 0.96831
Glasgow Scientific Services | 20% TEA in water | 2015 | R | Glasgow City Council |  | 9 | 30.2222 | 31.0889 | -2.788% | P | 1.0287
Glasgow Scientific Services | 20% TEA in water | 2015 | R | Glasgow City Council |  | 12 | 43.4556 | 38.1083 | 14.03% | P | 0.87695
Lambeth Scientific Services | 50% TEA in acetone | 2015 | R | LB Haringey |  | 12 | 36.5833 | 40.25 | -9.11% | S | 1.1002
Gradko | 50% TEA in acetone | 2015 | KS | London Borough of Croydon |  | 12 | 54.455 | 52.0083 | 4.704% | G | 0.95507
Gradko | 50% TEA in acetone | 2015 | B | London Borough of Richmond upon Thames |  | 12 | 21.214 | 21.25 | -0.1696% | G | 1.0017
Gradko | 50% TEA in acetone | 2015 | R | London Borough of Richmond upon Thames |  | 12 | 36.2083 | 33.25 | 8.897% | G | 0.9183
ESG Glasgow | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 104.5901 | 81.0404 | 29.06% | G | 0.77484
ESG Didcot | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 103.6181 | 81.0404 | 27.86% | G | 0.78211
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 111.4974 | 81.0404 | 37.58% | G | 0.72684
Lambeth Scientific Services | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 96.4 | 81.0404 | 18.95% | G | 0.84067
Milton Keynes Council | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 106.4778 | 81.0404 | 31.39% | G | 0.7611
Northampton BC | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 123.6328 | 81.0404 | 52.56% | G | 0.65549
ESG Didcot | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 108.4066 | 81.0404 | 33.77% | G | 0.74756
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 98.3278 | 81.0404 | 21.33% | G | 0.82419
Glasgow Scientific Services | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 98.1333 | 81.0404 | 21.09% | G | 0.82582
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 115.8333 | 81.0404 | 42.93% | G | 0.69963
Gradko | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 102.2429 | 81.0404 | 26.16% | G | 0.79263
Gradko | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 86.2271 | 81.0404 | 6.4% | G | 0.93985
Aberdeen Scientific Services | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 99.5222 | 81.0404 | 22.81% | G | 0.81429
Staffordshire Scientific Services | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 11 | 95.4517 | 80.8997 | 17.99% | G | 0.84755
Kirklees Council | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 108.4582 | 81.0404 | 33.83% | G | 0.7472
Tayside Scientific Services | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 106.5417 | 81.0404 | 31.47% | G | 0.76064
Somerset County Council | 20% TEA in water | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 106.6903 | 81.0404 | 31.65% | G | 0.75959
Gradko | 50% TEA in acetone | 2015 | UI | Middlesbrough |  | 11 | 16.1485 | 15.2182 | 6.113% | G | 0.94239
Gradko | 50% TEA in acetone | 2015 | SI | Redcar & Cleveland |  | 12 | 12.4861 | 12.475 | 0.08907% | G | 0.99911
Staffordshire Scientific Services | 20% TEA in water | 2015 | R | stockport SMBC |  | 12 | 30.8306 | 24.25 | 27.14% | G | 0.78656
Staffordshire Scientific Services | 20% TEA in water | 2015 | R | Stoke-on-Trent City Council |  | 12 | 59.4694 | 47.3428 | 25.61% | G | 0.79609
Staffordshire Scientific Services | 20% TEA in water | 2015 | R | Stoke-on-Trent City Council |  | 10 | 37.4 | 28.7885 | 29.91% | G | 0.76975
Staffordshire Scientific Services | 20% TEA in water | 2015 | UC | Stoke-on-Trent City Council |  | 12 | 31.6694 | 27.8502 | 13.71% | G | 0.8794
Lambeth Scientific Services | 50% TEA in acetone | 2015 | R | Surrey Heath Borough Council |  | 12 | 29.5417 | 40.125 | -26.38% | G | 1.3583
ESG Didcot | 50% TEA in acetone | 2015 | R | Vale of White Horse District Council |  | 11 | 33.9212 | 29.3182 | 15.7% | G | 0.8643
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 | R | Barnsley MBC |  | 12 | 40.2004 | 38.4167 | 4.643% | G | 0.95563
Staffordshire Scientific Services | 20% TEA in water | 2015 | R | East Staffordshire Borough Council |  | 10 | 50.73 | 39.27 | 29.18% | G | 0.7741
Tayside Scientific Services | 20% TEA in water | 2015 | R | Fife Council |  | 11 | 30.0909 | 22.7273 | 32.4% | G | 0.75529
Tayside Scientific Services | 20% TEA in water | 2015 | R | Fife Council |  | 12 | 25.4028 | 18.9167 | 34.29% | G | 0.74467
Tayside Scientific Services | 20% TEA in water | 2015 | R | Fife Council |  | 12 | 32.8861 | 25.3333 | 29.81% | G | 0.77034
Tayside Scientific Services | 20% TEA in water | 2015 | R | Fife Council |  | 12 | 35.2139 | 28.0833 | 25.39% | G | 0.79751
Gradko | 20% TEA in water | 2015 | UB | Liverpool |  | 12 | 20.2953 | 22.3 | -8.99% | G | 1.0988
ESG Glasgow | 50% TEA in acetone | 2015 | KS | Marylebone Road Intercomparison |  | 12 | 103.9373 | 81.0404 | 28.25% | G | 0.7797
Glasgow Scientific Services | 20% TEA in water | 2015 | R | North Ayrshire Council |  | 12 | 28.7639 | 28.25 | 1.819% | P | 0.98213
Staffordshire Scientific Services | 20% TEA in water | 2015 | R | Oldham MBC |  | 9 | 41.6833 | 32.5667 | 27.99% | G | 0.78129
Gradko | 20% TEA in water | 2015 | R | Preston City Council |  | 12 | 29.2831 | 26.8833 | 8.926% | G | 0.91805
Somerset County Council | 20% TEA in water | 2015 | R | South Gloucestershire Council |  | 12 | 26.292 | 21.5083 | 22.24% | G | 0.81806
ESG Didcot | 20% TEA in water | 2015 | KS | South Lakeland District Council |  | 11 | 42.4455 | 30.9982 | 36.93% | G | 0.73031
ESG Didcot | 50% TEA in acetone | 2015 | UI | Stockton on Tees |  | 12 | 23.9222 | 19.0667 | 25.47% | G | 0.79703
ESG Didcot | 50% TEA in acetone | 2015 | R | Stockton on Tees |  | 12 | 17.2889 | 14.475 | 19.44% | G | 0.83724
ESG Didcot | 50% TEA in acetone | 2015 | KS | Suffolk Coastal DC |  | 12 | 43.8917 | 34.8333 | 26% | P | 0.79362
Gradko | 20% TEA in water | 2015 | R | Thurrock Borough Council |  | 12 | 28.2366 | 23.0443 | 22.53% | G | 0.81611
Gradko | 20% TEA in water | 2015 | R | Gateshead Council |  | 11 | 33.1879 | 33.5818 | -1.173% | G | 1.0119
Gradko | 20% TEA in water | 2015 | R | Gateshead Council |  | 12 | 27.7167 | 26.675 | 3.905% | G | 0.96242
Gradko | 20% TEA in water | 2015 | R | Gateshead Council |  | 10 | 36.2303 | 32.48 | 11.55% | G | 0.89649
Kirklees Council | 50% TEA in acetone | 2015 | R | Kirklees Council |  | 12 | 52.0361 | 40.9025 | 27.22% | G | 0.78604
Kirklees Council | 50% TEA in acetone | 2015 | R | Kirklees Council |  | 12 | 50.9225 | 38.4167 | 32.55% | S | 0.75441
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | UC | Leeds City Council |  | 12 | 44.6944 | 30.975 | 44.29% | G | 0.69304
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | R | Leeds City Council |  | 9 | 48.7778 | 41.7333 | 16.88% | S | 0.85558
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | R | Leeds City Council |  | 12 | 45.5833 | 38.4167 | 18.66% | S | 0.84278
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | R | Leeds City Council |  | 9 | 39.5 | 28.4222 | 38.98% | P | 0.71955
Gradko | 20% TEA in water | 2015 | KS | New Forest DC |  | 11 | 46.7944 | 35.6909 | 31.11% | P | 0.76272
Gradko | 20% TEA in water | 2015 | R | New Forest DC |  | 11 | 33.2491 | 25.2455 | 31.7% | G | 0.75928
ESG Didcot | 20% TEA in water | 2015 | R | Rhondda Cynon Taf CBC |  | 12 | 30.3431 | 29.1958 | 3.929% | G | 0.96219
ESG Didcot | 50% TEA in acetone | 2015 | SU | Thanet District Council |  | 9 | 16.8407 | 15.2222 | 10.63% | G | 0.90389
ESG Didcot | 50% TEA in acetone | 2015 | R | Thanet District Council |  | 12 | 26.9056 | 22.8333 | 17.83% | G | 0.84865
Gradko | 50% TEA in acetone | 2015 | R | West Dorset District Council |  | 12 | 12.3714 | 10.7067 | 15.55% | G | 0.86544
Gradko | 50% TEA in acetone | 2015 | R | Worthing Borough Council |  | 11 | 41.8 | 36.5182 | 14.46% | G | 0.87364
ESG Didcot | 50% TEA in acetone | 2015 | B | Medway Council |  | 12 | 21.2806 | 12 | 77.34% | G | 0.5639
ESG Didcot | 50% TEA in acetone | 2015 | R | Medway Council |  | 11 | 32.1424 | 22.5455 | 42.57% | G | 0.70142
Staffordshire Scientific Services | 20% TEA in water | 2015 | R | Salford Council |  | 10 | 53.1333 | 53.32 | -0.3501% | G | 1.0035
Staffordshire Scientific Services | 20% TEA in water | 2015 | UB | Salford Council |  | 12 | 31.5611 | 27.1917 | 16.07% | G | 0.86156
Gradko | 20% TEA in water | 2015 | UC | Southampton City Council |  | 12 | 28.3569 | 29.3917 | -3.52% | G | 1.0365
Gradko | 20% TEA in water | 2015 | R | Wokingham Borough Council |  | 11 | 35.5202 | 32.9091 | 7.934% | G | 0.92649
Gradko | 50% TEA in acetone | 2015 | R | Royal Borough of Windsor and Maidenhead |  | 12 | 33.6771 | 36.775 | -8.424% | G | 1.092
Gradko | 50% TEA in acetone | 2015 | R | Royal Borough of Windsor and Maidenhead |  | 12 | 39.5939 | 38 | 4.194% | G | 0.95974
ESG Didcot | 50% TEA in acetone | 2015 | R | North East Lincolnshire Council |  | 10 | 34.27 | 28.28 | 21.18% | P | 0.82521
ESG Didcot | 50% TEA in acetone | 2015 | R | North East Lincolnshire Council |  | 11 | 38.5515 | 27.8091 | 38.63% | G | 0.72135
ESG Didcot | 50% TEA in acetone | 2015 | R | North East Lincolnshire Council |  | 11 | 54.753 | 47.1364 | 16.16% | G | 0.86089
Gradko | 20% TEA in water | 2015 | R | Brighton & Hove City Council |  | 9 | 46.5667 | 37.5379 | 24.05% | G | 0.80611
ESG Didcot | 50% TEA in acetone | 2015 | R | Hambleton District Council |  | 10 | 21.8733 | 18.6 | 17.6% | G | 0.85035
Somerset County Council | 20% TEA in water | 2015 | UB | Bristol City Council |  | 9 | 34.2619 | 31.4444 | 8.96% | G | 0.91777
Somerset County Council | 20% TEA in water | 2015 | R | Bristol City Council |  | 11 | 36.6212 | 41.2727 | -11.27% | G | 1.127
Somerset County Council | 20% TEA in water | 2015 | R | Bristol City Council |  | 10 | 50.44 | 40.8 | 23.63% | G | 0.80888
Somerset County Council | 20% TEA in water | 2015 | R | Bristol City Council |  | 12 | 44.0306 | 44.25 | -0.4959% | G | 1.005
Somerset County Council | 20% TEA in water | 2015 | UC | Bristol City Council |  | 11 | 97.6333 | 85.8182 | 13.77% | G | 0.87898
Somerset County Council | 20% TEA in water | 2015 | KS | Bristol City Council |  | 12 | 46.2722 | 39.4167 | 17.39% | G | 0.85184
Somerset County Council | 20% TEA in water | 2015 | UB | Bristol City Council |  | 11 | 29.203 | 26.0818 | 11.97% | G | 0.89312
ESG Didcot | 50% TEA in acetone | 2015 | UB | City of York Council |  | 11 | 23.8705 | 15.8509 | 50.59% | G | 0.66404
ESG Didcot | 50% TEA in acetone | 2015 | R | City of York Council |  | 11 | 35.9058 | 27.2261 | 31.88% | G | 0.75826
ESG Didcot | 50% TEA in acetone | 2015 | R | City of York Council |  | 11 | 33.9702 | 25.201 | 34.8% | G | 0.74186
ESG Didcot | 50% TEA in acetone | 2015 | R | City of York Council |  | 12 | 39.2165 | 27.7847 | 41.14% | G | 0.7085
Staffordshire Scientific Services | 20% TEA in water | 2015 | O | South Staffordshire Council |  | 10 | 32.8833 | 33.465 | -1.738% | G | 1.0177
Aberdeen Scientific Services | 20% TEA in water | 2015 | R | Aberdeen City Council |  | 13 | 54.2564 | 46.7508 | 16.05% | G | 0.86166
Aberdeen Scientific Services | 20% TEA in water | 2015 | R | Aberdeen City Council |  | 13 | 43.3333 | 40.8623 | 6.047% | G | 0.94298
Aberdeen Scientific Services | 20% TEA in water | 2015 | UB | Aberdeen City Council |  | 13 | 29.1026 | 22.9169 | 26.99% | G | 0.78745
Aberdeen Scientific Services | 20% TEA in water | 2015 | R | Aberdeen City Council |  | 13 | 44.5385 | 36.2077 | 23.01% | G | 0.81295
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | R | City of Edinburgh |  | 12 | 35.7153 | 28.46 | 25.49% | G | 0.79686
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | KS | City of Edinburgh |  | 10 | 65.65 | 61.518 | 6.717% | G | 0.93706
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | R | City of Edinburgh |  | 10 | 61.92 | 41.048 | 50.85% | G | 0.66292
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | R | City of Edinburgh |  | 12 | 39.2083 | 27.3892 | 43.15% | G | 0.69855
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | R | City of Edinburgh |  | 12 | 43.3444 | 25.8475 | 67.69% | G | 0.59633
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | R | City of Edinburgh |  | 10 | 36.7583 | 32.204 | 14.14% | P | 0.8761
Gradko | 50% TEA in acetone | 2015 | R | LB Newham |  | 11 | 34.5656 | 38.3636 | -9.9% | G | 1.1099
Aberdeen Scientific Services | 20% TEA in water | 2015 | R | Aberdeen City Council |  | 13 | 33.0769 | 27.8523 | 18.76% | G | 0.84205
ESG Didcot | 50% TEA in acetone | 2015 | R | Rugby Borough Council |  | 12 | 22.8653 | 20.6667 | 10.64% | G | 0.90384
Glasgow Scientific Services | 20% TEA in water | 2015 | R | North Lanarkshire Council |  | 12 | 15.775 | 18.575 | -15.07% | P | 1.1775
Gradko | 20% TEA in water | 2015 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 40.2261 | 38.5667 | 4.303% | G | 0.95875
Lambeth Scientific Services | 50% TEA in acetone | 2015 | R | Reigate and Banstead |  | 12 | 27.5556 | 26.325 | 4.674% | G | 0.95534
Lambeth Scientific Services | 50% TEA in acetone | 2015 | B | Reigate and Banstead |  | 12 | 14.5833 | 13.9417 | 4.603% | P | 0.956
Lambeth Scientific Services | 50% TEA in acetone | 2015 | UB | Reigate and Banstead |  | 12 | 22.5278 | 21.0167 | 7.19% | G | 0.93292
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | KS | Leeds City Council |  | 11 | 67.4545 | 54.2 | 24.45% | S | 0.8035
ESG Didcot | 50% TEA in acetone | 2015 | R | Cambridge City Council |  | 12 | 44.6375 | 35.6667 | 25.15% | G | 0.79903
Tayside Scientific Services | 20% TEA in water | 2015 | R | Dundee City Council |  | 10 | 56.37 | 48.49 | 16.25% | G | 0.86021
Tayside Scientific Services | 20% TEA in water | 2015 | R | Dundee City Council |  | 9 | 50.1 | 38.5889 | 29.83% | G | 0.77024
Tayside Scientific Services | 20% TEA in water | 2015 | R | Dundee City Council |  | 12 | 51.5528 | 50.7417 | 1.599% | G | 0.98427
Tayside Scientific Services | 20% TEA in water | 2015 | R | Dundee City Council |  | 12 | 39.5222 | 28.1583 | 40.36% | G | 0.71247
Tayside Scientific Services | 20% TEA in water | 2015 | R | Dundee City Council |  | 11 | 44.503 | 37.0909 | 19.98% | G | 0.83345
Glasgow Scientific Services | 20% TEA in water | 2015 | R | East Ayrshire Council |  | 9 | 23.3185 | 24.5556 | -5.038% | P | 1.053
Gradko | 20% TEA in water | 2015 | R | Lancaster City Council |  | 11 | 34.1055 | 35.1756 | -3.042% | G | 1.0314
Lambeth Scientific Services | 50% TEA in acetone | 2015 | UB | Spelthorne Borough Council |  | 12 | 27.4167 | 27.6102 | -0.7009% | G | 1.0071
ESG Didcot | 50% TEA in acetone | 2015 | R | Derry City and Strabane District Council |  | 12 | 37.5694 | 28.5342 | 31.66% | P | 0.7595
ESG Didcot | 50% TEA in acetone | 2015 | R | Watford Borough Council |  | 12 | 40.0417 | 35.1333 | 13.97% | S | 0.87742
Gradko | 20% TEA in water | 2015 | R | Hounslow Council |  | 12 | 70.5396 | 53.6687 | 31.44% | G | 0.76083
Gradko | 20% TEA in water | 2015 | R | Hounslow Council |  | 12 | 66.0231 | 44.8917 | 47.07% | G | 0.67994
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | R | Calderdale MBC |  | 12 | 57.3056 | 46.55 | 23.11% | G | 0.81231
ESG Didcot | 50% TEA in acetone | 2016 | UC | Cardiff City Council |  | 12 | 29.3861 | 22.4167 | 31.09% | G | 0.76283
ESG Didcot | 50% TEA in acetone | 2016 | R | Dumfries and Galloway Council |  | 12 | 35.7361 | 31.023 | 15.19% | G | 0.86811
Staffordshire Scientific Services | 20% TEA in water | 2016 | R | East Staffordshire Borough Council |  | 12 | 49.45 | 50.75 | -2.562% | G | 1.0263
Gradko | 20% TEA in water | 2016 | R | Gateshead Council |  | 12 | 28.5125 | 25.7933 | 10.54% | G | 0.90463
Gradko | 20% TEA in water | 2016 | R | Gateshead Council |  | 11 | 35.1339 | 37.3636 | -5.968% | G | 1.0635
Gradko | 20% TEA in water | 2016 | R | Gateshead Council |  | 12 | 37.1758 | 31.2442 | 18.98% | G | 0.84044
Glasgow Scientific Services | 20% TEA in water | 2016 | KS | Glasgow City Council |  | 12 | 59.9861 | 64.8667 | -7.524% | P | 1.0814
Glasgow Scientific Services | 20% TEA in water | 2016 | R | Glasgow City Council |  | 11 | 36.2545 | 33.3364 | 8.754% | P | 0.91951
Glasgow Scientific Services | 20% TEA in water | 2016 | R | Glasgow City Council |  | 12 | 39.8069 | 35.95 | 10.73% | P | 0.90311
Glasgow Scientific Services | 20% TEA in water | 2016 | UB | Glasgow City Council |  | 12 | 28.85 | 26.2667 | 9.835% | P | 0.91046
Gradko | 50% TEA in Acetone | 2016 | R | LB Newham |  | 12 | 35.6595 | 41.75 | -14.59% | G | 1.1708
Gradko | 50% TEA in acetone | 2016 | UB | London Borough of Camden |  | 12 | 42.1292 | 42.6908 | -1.316% | G | 1.0133
Gradko | 50% TEA in acetone | 2016 | R | London Borough of Richmond upon Thames |  | 12 | 36.3966 | 35.5417 | 2.405% | G | 0.97651
Gradko | 50% TEA in acetone | 2016 | B | London Borough of Richmond upon Thames |  | 11 | 23.9137 | 25.8818 | -7.604% | G | 1.0823
Gradko | 50% TEA in acetone | 2016 | R | Royal Borough of Greenwich |  | 11 | 50.6721 | 44.7364 | 13.27% | G | 0.88286
Gradko | 50% TEA in acetone | 2016 | SU | Royal Borough of Greenwich |  | 12 | 20.1747 | 21.4333 | -5.872% | G | 1.0624
Gradko | 50% TEA in acetone | 2016 | R | Royal Borough of Greenwich |  | 11 | 45.4397 | 45.0455 | 0.8752% | G | 0.99132
Gradko | 50% TEA in acetone | 2016 | R | Royal Borough of Greenwich |  | 12 | 68.8333 | 60.8833 | 13.06% | G | 0.8845
Gradko | 50% TEA in acetone | 2016 | R | Royal Borough of Greenwich |  | 9 | 40.167 | 41.2333 | -2.586% | G | 1.0265
Gradko | 50% TEA in acetone | 2016 | R | Royal Borough of Greenwich |  | 12 | 41.2756 | 38.0833 | 8.382% | P | 0.92266
ESG Didcot | 20% TEA in water | 2016 | KS | South Lakeland District Council |  | 12 | 40.8528 | 30.3583 | 34.57% | G | 0.74312
ESG Didcot | 50% TEA in acetone | 2016 | UI | Stockton on Tees |  | 10 | 23.2283 | 18.4 | 26.24% | P | 0.79214
ESG Didcot | 50% TEA in acetone | 2016 | R | Stockton on Tees |  | 11 | 19.103 | 14.8091 | 29% | G | 0.77522
ESG Didcot | 50% TEA in acetone | 2016 | R | Vale of Glamorgan |  | 10 | 40.23 | 28 | 43.68% | G | 0.696
ESG Didcot | 50% TEA in acetone | 2016 | R | Vale of White Horse District Council |  | 11 | 33.3727 | 28.9545 | 15.26% | G | 0.86761
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | UB | Wakefield Council |  | 11 | 39.6061 | 27.6829 | 43.07% | G | 0.69896
Gradko | 50% TEA in acetone | 2016 | R | West Berkshire Council |  | 12 | 37.8197 | 41.5026 | -8.874% | G | 1.0974
Staffordshire Scientific Services | 20% TEA in water | 2016 | UB | Wigan Council |  | 10 | 26.3867 | 20.51 | 28.65% | G | 0.77729
Gradko | 20% TEA in water | 2016 | R | Wokingham Borough Council |  | 11 | 44.6198 | 40.9174 | 9.049% | G | 0.91702
Gradko | 20% TEA in water | 2016 | R | Wokingham Borough Council |  | 11 | 36.7023 | 33.5245 | 9.479% | G | 0.91342
Gradko | 20% TEA in water | 2016 | R | Cheshire West and Chester |  | 12 | 37.2965 | 39.3917 | -5.319% | G | 1.0562
Gradko | 50% TEA in acetone | 2016 | R | East Hampshire District Council |  | 12 | 21.2197 | 22.625 | -6.211% | G | 1.0662
Gradko | 20% TEA in water | 2016 | R | Thurrock Borough Council |  | 12 | 29.3111 | 26.4119 | 10.98% | G | 0.90109
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2016 | R | Barnsley MBC |  | 12 | 37.9722 | 35.75 | 6.216% | G | 0.94148
Gradko | 20% TEA in water | 2016 | R | Borough Council of King's Lynn & West Norfolk |  | 11 | 29.5303 | 24.9909 | 18.16% | G | 0.84628
Gradko | 50% TEA in acetone | 2016 | B | City of London |  | 12 | 38.0107 | 41.5669 | -8.555% | G | 1.0936
Gradko | 50% TEA in acetone | 2016 | R | City of London |  | 12 | 82.5364 | 90.3958 | -8.694% | G | 1.0952
Glasgow Scientific Services | 20% TEA in water | 2016 | R | East Dunbartonshire Council |  | 11 | 33.5091 | 36.7273 | -8.762% | P | 1.096
Glasgow Scientific Services | 20% TEA in water | 2016 | R | East Dunbartonshire Council |  | 12 | 35.1583 | 36.0833 | -2.564% | G | 1.0263
Glasgow Scientific Services | 20% TEA in water | 2016 | R | East Dunbartonshire Council |  | 12 | 28.3417 | 26.75 | 5.95% | P | 0.94384
Glasgow Scientific Services | 20% TEA in water | 2016 | R | East Dunbartonshire Council |  | 12 | 23.5833 | 24.25 | -2.749% | P | 1.0283
Gradko | 20% TEA in water | 2016 | UB | Eastleigh Borough Council |  | 11 | 28.6445 | 30.0457 | -4.663% | G | 1.0489
Gradko | 20% TEA in water | 2016 | R | Eastleigh Borough Council |  | 12 | 43.6172 | 42.3786 | 2.923% | G | 0.9716
Staffordshire Scientific Services | 20% TEA in water | 2016 | UB | Manchester City Council |  | 11 | 43.0212 | 40.506 | 6.209% | G | 0.94154
Staffordshire Scientific Services | 20% TEA in water | 2016 | SI | Manchester City Council |  | 10 | 23.2517 | 22.5451 | 3.134% | G | 0.96961
Northampton BC | 20% TEA in water | 2016 | UB | Northampton |  | 9 | 17.6148 | 15.7222 | 12.04% | G | 0.89256
Northampton BC | 20% TEA in water | 2016 | R | Northampton BC |  | 12 | 42.975 | 37.8333 | 13.59% | G | 0.88036
Somerset County Council | 20% TEA in water | 2016 | R | South Gloucestershire Council |  | 12 | 27.9089 | 24.3167 | 14.77% | G | 0.87129
Staffordshire Scientific Services | 20% TEA in water | 2016 | UC | Stoke-on-Trent City Council |  | 12 | 32.7611 | 29.1011 | 12.58% | G | 0.88828
Staffordshire Scientific Services | 20% TEA in water | 2016 | R | Stoke-on-Trent City Council |  | 9 | 55.7741 | 44.3938 | 25.63% | G | 0.79596
Gradko | 50% TEA in acetone | 2016 | UB | Middlesbrough |  | 12 | 16.6603 | 18.0667 | -7.784% | G | 1.0844
Somerset County Council | 20% TEA in water | 2016 | R | Bath & North East Somerset |  | 11 | 48.967 | 48.4515 | 1.064% | G | 0.98947
Gradko | 20% TEA in water | 2016 | R | Brighton & Hove City Council |  | 12 | 51.7594 | 47.5652 | 8.818% | G | 0.91897
Gradko | 20% TEA in water | 2016 | R | Eastleigh Borough Council |  | 11 | 29.2305 | 37.4818 | -22.01% | G | 1.2823
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | UC | Leeds City Council |  | 12 | 43.0278 | 32.1417 | 33.87% | G | 0.747
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | R | Leeds City Council |  | 9 | 56.8889 | 42.0889 | 35.16% | S | 0.73984
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | R | Leeds City Council |  | 11 | 48.0909 | 39.6455 | 21.3% | S | 0.82439
ESG Didcot | 50% TEA in acetone | 2016 | KS | Leeds City Council |  | 9 | 65.6667 | 54.6667 | 20.12% | S | 0.83249
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | R | Leeds City Council |  | 11 | 45.5455 | 36.0909 | 26.2% | S | 0.79242
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 91.5611 | 79.2805 | 15.49% | G | 0.86587
Glasgow Scientific Services | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 93.0389 | 79.2805 | 17.35% | G | 0.85212
Lambeth Scientific Services | 50% TEA in acetone | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 84.0778 | 79.2805 | 6.051% | G | 0.94294
ESG Glasgow | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 11 | 98.796 | 78.414 | 25.99% | G | 0.7937
ESG Glasgow | 50% TEA in acetone | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 101.767 | 79.2805 | 28.36% | G | 0.77904
ESG Didcot | 50% TEA in acetone | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 103.69 | 79.2805 | 30.79% | G | 0.76459
ESG Didcot | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 104.268 | 79.2805 | 31.52% | G | 0.76035
Milton Keynes Council | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 107.3972 | 79.2805 | 35.46% | G | 0.7382
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 107.247 | 79.2805 | 35.28% | G | 0.73923
Northampton BC | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 101.5608 | 79.2805 | 28.1% | G | 0.78062
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 114.6111 | 79.2805 | 44.56% | G | 0.69173
Gradko | 50% TEA in acetone | 2016 | KS | Marylebone Road Intercomparison |  | 11 | 79.8208 | 78.0601 | 2.256% | G | 0.97794
Gradko | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 99.2362 | 79.2805 | 25.17% | G | 0.79891
Aberdeen Scientific Services | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 92.4928 | 79.2805 | 16.67% | G | 0.85715
Staffordshire Scientific Services | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 92.2746 | 79.2805 | 16.39% | G | 0.85918
Tayside Scientific Services | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 103.0639 | 79.2805 | 30% | G | 0.76924
Somerset County Council | 20% TEA in water | 2016 | KS | Marylebone Road Intercomparison |  | 12 | 100.0408 | 79.2805 | 26.19% | G | 0.79248
Gradko | 20% TEA in water | 2016 | R | Monmouthshire County Council |  | 11 | 39.2913 | 33.6879 | 16.63% | G | 0.85739
Gradko | 20% TEA in Water | 2016 | R | Preston City Council |  | 10 | 29.601 | 26.92 | 9.959% | G | 0.90943
ESG Didcot | 50% TEA in acetone | 2016 | R | Slough Borough Council |  | 12 | 42.7806 | 40.0833 | 6.729% | G | 0.93695
ESG Didcot | 50% TEA in acetone | 2016 | UB | Slough Borough Council |  | 11 | 37.8303 | 29.9091 | 26.48% | G | 0.79061
ESG Didcot | 50% TEA in acetone | 2016 | UB | Slough Borough Council |  | 12 | 34.1944 | 28.5833 | 19.63% | G | 0.83591
Staffordshire Scientific Services | 20% TEA in water | 2016 | UB | Trafford |  | 10 | 24.245 | 20.8 | 16.56% | G | 0.85791
Staffordshire Scientific Services | 20% TEA in water | 2016 | KS | Trafford |  | 11 | 37.65 | 32.1818 | 16.99% | G | 0.85476
ESG Didcot | 50% TEA in acetone | 2016 | R | Tunbridge Wells |  | 12 | 57.3486 | 43.9167 | 30.59% | G | 0.76578
ESG Didcot | 50% TEA in acetone | 2016 | R | Cambridge City Council |  | 10 | 48.6533 | 36.7 | 32.57% | G | 0.75432
ESG Didcot | 50% TEA in acetone | 2016 | R | City of Wolverhampton Council |  | 12 | 44.2778 | 39 | 13.53% | G | 0.8808
ESG Didcot | 50% TEA in acetone | 2016 | R | City of Wolverhampton Council |  | 11 | 53.2273 | 43.3636 | 22.75% | G | 0.81469
Gradko | 20% TEA in water | 2016 | R | Dudley MBC |  | 12 | 37.4832 | 33.7583 | 11.03% | G | 0.90063
Gradko | 20% TEA in water | 2016 | UB | Dudley MBC |  | 12 | 25.8955 | 21.8333 | 18.61% | G | 0.84313
Gradko | 20% TEA in water | 2016 | R | Dudley MBC |  | 11 | 43.1682 | 38.4 | 12.42% | G | 0.88954
Gradko | 20% TEA in water | 2016 | R | Dudley MBC |  | 12 | 51.3966 | 54.4583 | -5.622% | G | 1.0596
ESG Didcot | 50% TEA in acetone | 2016 | B | Gravesham Borough Council |  | 12 | 30.5861 | 22.9167 | 33.47% | G | 0.74925
ESG Didcot | 50% TEA in acetone | 2016 | B | Gravesham Borough Council |  | 12 | 40.2542 | 29.5833 | 36.07% | G | 0.73491
ESG Didcot | 50% TEA in acetone | 2016 | R | Horsham District Council |  | 12 | 35.2306 | 27.0417 | 30.28% | G | 0.76756
ESG Didcot | 50% TEA in acetone | 2016 | KS | Horsham District Council |  | 11 | 32.7121 | 28.7091 | 13.94% | G | 0.87763
ESG Didcot | 50% TEA in acetone | 2016 | R | Horsham District Council |  | 10 | 33.87 | 24.29 | 39.44% | G | 0.71715
ESG Didcot | 50% TEA in acetone | 2016 | B | Maidstone Borough Council |  | 11 | 15.2606 | 12.1818 | 25.27% | G | 0.79825
ESG Didcot | 50% TEA in acetone | 2016 | R | Medway Council |  | 12 | 35.0542 | 25.6667 | 36.57% | G | 0.7322
ESG Didcot | 50% TEA in acetone | 2016 | B | Medway Council |  | 9 | 20.8981 | 11.1111 | 88.08% | G | 0.53168
Staffordshire Scientific Services | 20% TEA in water | 2016 | UB | Salford City Council |  | 11 | 33.8576 | 29.1273 | 16.24% | G | 0.86029
Staffordshire Scientific Services | 20% TEA in water | 2016 | R | Salford City Council |  | 11 | 50.4273 | 45.6727 | 10.41% | G | 0.90571
ESG Didcot | 50% TEA in acetone | 2016 | KS | Suffolk Coastal DC |  | 12 | 43.3861 | 37 | 17.26% | G | 0.85281
ESG Didcot | 50% TEA in acetone | 2016 | UB | City of York Council |  | 9 | 21.5037 | 15.5116 | 38.63% | G | 0.72135
ESG Didcot | 50% TEA in acetone | 2016 | R | City of York Council |  | 12 | 38.9556 | 29.0483 | 34.11% | G | 0.74568
ESG Didcot | 50% TEA in acetone | 2016 | R | City of York Council |  | 12 | 32.9028 | 24.6667 | 33.39% | G | 0.74968
ESG Didcot | 50% TEA in acetone | 2016 | R | City of York Council |  | 12 | 41.0778 | 27.1667 | 51.21% | G | 0.66135
Gradko | 20% TEA in water | 2016 | B | LB Waltham Forest |  | 12 | 30.6061 | 29.9167 | 2.305% | G | 0.97747
ESG Didcot | 50% TEA in acetone | 2016 | KS | Leeds City Council |  | 9 | 65.6667 | 54.6667 | 20.12% | S | 0.83249
ESG Didcot | 50% TEA in acetone | 2016 | R | Leeds City Council |  | 12 | 56.75 | 44.4583 | 27.65% | S | 0.78341
Gradko | 20% TEA in water | 2016 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 36.8806 | 39 | -5.434% | G | 1.0575
Aberdeen Scientific Services | 20% TEA in water | 2016 | UB | Aberdeen City Council |  | 12 | 27.0833 | 21.3108 | 27.09% | G | 0.78686
Aberdeen Scientific Services | 20% TEA in water | 2016 | R | Aberdeen City Council |  | 13 | 31.6923 | 27.6904 | 14.45% | G | 0.87373
Aberdeen Scientific Services | 20% TEA in water | 2016 | R | Aberdeen City Council |  | 12 | 43.9722 | 35.3892 | 24.25% | G | 0.80481
Aberdeen Scientific Services | 20% TEA in water | 2016 | R | Aberdeen City Council |  | 12 | 49.8333 | 42.5908 | 17% | G | 0.85467
Aberdeen Scientific Services | 20% TEA in water | 2016 | R | Aberdeen City Council |  | 13 | 42.8462 | 45.8269 | -6.504% | G | 1.0696
Aberdeen Scientific Services | 20% TEA in water | 2016 | R | Aberdeen City Council |  | 12 | 28.0833 | 21.3747 | 31.39% | G | 0.76112
ESG Didcot | 50% TEA in acetone | 2016 | R | City and County Swansea |  | 9 | 35.0259 | 31.0756 | 12.71% | G | 0.88722
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | R | City of Edinburgh Council |  | 12 | 40.7028 | 27.9667 | 45.54% | G | 0.68709
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | R | City of Edinburgh Council |  | 12 | 39.8806 | 33.1 | 20.49% | G | 0.82998
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | R | City of Edinburgh Council |  | 9 | 65.0037 | 43.4667 | 49.55% | G | 0.66868
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | R | City of Edinburgh Council |  | 11 | 34.5848 | 26.4545 | 30.73% | G | 0.76492
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | KS | City of Edinburgh Council |  | 12 | 64.4333 | 53.775 | 19.82% | G | 0.83458
Tayside Scientific Services | 20% TEA in water | 2016 | R | Dundee City Council |  | 12 | 52.7583 | 45.3 | 16.46% | G | 0.85863
Tayside Scientific Services | 20% TEA in water | 2016 | R | Dundee City Council |  | 10 | 50.7267 | 36.37 | 39.47% | G | 0.71698
Tayside Scientific Services | 20% TEA in water | 2016 | R | Dundee City Council |  | 12 | 49.7417 | 46.95 | 5.946% | G | 0.94388
Tayside Scientific Services | 20% TEA in water | 2016 | R | Dundee City Council |  | 11 | 43.7576 | 36.9545 | 18.41% | G | 0.84453
Gradko | 20% TEA in water | 2016 | R | LB Hounslow |  | 9 | 74.8074 | 58.4444 | 28% | G | 0.78127
Gradko | 20% TEA in water | 2016 | UB | LB Hounslow |  | 9 | 33.1529 | 33.1111 | 0.1262% | G | 0.99874
Gradko | 20% TEA in water | 2016 | R | Lisburn & Castlereagh City Council |  | 12 | 38.7322 | 26.4583 | 46.39% | G | 0.68311
Lambeth Scientific Services | 50% TEA in acetone | 2016 | R | London Borough of Haringey |  | 11 | 35.5455 | 42.8182 | -16.99% | S | 1.2046
ESG Didcot | 50% TEA in acetone | 2016 | R | North East Lincolnshire Council |  | 10 | 36.2633 | 30.23 | 19.96% | G | 0.83362
ESG Didcot | 50% TEA in acetone | 2016 | R | North East Lincolnshire Council |  | 10 | 57.28 | 41.73 | 37.26% | G | 0.72853
ESG Didcot | 50% TEA in acetone | 2016 | R | North East Lincolnshire Council |  | 11 | 44.003 | 28.9455 | 52.02% | G | 0.65781
Gradko | 20% TEA in water | 2016 | B | Pembrokeshire Council |  | 11 | 4.4061 | 3.4545 | 27.54% | G | 0.78404
Gradko | 50% TEA in acetone | 2016 | SU | Redcar and Cleveland Borough Council |  | 10 | 13.3281 | 11.9 | 12% | G | 0.89285
Lambeth Scientific Services | 50% TEA in acetone | 2016 | SU | Reigate and Banstead BC |  | 12 | 20.9167 | 20.0663 | 4.238% | G | 0.95935
ESG Didcot | 50% TEA in acetone | 2016 | SU | Reigate and Banstead BC |  | 12 | 26.8139 | 20.0663 | 33.63% | G | 0.74835
Lambeth Scientific Services | 50% TEA in acetone | 2016 | SU | Reigate and Banstead BC |  | 12 | 26.8611 | 28.5118 | -5.79% | P | 1.0615
Lambeth Scientific Services | 50% TEA in acetone | 2016 | B | Reigate and Banstead BC |  | 12 | 16.1111 | 16.6421 | -3.19% | P | 1.033
ESG Didcot | 50% TEA in acetone | 2016 | B | Reigate and Banstead BC |  | 12 | 20.0861 | 16.6421 | 20.69% | G | 0.82854
ESG Didcot | 20% TEA in water | 2016 | R | Rhondda Cynon Taf CBC |  | 11 | 31.6242 | 31.6645 | -0.1273% | G | 1.0013
ESG Didcot | 50% TEA in acetone | 2016 | KS | Slough Borough Council |  | 11 | 42.1167 | 32.9982 | 27.63% | G | 0.78349
Staffordshire Scientific Services | 20% TEA in water | 2016 | UB | South Staffordshire Council |  | 10 | 35.39 | 43.18 | -18.04% | G | 1.2201
Gradko | 20% TEA in water | 2016 | R | Cheltenham Borough Council |  | 11 | 31.9579 | 32.2545 | -0.9198% | G | 1.0093
Gradko | 20% TEA in water | 2016 | R | Lancaster City Council |  | 11 | 33.3011 | 32.4 | 2.781% | G | 0.97294
ESG Didcot | 50% TEA in acetone | 2016 | R | Wrexham County Borough Council |  | 9 | 19.5444 | 18.0578 | 8.233% | G | 0.92393
Gradko | 50% TEA in acetone | 2016 | R | London Borough of Croydon |  | 11 | 53.3259 | 46.7282 | 14.12% | G | 0.87628
Staffordshire Scientific Services | 20% TEA in water | 2016 | R | Cannock Chase Council |  | 10 | 29.7533 | 26.96 | 10.36% | P | 0.90612
Gradko | 20% TEA in Water | 2016 | R | City of Lincoln Council |  | 11 | 46.472 | 38.45 | 20.86% | G | 0.82738
Lambeth Scientific Services | 50% TEA in acetone | 2016 | R | Epsom & Ewell Borough Council |  | 12 | 38.6111 | 45.25 | -14.67% | P | 1.1719
Gradko | 20% TEA in WATER | 2016 | R | Fareham Borough Council |  | 12 | 33.0044 | 25.9792 | 27.04% | G | 0.78714
Gradko | 20% TEA in WATER | 2016 | R | Fareham Borough Council |  | 12 | 38.5292 | 36.5783 | 5.333% | G | 0.94937
Gradko | 20% TEA in WATER | 2016 | R | Fareham Borough Council |  | 9 | 27.0946 | 32.3333 | -16.2% | G | 1.1933
Milton Keynes Council | 20% TEA in water | 2016 | R | Milton Keynes Council |  | 10 | 28.6186 | 18.3796 | 55.71% | G | 0.64223
Milton Keynes Council | 20% TEA in water | 2016 | R | Milton Keynes Council |  | 11 | 36.8097 | 22.9974 | 60.06% | G | 0.62476
GRADKO | 50% TEA in acetone | 2016 | UB | Norwich City Council |  | 10 | 12.8967 | 14.64 | -11.91% | G | 1.1352
GRADKO | 20% TEA IN WATER | 2016 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 34.4972 | 38.3333 | -10.01% | P | 1.1112
Gradko | 20% TEA in water | 2017 | UB | Bracknell Forest Borough Council |  | 11 | 19.4589 | 15.8163 | 23.03% | G | 0.81281
Gradko | 20% TEA in water | 2017 | R | Bracknell Forest Borough Council |  | 12 | 46.8898 | 38.5366 | 21.68% | G | 0.82185
Gradko | 20% TEA in water | 2017 | R | Brighton & Hove City Council |  | 11 | 50.7324 | 49.9416 | 1.583% | G | 0.98441
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | East Staffordshire Borough Council |  | 12 | 50.1556 | 47.1583 | 6.356% | G | 0.94024
ESG Didcot | 20% TEA in water | 2017 | KS | South Lakeland District Council |  | 11 | 43.1667 | 28.2727 | 52.68% | G | 0.65497
ESG Glasgow | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 98.0627 | 78.6969 | 24.61% | G | 0.80252
ESG Didcot | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 103.8912 | 78.6969 | 32.01% | G | 0.75749
Glasgow Scientific Services | 20% TEA in water | 2017 | UB | Glasgow City Council |  | 12 | 33.5125 | 25.225 | 32.85% | G | 0.7527
Glasgow Scientific Services | 20% TEA in water | 2017 | R | Glasgow City Council |  | 12 | 38.2583 | 37.1917 | 2.868% | G | 0.97212
Glasgow Scientific Services | 20% TEA in water | 2017 | R | Glasgow City Council |  | 10 | 35.0467 | 33.84 | 3.566% | P | 0.96557
Glasgow Scientific Services | 20% TEA in water | 2017 | KS | Glasgow City Council |  | 12 | 62.5069 | 58.8833 | 6.154% | G | 0.94203
Glasgow Scientific Services | 20% TEA in water | 2017 | R | Glasgow City Council |  | 12 | 45.2861 | 36.375 | 24.5% | P | 0.80323
Gradko | 20% TEA in water | 2017 | R | Wokingham Borough Council |  | 11 | 39.02 | 37.3046 | 4.598% | G | 0.95604
Gradko | 20% TEA in water | 2017 | UC | Southampton City Council |  | 11 | 31.0121 | 29.4455 | 5.321% | G | 0.94948
Gradko | 20% TEA in water | 2017 | R | Preston City Council |  | 12 | 31.4578 | 25.5167 | 23.28% | G | 0.81114
Gradko | 20% TEA in water | 2017 | R | Monmouthshire County Council |  | 9 | 42.039 | 33.2182 | 26.55% | G | 0.79017
Gradko | 20% TEA in water | 2017 | R | Cheshire West and Chester |  | 11 | 36.2403 | 35.7545 | 1.359% | G | 0.98659
Gradko | 20% TEA in water | 2017 | UI | Crawley Borough Council |  | 12 | 28.0775 | 28.405 | -1.153% | G | 1.0117
Gradko | 20% TEA in water | 2017 | R | Borough Council of King's Lynn & West Norfolk |  | 12 | 28.8664 | 24.875 | 16.05% | G | 0.86173
Gradko | 20% TEA in water | 2017 | R | Bath & North East Somerset |  | 12 | 44.6759 | 44.7716 | -0.2137% | G | 1.0021
Gradko | 20% TEA in water | 2017 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 38.4011 | 41.125 | -6.623% | G | 1.0709
Gradko | 20% TEA in water | 2017 | R | Lancaster City Council |  | 12 | 35.3981 | 32.2583 | 9.733% | G | 0.9113
Milton Keynes Council | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 88.1833 | 78.6969 | 12.05% | G | 0.89242
Staffordshire Scientific Services | 20% TEA in water | 2017 | O | South Staffordshire Council |  | 11 | 28.8303 | 41.4036 | -30.37% | G | 1.4361
Staffordshire Scientific Services | 20% TEA in water | 2017 | UB | Wigan Council |  | 11 | 25.703 | 17.5182 | 46.72% | G | 0.68156
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Stoke-on-Trent City Council |  | 12 | 54.55 | 51.9157 | 5.074% | G | 0.95171
Staffordshire Scientific Services | 20% TEA in water | 2017 | UB | Stoke-on-Trent City Council |  | 12 | 30.2139 | 25.7544 | 17.32% | G | 0.8524
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Salford City Council |  | 12 | 46.0083 | 43.8833 | 4.842% | G | 0.95381
Staffordshire Scientific Services | 20% TEA in water | 2017 | UB | Salford City Council |  | 12 | 31.6944 | 26.1083 | 21.4% | G | 0.82375
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Cannock Chase Council |  | 10 | 26.3367 | 21.89 | 20.31% | G | 0.83116
Tayside Scientific Services | 20% TEA in water | 2017 | R | Fife Council |  | 12 | 29.9556 | 21.875 | 36.94% | G | 0.73025
Tayside Scientific Services | 20% TEA in water | 2017 | R | Fife Council |  | 12 | 25.4222 | 17.9167 | 41.89% | G | 0.70476
Tayside Scientific Services | 20% TEA in water | 2017 | R | Fife Council |  | 12 | 31.8528 | 21.875 | 45.61% | G | 0.68675
Tayside Scientific Services | 20% TEA in water | 2017 | KS | Fife Council |  | 12 | 36.3722 | 26.2083 | 38.78% | G | 0.72056
Gradko | 20% TEA in water | 2017 | R | Thurrock Borough Council |  | 12 | 53.7208 | 52.0132 | 3.283% | S | 0.96821
Gradko | 20% TEA in water | 2017 | R | Thurrock Borough Council |  | 11 | 35.4973 | 33.1686 | 7.021% | G | 0.9344
Gradko | 20% TEA in water | 2017 | R | Thurrock Borough Council |  | 9 | 32.7641 | 28.6589 | 14.32% | G | 0.87471
Gradko | 20% TEA in water | 2017 | UB | Thurrock Borough Council |  | 11 | 30.04 | 27.8052 | 8.038% | S | 0.9256
Gradko | 20% TEA in water | 2017 | R | Dudley MBC |  | 12 | 50.234 | 49.8333 | 0.8041% | G | 0.99202
Gradko | 20% TEA in water | 2017 | UB | Dudley MBC |  | 12 | 23.9002 | 18.875 | 26.62% | G | 0.78974
Gradko | 20% TEA in water | 2017 | R | City of Lincoln Council |  | 12 | 41.5972 | 31.235 | 33.18% | G | 0.75089
Gradko | 20% TEA in water | 2017 | R | Gedling Borough Council |  | 12 | 34.5417 | 31.3667 | 10.12% | G | 0.90808
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | R | Stirling Council |  | 10 | 24.7467 | 22.5 | 9.985% | G | 0.90921
ESG Didcot | 50% TEA in acetone | 2017 | R | Suffolk Coastal DC |  | 12 | 44.5 | 36.5678 | 21.69% | G | 0.82175
ESG Didcot | 50% TEA in acetone | 2017 | R | Dumfries and Galloway Council |  | 12 | 36.1597 | 30.2248 | 19.64% | G | 0.83587
ESG Glasgow | 50% TEA in acetone | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 101.1296 | 78.6969 | 28.51% | G | 0.77818
ESG Didcot | 50% TEA in acetone | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 105.6803 | 78.6969 | 34.29% | G | 0.74467
Gradko | 50% TEA in acetone | 2017 | R | West Berkshire |  | 12 | 39.9 | 39.7851 | 0.2889% | G | 0.99712
Gradko | 50% TEA in acetone | 2017 | UB | London Borough of Camden |  | 12 | 40.4 | 39.6517 | 1.887% | G | 0.98148
Gradko | 50% TEA in acetone | 2017 | B | London Borough of Richmond upon Thames |  | 12 | 21.4197 | 21.325 | 0.4442% | G | 0.99558
Gradko | 50% TEA in acetone | 2017 | R | London Borough of Richmond upon Thames |  | 12 | 34.6398 | 31.0333 | 11.62% | G | 0.89589
Gradko | 50% TEA in acetone | 2017 | R | Royal Borough of Greenwich |  | 12 | 33.2028 | 35.9267 | -7.582% | G | 1.082
Gradko | 50% TEA in acetone | 2017 | R | Royal Borough of Greenwich |  | 12 | 40.0247 | 39.1958 | 2.115% | G | 0.97929
Gradko | 50% TEA in acetone | 2017 | R | Royal Borough of Greenwich |  | 10 | 74.7363 | 66.29 | 12.74% | G | 0.88698
Gradko | 50% TEA in acetone | 2017 | R | Royal Borough of Greenwich |  | 12 | 43.956 | 38.6817 | 13.64% | G | 0.88001
Gradko | 50% TEA in acetone | 2017 | R | Royal Borough of Greenwich |  | 11 | 48.5739 | 41.4645 | 17.15% | G | 0.85364
Gradko | 50% TEA in acetone | 2017 | SU | Royal Borough of Greenwich |  | 12 | 20.3869 | 20.6067 | -1.066% | G | 1.0108
Gradko | 50% TEA in acetone | 2017 | UI | Falkirk Council |  | 12 | 18.0333 | 17.1667 | 5.049% | G | 0.95194
Gradko | 50% TEA in acetone | 2017 | R | Falkirk Council |  | 10 | 38.0167 | 35.1 | 8.31% | G | 0.92328
ESG Didcot | 50% TEA in acetone | 2017 | R | Vale of White Horse District Council |  | 11 | 31.2121 | 24.7636 | 26.04% | G | 0.7934
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2017 | R | Barnsley MBC |  | 12 | 33.7778 | 35.1583 | -3.927% | G | 1.0409
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 103.9472 | 78.6969 | 32.09% | G | 0.75709
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | R | Calderdale MBC |  | 12 | 53.0833 | 48.3417 | 9.809% | G | 0.91068
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | R | Wakefield Council |  | 12 | 35.7917 | 27.6583 | 29.41% | G | 0.77276
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | UC | Wakefield Council |  | 10 | 40.3333 | 28.5037 | 41.5% | G | 0.7067
ESG Didcot | 50% TEA in acetone | 2017 | UB | Cardiff City Council |  | 10 | 28.9567 | 21.43 | 35.12% | G | 0.74007
Gradko | 50% TEA in acetone | 2017 | R | LB Newham |  | 12 | 38.3847 | 37.5752 | 2.154% | G | 0.97891
ESG Didcot | 50% TEA in acetone | 2017 | R | Cambridge City Council |  | 12 | 44.9681 | 32.6667 | 37.66% | G | 0.72644
Gradko | 50% TEA in acetone | 2017 | B | The City of London Corporation |  | 12 | 41.4294 | 38.1167 | 8.691% | P | 0.92004
ESG Didcot | 50% TEA in acetone | 2017 | R | Wrexham County Borough Council |  | 12 | 19.8194 | 17.3167 | 14.45% | G | 0.87372
ESG Didcot | 50% TEA in acetone | 2017 | UI | North Lincolnshire Council |  | 12 | 22.3944 | 15.9167 | 40.7% | G | 0.71074
Somerset County Council | 20% TEA in water | 2017 | R | Wiltshire Council |  | 12 | 41 | 29.9092 | 37.08% | G | 0.72949
Aberdeen Scientific Services | 20% TEA IN WATER | 2017 | R | Aberdeen City Council |  | 13 | 23.3846 | 18.24 | 28.21% | G | 0.78
Aberdeen Scientific Services | 20% TEA IN WATER | 2017 | UB | Aberdeen City Council |  | 13 | 26.4359 | 21.2815 | 24.22% | P | 0.80502
Aberdeen Scientific Services | 20% TEA IN WATER | 2017 | R | Aberdeen City Council |  | 12 | 30.6389 | 23.0233 | 33.08% | G | 0.75144
Aberdeen Scientific Services | 20% TEA IN WATER | 2017 | R | Aberdeen City Council |  | 12 | 41.4167 | 30.3325 | 36.54% | G | 0.73237
Aberdeen Scientific Services | 20% TEA IN WATER | 2017 | R | Aberdeen City Council |  | 12 | 48.4722 | 39.3608 | 23.15% | P | 0.81203
Aberdeen Scientific Services | 20% TEA IN WATER | 2017 | R | Aberdeen City Council |  | 12 | 44.7778 | 38.9958 | 14.83% | G | 0.87087
ESG Didcot | 50% TEA in acetone | 2017 | KS | Caerphilly CBC |  | 12 | 37.3472 | 32.25 | 15.81% | G | 0.86352
ESG Didcot | 50% TEA in acetone | 2017 | R | Caerphilly CBC |  | 11 | 44.3939 | 29.3636 | 51.19% | G | 0.66143
ESG Didcot | 50% TEA in acetone | 2017 | UB | City of York Council |  | 12 | 22.9903 | 14.9892 | 53.38% | G | 0.65198
ESG Didcot | 50% TEA in acetone | 2017 | R | City of York Council |  | 10 | 37.2267 | 28.466 | 30.78% | G | 0.76467
ESG Didcot | 50% TEA in acetone | 2017 | R | City of York Council |  | 11 | 32.4091 | 22.9855 | 41% | G | 0.70923
ESG Didcot | 50% TEA in acetone | 2017 | R | City of York Council |  | 12 | 40.3139 | 25.4142 | 58.63% | G | 0.63041
Gradko | 20% TEA in water | 2017 | R | Gateshead Council |  | 12 | 36.2544 | 37.2583 | -2.694% | G | 1.0277
Gradko | 20% TEA in water | 2017 | R | Gateshead Council |  | 12 | 29.399 | 25.015 | 17.53% | G | 0.85088
Gradko | 20% TEA in water | 2017 | R | Gateshead Council |  | 12 | 33.5025 | 35.3667 | -5.271% | G | 1.0556
ESG Didcot | 50% TEA in acetone | 2017 | R | Hambleton District Council |  | 10 | 20.8567 | 20.06 | 3.971% | G | 0.9618
ESG Didcot | 50% TEA in acetone | 2017 | R | Horsham District Council |  | 11 | 34.797 | 29.4571 | 18.13% | G | 0.84654
ESG Didcot | 50% TEA in acetone | 2017 | R | Horsham District Council |  | 12 | 31.2278 | 25.7478 | 21.28% | G | 0.82452
ESG Didcot | 50% TEA in acetone | 2017 | R | Horsham District Council |  | 11 | 33.0758 | 22.4525 | 47.31% | G | 0.67882
Gradko | 20% TEA in water | 2017 | R | LB Hounslow |  | 12 | 65.4038 | 53.5259 | 22.19% | G | 0.81839
Gradko | 20% TEA in water | 2017 | R | LB Hounslow |  | 12 | 58.8604 | 53.2236 | 10.59% | G | 0.90423
Gradko | 20% TEA in water | 2017 | B | LB Hounslow |  | 11 | 28.0869 | 29.8929 | -6.042% | G | 1.0643
Gradko | 20% TEA in water | 2017 | R | LB Hounslow |  | 11 | 43.3311 | 33.6327 | 28.84% | G | 0.77618
Gradko | 20% TEA in water | 2017 | B | LB Hounslow |  | 9 | 38.287 | 33.3268 | 14.88% | G | 0.87045
Gradko | 20% TEA in water | 2017 | R | LB Hounslow |  | 11 | 51.6577 | 41.5407 | 24.35% | G | 0.80415
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | UC | Leeds City Council 1 |  | 12 | 41.25 | 32.1 | 28.5% | G | 0.77818
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | R | Leeds City Council 10 |  | 11 | 47.5455 | 38.0182 | 25.06% | S | 0.79962
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | KS | Leeds City Council 2 |  | 12 | 47.25 | 35.1667 | 34.36% | S | 0.74427
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | R | Leeds City Council 4 |  | 11 | 55.7273 | 43.1545 | 29.13% | S | 0.77439
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | R | Leeds City Council 7 |  | 11 | 37.8182 | 27.0455 | 39.83% | S | 0.71514
Gradko | 20% TEA in water | 2017 | UB | Liverpool |  | 11 | 19.7273 | 17.1273 | 15.18% | G | 0.8682
Staffordshire Scientific Services | 20% TEA in water | 2017 | KS | Manchester City Council |  | 11 | 63.8242 | 64.7355 | -1.408% | G | 1.0143
Staffordshire Scientific Services | 20% TEA in water | 2017 | UC | Manchester City Council |  | 11 | 39.8545 | 35.0027 | 13.86% | G | 0.87826
Staffordshire Scientific Services | 20% TEA in water | 2017 | SI | Manchester City Council |  | 12 | 25.0028 | 23.51 | 6.35% | G | 0.9403
Lambeth Scientific Services | 50% TEA in acetone | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 87.7278 | 78.6969 | 11.48% | G | 0.89706
Gradko | 50% TEA in acetone | 2017 | UI | Middlesbrough |  | 10 | 16.8055 | 13.825 | 21.56% | G | 0.82265
Gradko | 20% TEA in water | 2017 | R | North Ayrshire Council |  | 12 | 26.3764 | 21.4167 | 23.16% | G | 0.81196
Gradko | 50% TEA in acetone | 2017 | UB | Norwich City Council |  | 11 | 13.2017 | 13.8364 | -4.587% | G | 1.0481
Gradko | 50% TEA in acetone | 2017 | R | RBWM |  | 12 | 38.7184 | 37.8083 | 2.407% | G | 0.9765
Gradko | 50% TEA in acetone | 2017 | R | RBWM |  | 12 | 34.7426 | 34.375 | 1.069% | G | 0.98942
Gradko | 50% TEA in acetone | 2017 | UB | Reading Borough Council |  | 12 | 19.7211 | 28.75 | -31.4% | G | 1.4578
Gradko | 50% TEA in acetone | 2017 | SU | Redcar and Cleveland Borough Council |  | 11 | 14.7561 | 11.4918 | 28.4% | P | 0.77879
ESG Didcot | 50% TEA in acetone | 2017 | R | Slough Borough Council |  | 12 | 44.6861 | 35.3519 | 26.4% | G | 0.79112
ESG Didcot | 50% TEA in acetone | 2017 | UB | Slough Borough Council |  | 12 | 32.4778 | 25.2619 | 28.56% | G | 0.77782
ESG Didcot | 50% TEA in acetone | 2017 | UB | Slough Borough Council |  | 11 | 39.2182 | 32.9044 | 19.19% | G | 0.83901
Gradko | 20% TEA in water | 2017 | R | South Gloucestershire Council |  | 12 | 25.0979 | 22.75 | 10.32% | G | 0.90645
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Stockport MBC |  | 9 | 48.7519 | 43.7374 | 11.46% | G | 0.89714
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Stockport MBC |  | 12 | 28.9 | 20.4923 | 41.03% | G | 0.70908
ESG Didcot | 50% TEA in acetone | 2017 | R | Tunbridge Wells |  | 12 | 55.6056 | 40.2333 | 38.21% | G | 0.72355
Gradko | 50% TEA in acetone | 2017 | R | Worthing Borough Council |  | 12 | 41.9269 | 38.4508 | 9.04% | G | 0.91709
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 107.3333 | 78.6969 | 36.39% | G | 0.7332
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 90.8472 | 78.6969 | 15.44% | G | 0.86626
Gradko | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 101.2194 | 78.6969 | 28.62% | G | 0.77749
Glasgow Scientific Services | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 76.9972 | 78.6969 | -2.16% | G | 1.0221
Gradko | 50% TEA in acetone | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 83.3856 | 78.6969 | 5.958% | G | 0.94377
Aberdeen Scientific Services | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 96.0986 | 78.6969 | 22.11% | G | 0.81892
Staffordshire Scientific Services | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 89.0333 | 78.6969 | 13.13% | G | 0.8839
Tayside Scientific Services | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 103.2583 | 78.6969 | 31.21% | G | 0.76214
Somerset County Council | 20% TEA in water | 2017 | KS | Marylebone Road Intercomparison |  | 12 | 97.4116 | 78.6969 | 23.78% | G | 0.80788
Gradko | 20% TEA in water | 2017 | R | Ards and North Down Borough Council |  | 11 | 40 | 25.1182 | 59.25% | G | 0.62795
Gradko | 20% TEA in water | 2017 | UC | Belfast City Council |  | 12 | 31.9233 | 25.0375 | 27.5% | G | 0.7843
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Bury Council |  | 10 | 34.4933 | 29.13 | 18.41% | G | 0.84451
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Cannock Chase Council |  | 11 | 26.6091 | 22.2636 | 19.52% | G | 0.83669
Staffordshire Scientific Services | 20% TEA in water | 2017 | R | Cannock Chase Council |  | 9 | 24.3963 | 20.2556 | 20.44% | G | 0.83027
Glasgow Scientific Services | 20% TEA in water | 2017 | R | East Dunbartonshire Council |  | 11 | 31.5394 | 28.2727 | 11.55% | P | 0.89643
Glasgow Scientific Services | 20% TEA in water | 2017 | R | East Dunbartonshire Council |  | 11 | 28.9727 | 25.0909 | 15.47% | G | 0.86602
Glasgow Scientific Services | 20% TEA in water | 2017 | R | East Dunbartonshire Council |  | 11 | 22.1545 | 21.2727 | 4.145% | G | 0.9602
Glasgow Scientific Services | 20% TEA in water | 2017 | R | East Dunbartonshire Council |  | 11 | 37.8909 | 32.8182 | 15.46% | P | 0.86612
ESG Didcot | 50% TEA in acetone | 2017 | UB | Kingston upon Hull City Council |  | 12 | 32.0972 | 23.2175 | 38.25% | G | 0.72335
Gradko | 20% TEA in water | 2017 | R | Lisburn & Castlereagh City Council |  | 12 | 40.5 | 26.6583 | 51.92% | G | 0.65823
Lambeth Scientific Services | 50% TEA in acetone | 2017 | R | London Borough of Islington |  | 11 | 48.1515 | 49.3273 | -2.384% | G | 1.0244
Milton Keynes Council | 20% TEA in water | 2017 | UC | Milton Keynes Council |  | 12 | 23.6824 | 17.0167 | 39.17% | G | 0.71854
Milton Keynes Council | 20% TEA in water | 2017 | R | Milton Keynes Council |  | 12 | 38.9903 | 30.6091 | 27.38% | G | 0.78505
Milton Keynes Council | 20% TEA in water | 2017 | R | Milton Keynes Council |  | 12 | 32.7449 | 22.2578 | 47.12% | G | 0.67973
ESG Didcot | 20% TEA in water | 2017 | R | New Forest DC |  | 9 | 37.0481 | 26.5103 | 39.75% | G | 0.71556
ESG Didcot | 20% TEA in water | 2017 | KS | New Forest DC |  | 10 | 53.0233 | 35.4331 | 49.64% | G | 0.66825
Gradko | 50% TEA in acetone | 2017 | UB | Norwich City Council |  | 11 | 13.1527 | 13.8364 | -4.941% | G | 1.052
Lambeth Scientific Services | 50% TEA in acetone | 2017 | SU | Reigate and Banstead BC |  | 12 | 23.3333 | 20.4699 | 13.99% | G | 0.87728
Lambeth Scientific Services | 50% TEA in acetone | 2017 | B | Reigate and Banstead BC |  | 12 | 15.3333 | 13.8482 | 10.72% | G | 0.90314
Lambeth Scientific Services | 50% TEA in acetone | 2017 | SU | Reigate and Banstead BC |  | 12 | 28.1389 | 26.7073 | 5.36% | G | 0.94912
ESG Didcot | 50% TEA in acetone | 2017 | R | Suffolk Coastal DC |  | 12 | 45.2583 | 36.5678 | 23.77% | G | 0.80798
Staffordshire Scientific Services | 20% TEA in water | 2017 | B | Trafford |  | 12 | 16.8028 | 14.49 | 15.96% | G | 0.86236
Staffordshire Scientific Services | 20% TEA in water | 2017 | UB | Trafford |  | 10 | 22.2133 | 18.58 | 19.56% | G | 0.83643
ESG Didcot | 50% TEA in acetone | 2017 | R | Dacorum Borough Council |  | 9 | 30.8222 | 26.8744 | 14.69% | G | 0.87192
ESG Didcot | 50% TEA in acetone | 2017 | R | North East Lincolnshire Council |  | 11 | 36.5242 | 23.7909 | 53.52% | G | 0.65137
ESG Didcot | 20% TEA in water | 2017 | R | Rhondda Cynon Taf CBC |  | 13 | 32.2462 | 29.1831 | 10.5% | G | 0.90501
Gradko | 50% TEA in acetone | 2017 | UB | Royal Borough of Kensington and Chelsea |  | 12 | 29.0139 | 32.7417 | -11.39% | G | 1.1285
Gradko | 50% TEA in acetone | 2017 | R | Royal Borough of Kensington and Chelsea |  | 10 | 57.5083 | 50.4746 | 13.94% | G | 0.87769
ESG Didcot | 50% TEA in acetone | 2017 | UB | Swansea Council |  | 10 | 16.84 | 13.642 | 23.44% | G | 0.8101
ESG Didcot | 50% TEA in acetone | 2017 | R | Swansea Council |  | 12 | 32.8056 | 24.3858 | 34.53% | G | 0.74334
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | R | The City of Edinburgh Council |  | 11 | 33.0621 | 29.2727 | 12.95% | G | 0.88539
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | R | The City of Edinburgh Council |  | 11 | 59.6606 | 52.4545 | 13.74% | G | 0.87922
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | R | The City of Edinburgh Council |  | 11 | 41.3242 | 26.2727 | 57.29% | G | 0.63577
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | R | The City of Edinburgh Council |  | 11 | 31.3439 | 24.9091 | 25.83% | P | 0.7947
Gradko | 20% TEA in water | 2017 | R | The Highland Council |  | 11 | 23.9212 | 20.78 | 15.12% | G | 0.86869
Gradko | 20% TEA in water | 2017 | R | The Highland Council |  | 10 | 36.2033 | 39.291 | -7.858% | G | 1.0853
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2017 | UB | Bournemouth Borough Council |  | 10 | 15.4175 | 13.0989 | 17.7% | G | 0.84961
ESG Didcot | 50% TEA in acetone | 2017 | UB | Derry City and Strabane District Council |  | 12 | 13.725 | 9.8667 | 39.1% | G | 0.71888
ESG Didcot | 50% TEA in acetone | 2017 | R | Derry City and Strabane District Council |  | 12 | 35.8833 | 35.55 | 0.9376% | G | 0.99071
Gradko | 20% TEA in water | 2018 | R | Ards and North Down Borough Council |  | 11 | 36.3561 | 28.5364 | 27.4% | G | 0.78491
Gradko | 20% TEA in water | 2018 | R | Gedling Borough Council |  | 12 | 33.3422 | 31.5833 | 5.569% | G | 0.94725
Gradko | 20% TEA in water | 2018 | R | Lisburn & Castlereagh City Council |  | 12 | 31.5278 | 23.8667 | 32.1% | G | 0.757
Milton Keynes Council | 20% TEA in water | 2018 | UC | Milton Keynes Council |  | 12 | 23.6198 | 16.3526 | 44.44% | G | 0.69233
Milton Keynes Council | 20% TEA in water | 2018 | R | Milton Keynes Council |  | 12 | 30.1734 | 20.1744 | 49.56% | G | 0.66862
Milton Keynes Council | 20% TEA in water | 2018 | R | Milton Keynes Council |  | 12 | 35.0852 | 25.7789 | 36.1% | G | 0.73475
Gradko | 20% TEA in water | 2018 | R | Monmouthshire County Council |  | 12 | 37.9747 | 36.269 | 4.703% | G | 0.95509
Gradko | 20% TEA in water | 2018 | UB | Northampton Borough Council |  | 12 | 16.3819 | 12.9167 | 26.83% | G | 0.78847
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | UB | Wakefield Council |  | 12 | 33.5278 | 24.5502 | 36.57% | G | 0.73223
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | R | Wakefield Council |  | 10 | 32.7333 | 25.5039 | 28.35% | G | 0.77914
Aberdeen Scientific Services | 20% TEA in water | 2018 | R | Aberdeen City Council |  | 13 | 24.4359 | 17.8138 | 37.17% | G | 0.729
Aberdeen Scientific Services | 20% TEA in water | 2018 | UB | Aberdeen City Council |  | 13 | 25.7692 | 20.3846 | 26.42% | G | 0.79104
Aberdeen Scientific Services | 20% TEA in water | 2018 | R | Aberdeen City Council |  | 11 | 27.4545 | 22.0909 | 24.28% | G | 0.80464
Aberdeen Scientific Services | 20% TEA in water | 2018 | R | Aberdeen City Council |  | 13 | 43.7949 | 32.9231 | 33.02% | G | 0.75176
Aberdeen Scientific Services | 20% TEA in water | 2018 | R | Aberdeen City Council |  | 13 | 46.5128 | 39 | 19.26% | G | 0.83848
Aberdeen Scientific Services | 20% TEA in water | 2018 | R | Aberdeen City Council |  | 13 | 45.3846 | 40.1538 | 13.03% | G | 0.88475
Gradko | 20% TEA in water | 2018 | R | Bedford Borough Council |  | 11 | 31.8636 | 29.1818 | 9.19% | G | 0.91583
Gradko | 20% TEA in water | 2018 | R | Borough Council of King's Lynn and West Norfolk |  | 12 | 25.6111 | 24.1583 | 6.014% | G | 0.94328
Socotec Didcot | 50% TEA in acetone | 2018 | R | Cambridge City Council |  | 12 | 42.1917 | 30.0908 | 40.21% | G | 0.71319
Socotec Didcot | 50% TEA in acetone | 2018 | R | Canterbury City Council |  | 11 | 37.8515 | 27.8636 | 35.85% | G | 0.73613
Socotec Didcot | 50% TEA in acetone | 2018 | UB | Canterbury City Council |  | 12 | 15.8278 | 11.6083 | 36.35% | G | 0.73342
Gradko | 20% TEA in water | 2018 | R | Cheshire West and Chester |  | 12 | 36.3268 | 37.25 | -2.478% | G | 1.0254
Gradko | 20% TEA in water | 2018 | R | Cheshire West and Chester |  | 12 | 42.7514 | 40.3083 | 6.061% | G | 0.94285
Gradko | 50% TEA in acetone | 2018 | R | City of London |  | 12 | 83.8574 | 93.8833 | -10.68% | G | 1.1196
Gradko | 50% TEA in acetone | 2018 | B | City of London |  | 10 | 38.4454 | 31.796 | 20.91% | G | 0.82704
Gradko | 20% TEA in water | 2018 | R | Fareham Borough Council |  | 12 | 27.6539 | 33.5075 | -17.47% | G | 1.2117
Gradko | 20% TEA in water | 2018 | R | Fareham Borough Council |  | 12 | 37.3725 | 34.33 | 8.863% | G | 0.91859
Gradko | 20% TEA in water | 2018 | R | Fareham Borough Council |  | 12 | 32.0097 | 28.4217 | 12.62% | G | 0.88791
Socotec Didcot | 50% TEA in acetone | 2018 | R | Hambleton District Council |  | 12 | 21.4667 | 17.775 | 20.77% | G | 0.82803
Socotec Didcot | 50% TEA in acetone | 2018 | R | Ipswich Borough Council |  | 12 | 33.8458 | 28.1167 | 20.38% | G | 0.83073
Gradko | 20% TEA in water | 2018 | R | NOTTINGHAM CITY COUNCIL |  | 12 | 34.5986 | 34.4917 | 0.3101% | G | 0.99691
Gradko | 50% TEA in acetone | 2018 | R | RBWM |  | 12 | 39.1619 | 36.3333 | 7.785% | G | 0.92777
Gradko | 50% TEA in acetone | 2018 | R | RBWM |  | 12 | 34.9934 | 34.225 | 2.245% | G | 0.97804
Socotec Didcot | 20% TEA in water | 2018 | KS | South Lakeland District Council |  | 12 | 38.1556 | 22.45 | 69.96% | G | 0.58838
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | R | Stirling Council |  | 10 | 24.005 | 22.2 | 8.131% | G | 0.92481
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | R | Wakefield Council |  | 12 | 31.6111 | 26.4083 | 19.7% | G | 0.83541
Gradko | 20% TEA in water | 2018 | R | Bracknell Forest Borough Council |  | 12 | 44.0329 | 36.8773 | 19.4% | G | 0.83749
Gradko | 20% TEA in water | 2018 | R | Brighton & Hove City Council |  | 9 | 48.2556 | 50.1333 | -3.746% | G | 1.0389
Socotec Didcot | 50% TEA in acetone | 2018 | R | City of York Council |  | 12 | 40.9278 | 26.5483 | 54.16% | G | 0.64866
Socotec Didcot | 50% TEA in acetone | 2018 | UB | City of York Council |  | 11 | 22.2364 | 14.6273 | 52.02% | G | 0.65781
Socotec Didcot | 50% TEA in acetone | 2018 | R | City of York Council |  | 12 | 34.3722 | 26.2775 | 30.8% | G | 0.7645
Socotec Didcot | 50% TEA in acetone | 2018 | R | City of York Council |  | 11 | 30.2788 | 22.7755 | 32.94% | G | 0.75219
Socotec Didcot | 50% TEA in acetone | 2018 | R | Dumfries and Galloway Council |  | 12 | 35.6361 | 29.7832 | 19.65% | G | 0.83576
Gradko | 20% TEA in water | 2018 | R | Eastleigh Borough Council |  | 11 | 27.7968 | 31.5828 | -11.99% | G | 1.1362
Gradko | 20% TEA in water | 2018 | R | Eastleigh Borough Council |  | 12 | 41.7755 | 37.9259 | 10.15% | G | 0.90785
Gradko | 20% TEA in water | 2018 | UB | Eastleigh Borough Council |  | 12 | 27.086 | 28.3314 | -4.396% | G | 1.046
Gradko | 20% TEA in water | 2018 | R | Gateshead Council |  | 12 | 28.5697 | 25.0792 | 13.92% | G | 0.87782
Gradko | 20% TEA in water | 2018 | R | Gateshead Council |  | 12 | 31.7825 | 28.6833 | 10.8% | G | 0.90249
Gradko | 20% TEA in water | 2018 | R | Gateshead Council |  | 9 | 40.0411 | 40.7867 | -1.828% | G | 1.0186
Socotec Didcot | 50% TEA in acetone | 2018 | R | Knowsley MBC |  | 12 | 47.4583 | 37.5167 | 26.5% | G | 0.79052
Staffordshire Scientific Services | 20% TEA in water | 2018 | UC | Manchester City Council |  | 12 | 36.7333 | 34.9725 | 5.035% | G | 0.95206
Staffordshire Scientific Services | 20% TEA in water | 2018 | KS | Manchester City Council |  | 12 | 60.6583 | 62.0708 | -2.276% | G | 1.0233
Staffordshire Scientific Services | 20% TEA in water | 2018 | SI | Manchester City Council |  | 12 | 23.0472 | 24.1833 | -4.698% | G | 1.0493
Gradko | 50% TEA in acetone | 2018 | SU | Redcar and Cleveland Borough Council |  | 9 | 17.7641 | 9.69 | 83.32% | G | 0.54548
Socotec Didcot | 50% TEA in acetone | 2018 | R | Suffolk Coastal DC |  | 11 | 43.6394 | 32.9596 | 32.4% | G | 0.75527
Gradko | 50% TEA in acetone | 2018 | R | West Berkshire |  | 10 | 40.3413 | 36.5159 | 10.48% | G | 0.90517
Somerset County Council | 20% TEA in water | 2018 | R | Wiltshire Council |  | 13 | 34.7949 | 31.4677 | 10.57% | G | 0.90438
Somerset County Council | 20% TEA in water | 2018 | R | Wiltshire Council |  | 13 | 39.2051 | 32.3669 | 21.13% | G | 0.82558
Gradko | 20% TEA in water | 2018 | R | Wokingham Borough Council |  | 12 | 37.5514 | 33.1872 | 13.15% | G | 0.88378
Gradko | 20% TEA in water | 2018 | R | Bath & North East Somerset |  | 12 | 40.1098 | 38.5496 | 4.047% | G | 0.9611
Gradko | 20% TEA in water | 2018 | R | Bedford Borough Council |  | 10 | 29.7833 | 27.37 | 8.817% | G | 0.91897
Staffordshire Scientific Services | 20% TEA in water | 2018 | R | Bury Council |  | 12 | 30.3306 | 25.4811 | 19.03% | G | 0.84011
Glasgow Scientific Services | 20% TEA in water | 2018 | R | East Dunbartonshire Council |  | 10 | 28.31 | 27.4 | 3.321% | G | 0.96786
Glasgow Scientific Services | 20% TEA in water | 2018 | R | East Dunbartonshire Council |  | 12 | 31.1333 | 28.6667 | 8.605% | P | 0.92077
Glasgow Scientific Services | 20% TEA in water | 2018 | R | East Dunbartonshire Council |  | 9 | 21.8963 | 18.6667 | 17.3% | P | 0.8525
Tayside Scientific Services | 20% TEA in water | 2018 | R | Fife Council |  | 12 | 26.0806 | 21.6958 | 20.21% | G | 0.83188
Tayside Scientific Services | 20% TEA in water | 2018 | R | Fife Council |  | 12 | 29.0139 | 21.4417 | 35.32% | G | 0.73901
Tayside Scientific Services | 20% TEA in water | 2018 | R | Fife Council |  | 11 | 21.503 | 16.9609 | 26.78% | G | 0.78877
Tayside Scientific Services | 20% TEA in water | 2018 | KS | Fife Council |  | 12 | 34.9653 | 25.6058 | 36.55% | G | 0.73232
Gradko | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 11 | 92.9948 | 85.0621 | 9.326% | G | 0.9147
Tayside Scientific Services | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 93.4444 | 85.0436 | 9.878% | G | 0.9101
Glasgow Scientific Services | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 86.1583 | 85.0436 | 1.311% | G | 0.98706
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 85.8081 | 85.0436 | 0.8989% | G | 0.99109
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 89.9917 | 85.0436 | 5.818% | G | 0.94502
Milton Keynes Council | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 84.2254 | 85.0436 | -0.9621% | G | 1.0097
Lambeth Scientific Services | 50% TEA in acetone | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 81.4222 | 85.0436 | -4.258% | G | 1.0445
Somerset County Council | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 90.4676 | 85.0436 | 6.378% | G | 0.94004
Staffordshire Scientific Services | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 84.4806 | 85.0436 | -0.6621% | G | 1.0067
Aberdeen Scientific Services | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 88.5567 | 85.0436 | 4.131% | G | 0.96033
Gradko | 50% TEA in acetone | 2018 | KS | Marylebone Road Intercomparison |  | 11 | 90.5936 | 85.0621 | 6.503% | G | 0.93894
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 99.5833 | 85.0436 | 17.1% | G | 0.85399
Gradko | 50% TEA in acetone | 2018 | UB | Reading Borough Council |  | 12 | 20.3151 | 26.2575 | -22.63% | G | 1.2925
Staffordshire Scientific Services | 20% TEA in water | 2018 | R | Salford City Council |  | 11 | 44.3515 | 40.1455 | 10.48% | G | 0.90517
Staffordshire Scientific Services | 20% TEA in water | 2018 | B | Salford City Council |  | 9 | 17.9574 | 14.4111 | 24.61% | G | 0.80252
Staffordshire Scientific Services | 20% TEA in water | 2018 | UB | Salford City Council |  | 12 | 29.6 | 25.0833 | 18.01% | G | 0.84741
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | UB | Sheffield City Council |  | 11 | 32.7273 | 27.9182 | 17.23% | S | 0.85306
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | UB | Sheffield City Council |  | 12 | 24.7222 | 22.8867 | 8.02% | G | 0.92575
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | R | Sheffield City Council |  | 12 | 32.75 | 36.5917 | -10.5% | S | 1.1173
Gradko | 20% TEA in water | 2018 | R | South Gloucestershire Council |  | 12 | 21.4464 | 20.1667 | 6.346% | G | 0.94033
Staffordshire Scientific Services | 20% TEA in water | 2018 | R | stockport |  | 12 | 43.2861 | 37.4583 | 15.56% | G | 0.86537
Socotec Didcot | 50% TEA in acetone | 2018 | R | Thanet District Council |  | 10 | 26.0817 | 20.8 | 25.39% | G | 0.7975
Gradko | 20% TEA in water | 2018 | R | Thurrock Borough Council |  | 12 | 53.2275 | 52.0066 | 2.348% | S | 0.97706
Gradko | 20% TEA in water | 2018 | R | Thurrock Borough Council |  | 12 | 34.0642 | 29.6053 | 15.06% | G | 0.86911
Gradko | 20% TEA in water | 2018 | R | Thurrock Borough Council |  | 12 | 30.6675 | 23.8144 | 28.78% | G | 0.77653
Gradko | 20% TEA in water | 2018 | UB | Thurrock Borough Council |  | 12 | 27.315 | 25.0212 | 9.167% | S | 0.91602
Glasgow Scientific Services | 20% TEA in water | 2018 | KS | Glasgow City Council |  | 12 | 63.7806 | 59.45 | 7.284% | P | 0.9321
Glasgow Scientific Services | 20% TEA in water | 2018 | R | Glasgow City Council |  | 11 | 44.3364 | 34 | 30.4% | P | 0.76686
Glasgow Scientific Services | 20% TEA in water | 2018 | R | Glasgow City Council |  | 12 | 32.3139 | 29.425 | 9.818% | P | 0.9106
Glasgow Scientific Services | 20% TEA in water | 2018 | R | Glasgow City Council |  | 10 | 36.2667 | 30.21 | 20.05% | P | 0.833
Glasgow Scientific Services | 20% TEA in water | 2018 | UB | Glasgow City Council |  | 11 | 32.5939 | 24.7364 | 31.77% | P | 0.75893
Socotec Didcot | 50% TEA in acetone | 2018 | R | Horsham District Council |  | 11 | 33.2258 | 23.3636 | 42.21% | G | 0.70318
Socotec Didcot | 50% TEA in acetone | 2018 | R | Horsham District Council |  | 12 | 33.4111 | 28.5 | 17.23% | G | 0.85301
Socotec Didcot | 50% TEA in acetone | 2018 | R | Horsham District Council |  | 12 | 30.3917 | 26.1667 | 16.15% | G | 0.86098
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | KS | Leeds City Council |  | 12 | 39.6667 | 32.7083 | 21.27% | S | 0.82458
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | R | Leeds City Council |  | 12 | 34.3333 | 27.8083 | 23.46% | S | 0.80995
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | UC | Leeds City Council |  | 12 | 38.25 | 29.8 | 28.36% | G | 0.77908
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | R | Leeds City Council |  | 11 | 49.8182 | 41.2364 | 20.81% | S | 0.82774
Staffordshire Scientific Services | 20% TEA in water | 2018 | KS | Oldham Council |  | 12 | 35.9042 | 28.9697 | 23.94% | G | 0.80686
Lambeth Scientific Services | 50% TEA in acetone | 2018 | SU | Reigate and Banstead BC |  | 12 | 23.7917 | 24.997 | -4.822% | G | 1.0507
Lambeth Scientific Services | 50% TEA in acetone | 2018 | SU | Reigate and Banstead BC |  | 12 | 21.6389 | 18.9577 | 14.14% | G | 0.87609
Lambeth Scientific Services | 50% TEA in acetone | 2018 | B | Reigate and Banstead BC |  | 12 | 15.8611 | 15.8088 | 0.3311% | P | 0.9967
Lambeth Scientific Services | 50% TEA in acetone | 2018 | R | Reigate and Banstead BC (Note tubes set up at Croydon 7 site) |  | 10 | 29.5333 | 30.8359 | -4.224% | G | 1.0441
Socotec Didcot | 50% TEA in acetone | 2018 | UB | Slough Borough Council |  | 10 | 38.3233 | 30.518 | 25.58% | G | 0.79633
Socotec Didcot | 50% TEA in acetone | 2018 | SU | Slough Borough Council |  | 11 | 31.953 | 21.7801 | 46.71% | G | 0.68163
Socotec Didcot | 50% TEA in acetone | 2018 | R | Slough Borough Council |  | 11 | 39.3485 | 32.1318 | 22.46% | G | 0.8166
Staffordshire Scientific Services | 20% TEA in water | 2018 | R | Stoke-on-Trent City Council |  | 10 | 57.2933 | 55.7967 | 2.682% | G | 0.97388
Staffordshire Scientific Services | 20% TEA in water | 2018 | UB | Stoke-on-Trent City Council |  | 11 | 27.7515 | 23.3015 | 19.1% | G | 0.83965
Socotec Didcot | 50% TEA in acetone | 2018 | R | Vale of Glamorgan |  | 12 | 39.0694 | 24.7667 | 57.75% | G | 0.63391
Staffordshire Scientific Services | 20% TEA in water | 2018 | UB | Wigan Council |  | 10 | 26.305 | 16.308 | 61.3% | G | 0.61996
Socotec Didcot | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 96.0583 | 85.0436 | 12.95% | G | 0.88533
Socotec Glasgow | 20% TEA in water | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 91.3306 | 85.0436 | 7.393% | G | 0.93116
Socotec Glasgow | 50% TEA in acetone | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 88.025 | 85.0436 | 3.506% | G | 0.96613
Socotec Didcot | 50% TEA in acetone | 2018 | KS | Marylebone Road Intercomparison |  | 12 | 94.6389 | 85.0436 | 11.28% | G | 0.89861
Lambeth Scientific Services | 50% TEA in acetone | 2018 | R | Elmbridge Borough Council |  | 12 | 28.6667 | 32.5 | -11.79% | G | 1.1337
Lambeth Scientific Services | 50% TEA in acetone | 2018 | R | Elmbridge Borough Council |  | 12 | 33.3611 | 37.5833 | -11.23% | G | 1.1266
Gradko | 20% TEA in water | 2018 | UC | Belfast City Council |  | 12 | 31.5519 | 27.1158 | 16.36% | G | 0.8594
Gradko | 20% TEA in water | 2018 | R | City of Lincoln Council |  | 12 | 44.2706 | 33.5058 | 32.13% | G | 0.75684
Staffordshire Scientific Services | 20% TEA in water | 2018 | R | East Staffordshire Borough Council |  | 12 | 44.1778 | 39.8833 | 10.77% | G | 0.90279
Gradko | 50% TEA in acetone | 2018 | UI | Falkirk Council |  | 12 | 21.3464 | 17.8333 | 19.7% | G | 0.83543
Gradko | 50% TEA in acetone | 2018 | UC | Falkirk Council |  | 10 | 41.4957 | 41.2 | 0.7176% | G | 0.99287
Socotec Didcot | 50% TEA in acetone | 2018 | B | Gravesham Borough Council |  | 12 | 36.625 | 29.9917 | 22.12% | G | 0.81889
Socotec Didcot | 50% TEA in acetone | 2018 | B | Gravesham Borough Council |  | 12 | 28.1931 | 23.7333 | 18.79% | G | 0.84181
Gradko | 20% TEA in water | 2018 | R | Lancaster City Council |  | 11 | 38.9463 | 34.6364 | 12.44% | G | 0.88934
Gradko | 20% TEA in water | 2018 | R | Lancaster City Council |  | 11 | 30.9391 | 33.8273 | -8.538% | G | 1.0933
Gradko | 20% TEA in water | 2018 | UB | Liverpool City Council |  | 12 | 19.6451 | 17.7 | 10.99% | G | 0.90099
Gradko | 50% TEA in acetone | 2018 | R | London Borough of Croydon |  | 12 | 55.2247 | 41.3942 | 33.41% | G | 0.74956
Gradko | 50% TEA in acetone | 2018 | KS | London Borough of Croydon |  | 12 | 56.4239 | 48.9767 | 15.21% | P | 0.86801
Gradko | 50% TEA in acetone | 2018 | UI | Middlesbrough |  | 12 | 20.8247 | 14.5 | 43.62% | G | 0.69629
Socotec Didcot | 50% TEA in acetone | 2018 | UI | North Lincolnshire Council |  | 12 | 23.8028 | 15.5083 | 53.48% | G | 0.65153
Gradko | 50% TEA in acetone | 2018 | UB | Norwich City Council |  | 12 | 13.9267 | 11.5667 | 20.4% | G | 0.83054
Gradko | 50% TEA in Acetone | 2018 | R | Sandwell Metropolitan Borough Council |  | 12 | 43.7149 | 40.0792 | 9.071% | G | 0.91683
Gradko | 50% TEA in Acetone | 2018 | UB | Sandwell Metropolitan Borough Council |  | 11 | 16.49 | 15.1091 | 9.14% | S | 0.91626
Gradko | 50% TEA in Acetone | 2018 | B | Sandwell Metropolitan Borough Council |  | 12 | 25.8544 | 24.925 | 3.729% | G | 0.96405
Gradko | 50% TEA in Acetone | 2018 | R | Sandwell Metropolitan Borough Council |  | 10 | 32.531 | 31.365 | 3.718% | S | 0.96416
Staffordshire Scientific Services | 20% TEA in water | 2018 | O | South Staffordshire Council |  | 12 | 32.4111 | 34.8942 | -7.116% | G | 1.0766
Socotec Didcot | 50% TEA in acetone | 2018 | R | Swansea Council |  | 12 | 33.1042 | 23.8108 | 39.03% | G | 0.71927
Socotec Didcot | 50% TEA in acetone | 2018 | UB | Swansea Council |  | 10 | 19.3283 | 15.666 | 23.38% | G | 0.81052
Gradko | 20% TEA in water | 2018 | R | Blackburn with Darwen Borogh Council |  | 12 | 25.6692 | 19.9308 | 28.79% | G | 0.77645
Socotec Didcot | 20% TEA in water | 2018 | R | Rhondda Cynon Taf CBC |  | 12 | 30.0222 | 25.48 | 17.83% | G | 0.8487
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | R | Barnsley MBC |  | 9 | 34.0741 | 32.3545 | 5.315% | G | 0.94953
Glasgow Scientific Services | 20% TEA in water | 2018 | R | East Ayrshire Council |  | 10 | 26.48 | 30.15 | -12.17% | P | 1.1386
Socotec Didcot | 50% TEA in Acetone | 2018 | R | Wrexham County Borough Council |  | 11 | 21.0061 | 18.0909 | 16.11% | G | 0.86122
Gradko | 20% TEA in water | 2018 | R | Dartford Borough Council |  | 11 | 50.1015 | 48.0455 | 4.279% | G | 0.95896
Socotec Didcot | 50% TEA in acetone | 2018 | R | Sevenoaks District Council |  | 12 | 33.6389 | 24.9583 | 34.78% | G | 0.74195
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | UB | BCP Council (Bournemmouth Borough Council) |  | 11 | 13.4545 | 12.0134 | 12% | P | 0.89289
Somerset County Council | 20% TEA in water | 2018 | KS | Bristol City Council |  | 9 | 39.8111 | 31.7778 | 25.28% | G | 0.79821
Somerset County Council | 20% TEA in water | 2018 | R | Bristol City Council |  | 11 | 46.7015 | 44.6364 | 4.627% | G | 0.95578
Somerset County Council | 20% TEA in water | 2018 | R | Bristol City Council |  | 12 | 41.0028 | 39.75 | 3.152% | G | 0.96945
Somerset County Council | 20% TEA in water | 2018 | R | Bristol City Council |  | 11 | 37.5955 | 41.5455 | -9.508% | G | 1.1051
Somerset County Council | 20% TEA in water | 2018 | UB | Bristol City Council |  | 9 | 28.2296 | 24.7778 | 13.93% | G | 0.87772
Somerset County Council | 20% TEA in water | 2018 | UB | Bristol City Council |  | 12 | 30.3389 | 25.8333 | 17.44% | G | 0.85149
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | R | City of Edinburgh Council |  | 10 | 34.895 | 25.59 | 36.36% | G | 0.73334
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | R | City of Edinburgh Council |  | 12 | 31.7417 | 27.9083 | 13.74% | G | 0.87923
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | R | City of Edinburgh Council |  | 12 | 26.5986 | 24.8417 | 7.073% | G | 0.93395
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | KS | City of Edinburgh Council |  | 12 | 46.8528 | 42.8833 | 9.256% | G | 0.91528
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | R | City of Edinburgh Council |  | 11 | 55.2833 | 52.2 | 5.907% | G | 0.94423
Socotec Didcot | 50% TEA in acetone | 2018 | R | Derry City and Strabane District Council |  | 12 | 37.2306 | 31.9583 | 16.5% | G | 0.85839
Socotec Didcot | 20% TEA in water | 2018 | R | New Forest District Council |  | 12 | 36.2639 | 26.2339 | 38.23% | G | 0.72342
Socotec Didcot | 50% TEA in acetone | 2018 | R | City of Wolverhampton Council |  | 12 | 39.7722 | 27.8333 | 42.89% | G | 0.69982
Staffordshire Scientific Services | 20% TEA in water | 2018 | R | Tameside MBC |  | 10 | 63.5967 | 42.8 | 48.59% | G | 0.67299
Gradko | 20% TEA in water | 2018 | UB | Dudley MBC |  | 11 | 24.1061 | 18.6455 | 29.29% | G | 0.77348
Gradko | 20% TEA in water | 2018 | R | Dudley MBC |  | 12 | 36.7353 | 35.4583 | 3.601% | G | 0.96524
Gradko | 20% TEA in water | 2018 | R | Dudley MBC |  | 11 | 46.7094 | 45.8273 | 1.925% | G | 0.98111
Aberdeen Scientific Services | 20% TEA in water | 2019 | UB | Aberdeen City Council |  | 13 | 23.5128 | 17.2308 | 36.46% | G | 0.73282
Aberdeen Scientific Services | 20% TEA in water | 2019 | R | Aberdeen City Council |  | 12 | 23.8333 | 19.6667 | 21.19% | G | 0.82517
Aberdeen Scientific Services | 20% TEA in water | 2019 | R | Aberdeen City Council |  | 13 | 40.0513 | 31.7692 | 26.07% | G | 0.79321
Aberdeen Scientific Services | 20% TEA in water | 2019 | R | Aberdeen City Council |  | 13 | 43.7179 | 35.6923 | 22.49% | G | 0.81642
Aberdeen Scientific Services | 20% TEA in water | 2019 | R | Aberdeen City Council |  | 13 | 41.6154 | 35.2308 | 18.12% | G | 0.84658
Gradko | 20% TEA in water | 2019 | R | Blackburn with darwen Borough Council |  | 10 | 29.0831 | 21.2464 | 36.89% | G | 0.73054
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | Cannock Chase Council |  | 13 | 24.4077 | 21.1077 | 15.63% | G | 0.8648
Gradko | 20% TEA in water | 2019 | R | Cheshire West and Chester |  | 12 | 39.0996 | 38.3333 | 1.999% | G | 0.9804
Gradko | 20% TEA in water | 2019 | R | Cheshire West and Chester |  | 11 | 33.6808 | 34.4091 | -2.117% | G | 1.0216
Gradko | 50% TEA in acetone | 2019 | R | City of London |  | 12 | 73.9536 | 71.025 | 4.123% | G | 0.9604
Gradko | 50% TEA in acetone | 2019 | UB | City of London |  | 12 | 37.2185 | 32.575 | 14.25% | G | 0.87524
Glasgow Scientific Services | 20% TEA in water | 2019 | R | East Dunbartonshire Council |  | 12 | 35.9833 | 31.8333 | 13.04% | P | 0.88467
Glasgow Scientific Services | 20% TEA in water | 2019 | R | East Dunbartonshire Council |  | 12 | 32.6722 | 26.5833 | 22.9% | G | 0.81364
Glasgow Scientific Services | 20% TEA in water | 2019 | R | East Dunbartonshire Council |  | 11 | 23.1364 | 19.3636 | 19.48% | P | 0.83694
Gradko | 20% TEA in water | 2019 | R | Gedling Borough Council |  | 12 | 31.9806 | 29.8 | 7.317% | G | 0.93182
Socotec Didcot | 50% TEA in acetone | 2019 | UB | Kingston upon Hull City Council |  | 12 | 29.7833 | 22.525 | 32.22% | G | 0.7563
Socotec Didcot | 50% TEA in acetone | 2019 | O | Kingston upon Hull City Council |  | 11 | 31.5091 | 26.4545 | 19.11% | G | 0.83958
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | Leicester City Council |  | 11 | 51.0273 | 43.8818 | 16.28% | G | 0.85997
Milton Keynes Council | 20% TEA in water | 2019 | R | Milton Keynes Council |  | 11 | 32.065 | 27.6517 | 15.96% | G | 0.86236
Gradko | 20% TEA in water | 2019 | R | NOTTINGHAM CITY COUNCIL |  | 10 | 37.045 | 39.82 | -6.969% | G | 1.0749
Staffordshire Scientific Services | 20% TEA in water | 2019 | UB | Salford City Council |  | 12 | 27.1472 | 25.2189 | 7.646% | G | 0.92897
Staffordshire Scientific Services | 20% TEA in water | 2019 | B | Salford City Council |  | 11 | 15.9333 | 14.7716 | 7.865% | G | 0.92709
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | Salford City Council |  | 12 | 44.4958 | 43.5694 | 2.126% | G | 0.97918
Socotec Didcot | 50% TEA in acetone | 2019 | R | Vale of Glamorgan |  | 11 | 39.7939 | 23.69 | 67.98% | G | 0.59532
Socotec Didcot | 50% TEA in acetone | 2019 | R | Watford Borough Council |  | 12 | 35.025 | 30.2167 | 15.91% | S | 0.86272
Gradko | 20% TEA in water | 2019 | R | Bedford Borough Council |  | 11 | 28.6818 | 28.9727 | -1.004% | G | 1.0101
Gradko | 20% TEA in water | 2019 | R | Bedford Borough Council |  | 12 | 36.5972 | 32.3917 | 12.98% | G | 0.88509
Socotec Didcot | 50% TEA in acetone | 2019 | R | Dumfries & Galloway Council |  | 13 | 35.0795 | 31.2431 | 12.28% | G | 0.89064
Socotec Glasgow | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 82.7639 | 65.3983 | 26.55% | G | 0.79018
Socotec Glasgow | 50% TEA in acetone | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 86.7556 | 65.3983 | 32.66% | G | 0.75382
Socotec Didcot | 50% TEA in acetone | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 91.8861 | 65.3983 | 40.5% | G | 0.71173
Socotec Didcot | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 93.5972 | 65.3983 | 43.12% | G | 0.69872
Milton Keynes Council | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 79.179 | 65.3983 | 21.07% | G | 0.82595
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 84.1722 | 65.3983 | 28.71% | G | 0.77696
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 75.4612 | 65.3983 | 15.39% | P | 0.86665
Glasgow Scientific Services | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 79.1528 | 65.3983 | 21.03% | G | 0.82623
West Yorkshire Analytical Services | 50% TEA in acetone | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 79.5722 | 65.3983 | 21.67% | G | 0.82187
Gradko | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 85.0767 | 65.3983 | 30.09% | G | 0.7687
Gradko | 50% TEA in acetone | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 82.6097 | 65.3983 | 26.32% | G | 0.79165
Aberdeen Scientific Services | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 77.3603 | 65.3983 | 18.29% | G | 0.84537
Staffordshire Scientific Services | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 75.0306 | 65.3983 | 14.73% | G | 0.87162
Tayside Scientific Services | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 81.9639 | 65.3983 | 25.33% | G | 0.79789
Somerset County Council | 20% TEA in water | 2019 | KS | Marylebone Road Intercomparison |  | 12 | 83.1 | 65.3983 | 27.07% | G | 0.78698
Lambeth Scientific Services | 50% TEA in acetone | 2019 | KS | Marylebone Road Intercomparison |  | 11 | 77.7333 | 65.9313 | 17.9% | G | 0.84817
Gradko | 20% TEA in water | 2019 | R | Borough Council of King's Lynn and West Norfolk |  | 9 | 27.2037 | 21.1889 | 28.39% | G | 0.7789
Socotec Didcot | 50% TEA in acetone | 2019 | UB | City of York Council |  | 12 | 21.9639 | 16.1942 | 35.63% | G | 0.73731
Socotec Didcot | 50% TEA in acetone | 2019 | R | City of York Council |  | 12 | 32.775 | 25.8408 | 26.83% | G | 0.78843
Socotec Didcot | 50% TEA in acetone | 2019 | R | City of York Council |  | 9 | 31.6593 | 23.0722 | 37.22% | G | 0.72877
Socotec Didcot | 50% TEA in acetone | 2019 | R | City of York Council |  | 11 | 39.7788 | 27.7455 | 43.37% | G | 0.69749
Socotec Didcot | 20% TEA in water | 2019 | KS | Fife Council |  | 12 | 30.4528 | 23.1417 | 31.59% | G | 0.75992
Socotec Didcot | 20% TEA in water | 2019 | R | Fife Council |  | 12 | 27.2583 | 20.7667 | 31.26% | G | 0.76185
Socotec Didcot | 20% TEA in water | 2019 | R | Fife Council |  | 12 | 24.8194 | 16.2917 | 52.34% | G | 0.65641
Socotec Didcot | 20% TEA in water | 2019 | R | Fife Council |  | 12 | 25.8389 | 21.75 | 18.8% | G | 0.84175
Socotec Didcot | 50% TEA in acetone | 2019 | R | Ipswich Boorough council |  | 11 | 34.2606 | 25.5445 | 34.12% | G | 0.7456
Gradko | 20% TEA in water | 2019 | R | Lancaster City Council |  | 13 | 39.6727 | 34.0692 | 16.45% | G | 0.85876
Gradko | 20% TEA in water | 2019 | R | Lancaster City Council |  | 12 | 31.3579 | 30.85 | 1.646% | G | 0.9838
Gradko | 20% TEA in Water | 2019 | R | Monmouthshire County Council |  | 12 | 39.3214 | 38.8058 | 1.329% | G | 0.98689
Staffordshire Scientific Services | 20% TEA in water | 2019 | UB | Stoke-on-Trent City Council |  | 12 | 26.5306 | 23.637 | 12.24% | G | 0.89093
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | Stoke-on-Trent City Council |  | 10 | 55.3467 | 63.3048 | -12.57% | G | 1.1438
Socotec Didcot | 50% TEA in acetone | 2019 | R | Swale BC |  | 12 | 51.4722 | 39.0917 | 31.67% | G | 0.75947
Socotec Didcot | 50% TEA in acetone | 2019 | R | Swale BC |  | 12 | 33.3472 | 26.925 | 23.85% | G | 0.80741
Socotec Didcot | 50% TEA in acetone | 2019 | R | Swale BC |  | 12 | 39.6944 | 31.3417 | 26.65% | G | 0.78957
Socotec Didcot | 50% TEA in acetone | 2019 | R | Wrexham County Borough Council |  | 10 | 19.81 | 16.21 | 22.21% | G | 0.81827
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | Bury Council |  | 11 | 27.1455 | 25.9236 | 4.713% | G | 0.95499
Socotec Didcot | 50% TEA in acetone | 2019 | R | City of Wolverhampton Council |  | 12 | 39.325 | 26.5 | 48.4% | G | 0.67387
Gradko | 20% TEA in water | 2019 | R | Dudley MBC |  | 12 | 33.3375 | 31.9 | 4.506% | G | 0.95688
Gradko | 20% TEA in water | 2019 | R | Dudley MBC |  | 12 | 43.6731 | 42.0417 | 3.88% | G | 0.96265
Gradko | 20% TEA in water | 2019 | UB | Dudley MBC |  | 12 | 22.8744 | 19.0917 | 19.81% | G | 0.83463
Gradko | 20% TEA in water | 2019 | UB | Eastleigh Borough Council |  | 12 | 24.3858 | 26.2524 | -7.11% | G | 1.0765
Gradko | 20% TEA in water | 2019 | R | Gateshead Council |  | 12 | 33.5589 | 27.1317 | 23.69% | P | 0.80848
Gradko | 20% TEA in water | 2019 | R | Gateshead Council |  | 11 | 39.5512 | 44.1727 | -10.46% | G | 1.1168
Gradko | 20% TEA in water | 2019 | R | Gateshead Council |  | 10 | 31.8813 | 34.348 | -7.181% | G | 1.0774
Gradko | 20% TEA in water | 2019 | R | Gateshead Council |  | 12 | 29.8472 | 25.2658 | 18.13% | G | 0.84651
Gradko | 50% TEA in acetone | 2019 | R | London Borough of Richmond upon Thames |  | 12 | 46.1922 | 35.4333 | 30.36% | G | 0.76708
Gradko | 50% TEA in acetone | 2019 | R | London Borough of Richmond upon Thames |  | 12 | 28.9392 | 27.0333 | 7.05% | G | 0.93414
Gradko | 50% TEA in acetone | 2019 | B | London Borough of Richmond upon Thames |  | 11 | 21.3749 | 21.1636 | 0.9984% | G | 0.99011
Staffordshire Scientific Services | 20% TEA in water | 2019 | KS | Manchester City Council |  | 12 | 55.175 | 58.9 | -6.324% | G | 1.0675
Staffordshire Scientific Services | 20% TEA in water | 2019 | UC | Manchester City Council |  | 12 | 34.8694 | 36.2708 | -3.864% | G | 1.0402
Staffordshire Scientific Services | 20% TEA in water | 2019 | SI | Manchester City Council |  | 12 | 20.2583 | 22.2792 | -9.071% | G | 1.0998
Socotec Didcot | 50% TEA in acetone | 2019 | R | North Herts DC |  | 12 | 59.3389 | 46.1917 | 28.46% | G | 0.77844
Socotec Didcot | 20% TEA in water | 2019 | R | Rhondda Cynon Taf CBC |  | 11 | 31.0879 | 25.9035 | 20.01% | G | 0.83323
Somerset County Council | 20% TEA in water | 2019 | R | South Gloucestershire Council |  | 11 | 24.871 | 19.0727 | 30.4% | G | 0.76687
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | Stockport MBC |  | 12 | 40.0444 | 35.509 | 12.77% | G | 0.88674
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | Stockport MBC |  | 11 | 26.5121 | 23.0608 | 14.97% | G | 0.86982
Gradko | 20% TEA in water | 2019 | R | Thurrock Borough Council |  | 12 | 29.0692 | 23.9029 | 21.61% | G | 0.82228
Staffordshire Scientific Services | 20% TEA in water | 2019 | UB | Trafford |  | 10 | 22.03 | 18.66 | 18.06% | G | 0.84703
Staffordshire Scientific Services | 20% TEA in water | 2019 | UB | Wigan Council |  | 11 | 24.2424 | 18.0795 | 34.09% | G | 0.74578
Gradko | 20% TEA in water | 2019 | R | Brighton & Hove City Council |  | 11 | 45.4091 | 46.0091 | -1.304% | G | 1.0132
Staffordshire Scientific Services | 20% TEA in water | 2019 | R | East Staffordshire Borough Council |  | 11 | 40.9606 | 35.9091 | 14.07% | G | 0.87667
Gradko | 50% TEA in acetone | 2019 | UB | Falkirk Council |  | 9 | 18.1363 | 15.3556 | 18.11% | G | 0.84668
Socotec Didcot | 50% TEA in acetone | 2019 | R | Horsham District Council |  | 12 | 29.8722 | 23.9942 | 24.5% | G | 0.80323
Socotec Didcot | 50% TEA in acetone | 2019 | R | Horsham District Council |  | 11 | 31.1818 | 21.5336 | 44.81% | G | 0.69058
Socotec Didcot | 50% TEA in acetone | 2019 | R | Horsham District Council |  | 11 | 32.0939 | 23.8827 | 34.38% | G | 0.74415
Gradko | 50% TEA in acetone | 2019 | R | LB Newham |  | 12 | 34.6107 | 29.7963 | 16.16% | G | 0.8609
West Yorkshire Analytical Services | 50% TEA in acetone | 2019 | KS | Leeds City Council |  | 10 | 39.6 | 27.97 | 41.58% | S | 0.70631
West Yorkshire Analytical Services | 50% TEA in acetone | 2019 | UC | Leeds City Council |  | 12 | 32.6667 | 28.325 | 15.33% | P | 0.86709
Socotec Didcot | 50% TEA in acetone | 2019 | B | Medway Council |  | 10 | 21.1017 | 13.23 | 59.5% | P | 0.62696
Socotec Didcot | 50% TEA in acetone | 2019 | R | Medway Council |  | 12 | 32.6819 | 24.1833 | 35.14% | G | 0.73996
Socotec Didcot | 50% TEA in acetone | 2019 | R | Waverley Borough Council |  | 10 | 38.0367 | 29.832 | 27.5% | G | 0.7843
Socotec Didcot | 50% TEA in acetone | 2019 | R | Waverley Borough Council |  | 12 | 34.7583 | 24.0292 | 44.65% | G | 0.69132
Gradko | 20% TEA in water | 2019 | R | Belfast City Council |  | 12 | 40.3781 | 33.3667 | 21.01% | G | 0.82636
Gradko | 20% TEA in water | 2019 | R | Belfast City Council |  | 12 | 44.0172 | 44.9908 | -2.164% | G | 1.0221
Gradko | 20% TEA in water | 2019 | R | Belfast City Council |  | 12 | 27.9203 | 26.4983 | 5.366% | G | 0.94907
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2019 | R | Barnsley MBC |  | 11 | 33.4848 | 32.9245 | 1.702% | G | 0.98327
Gradko | 20% TEA in water | 2019 | UB | Southampton City Council |  | 12 | 30.3192 | 27.9221 | 8.585% | G | 0.92094
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2019 | SU | Blaby District Council |  | 10 | 23 | 30.59 | -24.81% | G | 1.33
Socotec Didcot | 50% TEA in acetone | 2019 | KS | Caerphilly CBC |  | 12 | 90.3889 | 63.4628 | 42.43% | G | 0.70211
Socotec Didcot | 50% TEA in acetone | 2019 | R | Caerphilly CBC |  | 11 | 41.7273 | 27.0866 | 54.05% | G | 0.64914
Socotec Didcot | 50% TEA in acetone | 2019 | KS | Caerphilly CBC |  | 11 | 34.4848 | 24.3642 | 41.54% | G | 0.70652
Socotec Didcot | 50% TEA in acetone | 2019 | R | Cambridge City Council |  | 11 | 41.6576 | 28.3187 | 47.1% | G | 0.6798
Gradko | 20% TEA in water | 2019 | UB | Liverpool City Council |  | 12 | 19.7769 | 19.45 | 1.681% | G | 0.98347
Gradko | 50% TEA in acetone | 2019 | UB | Middlesbrough |  | 12 | 19.7256 | 16.1167 | 22.39% | G | 0.81705
Gradko | 50% TEA in acetone | 2019 | UB | Norwich City Council |  | 12 | 14.0992 | 13.1083 | 7.559% | G | 0.92972
Gradko | 50% TEA in acetone | 2019 | R | RBWM |  | 12 | 34.2649 | 32.4 | 5.756% | G | 0.94558
Gradko | 50% TEA in acetone | 2019 | R | RBWM |  | 12 | 37.5578 | 35.225 | 6.622% | G | 0.93789
Lambeth Scientific Services | 50% TEA in acetone | 2019 | UB | Spelthorne Borough Council |  | 12 | 32.4167 | 33.36 | -2.828% | G | 1.0291
Lambeth Scientific Services | 50% TEA in acetone | 2019 | UB | Spelthorne Borough Council |  | 9 | 29.5556 | 27.0495 | 9.265% | G | 0.91521
Somerset County Council | 20% TEA in water | 2019 | R | Wiltshire Council |  | 13 | 36.0769 | 31.1246 | 15.91% | G | 0.86273
Somerset County Council | 20% TEA in water | 2019 | UB | Wiltshire Council |  | 13 | 33.6154 | 30.2594 | 11.09% | G | 0.90016
Lambeth Scientific Services | 50% TEA in acetone | 2019 | R | Elmbridge Borough Council |  | 12 | 39.0833 | 41.175 | -5.08% | G | 1.0535
Lambeth Scientific Services | 50% TEA in acetone | 2019 | R | Elmbridge Borough Council |  | 12 | 32.9722 | 30.9167 | 6.649% | G | 0.93766
Lambeth Scientific Services | 50% TEA in acetone | 2019 | B | Reigate and Banstead BC (RG3) |  | 12 | 16.3056 | 15.0932 | 8.033% | G | 0.92565
Lambeth Scientific Services | 50% TEA in acetone | 2019 | SU | Reigate and Banstead BC (RG1) |  | 12 | 24.75 | 19.1626 | 29.16% | G | 0.77425
Lambeth Scientific Services | 50% TEA in acetone | 2019 | SU | Reigate and Banstead BC RG2(6) |  | 12 | 26.9167 | 24.3061 | 10.74% | G | 0.90301
Gradko | 20% TEA in water | 2019 | R | Ards and North Down Borough Council |  | 12 | 33.3603 | 25.45 | 31.08% | G | 0.76288
Gradko | 50% TEA in acetone | 2019 | UB | Bureau Veritas |  | 11 | 38.3871 | 31.5628 | 21.62% | G | 0.82222
Gradko | 50% TEA in acetone | 2019 | KS | Croydon |  | 11 | 50.3164 | 43.3065 | 16.19% | G | 0.86068
Gradko | 50% TEA in acetone | 2019 | R | Croydon |  | 11 | 53.3603 | 44.288 | 20.48% | G | 0.82998
Gradko | 20% TEA in water | 2019 | R | Eastleigh Borough Council |  | 12 | 24.8787 | 25.7248 | -3.289% | G | 1.034
Glasgow Scientific Services | 20% TEA in water | 2019 | R | Glasgow City Council |  | 12 | 43.6778 | 34.675 | 25.96% | G | 0.79388
Glasgow Scientific Services | 20% TEA in water | 2019 | R | Glasgow City Council |  | 10 | 32.6083 | 29.58 | 10.24% | G | 0.90713
Glasgow Scientific Services | 20% TEA in water | 2019 | R | Glasgow City Council |  | 12 | 36.3139 | 29.5833 | 22.75% | G | 0.81466
Glasgow Scientific Services | 20% TEA in water | 2019 | KS | Glasgow City Council |  | 12 | 64.8708 | 55.9 | 16.05% | G | 0.86171
Glasgow Scientific Services | 20% TEA in water | 2019 | UB | Glasgow City Council |  | 12 | 28.5556 | 24.0917 | 18.53% | G | 0.84368
Gradko | 50% TEA in acetone | 2019 | SU | Greenwich |  | 12 | 20.2167 | 17.2721 | 17.05% | G | 0.85435
Gradko | 50% TEA in acetone | 2019 | R | Greenwich |  | 12 | 60.1133 | 52.5986 | 14.29% | G | 0.87499
Gradko | 50% TEA in acetone | 2019 | R | Greenwich |  | 12 | 44.3364 | 35.574 | 24.63% | G | 0.80236
Gradko | 50% TEA in acetone | 2019 | R | Greenwich |  | 12 | 40.7304 | 38.5032 | 5.784% | G | 0.94532
Gradko | 50% TEA in acetone | 2019 | R | Greenwich |  | 12 | 37.4853 | 33.8431 | 10.76% | G | 0.90284
Gradko | 50% TEA in acetone | 2019 | R | Greenwich |  | 12 | 32.6573 | 32.6259 | 0.09613% | G | 0.99904
Gradko | 50% TEA in acetone | 2019 | R | Greenwich |  | 12 | 36.458 | 33.2184 | 9.752% | G | 0.91114
Gradko | 50% TEA in acetone | 2019 | R | Hammersmith and Fulham |  | 11 | 57.001 | 58.3528 | -2.317% | G | 1.0237
Gradko | 20% TEA in water | 2019 | R | Lisburn & Castlereagh City Council |  | 12 | 28.2342 | 22 | 28.34% | G | 0.7792
Socotec Didcot | 20% TEA in water | 2019 | KS | New Forest DC |  | 11 | 46.4394 | 31.6627 | 46.67% | G | 0.68181
Gradko | 50% TEA in acetone | 2019 | R | Newham |  | 12 | 34.6107 | 29.7498 | 16.34% | G | 0.85955
Gradko | 50% TEA in acetone | 2019 | UB | Royal Borough of Kensington and Chelsea |  | 12 | 26.7945 | 27.4035 | -2.222% | G | 1.0227
Socotec Didcot | 20% TEA in water | 2019 | R | South Oxfordshire Distric Council |  | 12 | 32.8194 | 28.2475 | 16.19% | G | 0.86069
Socotec Didcot | 20% TEA in water | 2019 | R | South Oxfordshire Distric Council |  | 11 | 41.6788 | 34.8964 | 19.44% | G | 0.83727
Somerset County Council | 20% TEA in water | 2019 | UB | Bristol City Council |  | 11 | 29.2006 | 25 | 16.8% | G | 0.85615
Somerset County Council | 20% TEA in water | 2019 | R | Bristol City Council |  | 12 | 83.6338 | 66.3333 | 26.08% | G | 0.79314
Somerset County Council | 20% TEA in water | 2019 | R | Bristol City Council |  | 12 | 38.7049 | 30.8333 | 25.53% | G | 0.79663
Somerset County Council | 20% TEA in water | 2019 | R | Bristol City Council |  | 12 | 42.0464 | 38.85 | 8.228% | G | 0.92398
Somerset County Council | 20% TEA in water | 2019 | KS | Bristol City Council |  | 12 | 38.8297 | 30.25 | 28.36% | G | 0.77904
Socotec Didcot | 50% TEA in acetone | 2019 | UB | Canterbury City Council |  | 12 | 15.7028 | 12.3083 | 27.58% | G | 0.78383
Socotec Didcot | 50% TEA in acetone | 2019 | R | Canterbury City Council |  | 12 | 34.2167 | 25.25 | 35.51% | G | 0.73794
Socotec Didcot | 50% TEA in acetone | 2019 | R | Dacorum Borough Council |  | 11 | 31.247 | 24 | 30.2% | G | 0.76807
Socotec Didcot | 50% TEA in acetone | 2019 | R | Derry City and Strabane District Council |  | 12 | 38.7942 | 32.3042 | 20.09% | G | 0.83271
Socotec Didcot | 50% TEA in acetone | 2019 | UB | Derry City and Strabane District Council |  | 12 | 14.9029 | 10.615 | 40.39% | G | 0.71228
Glasgow Scientific Services | 20% TEA in water | 2019 | R | East Ayrshire Council |  | 12 | 25.9056 | 24.2317 | 6.908% | G | 0.93538
Glasgow Scientific Services | 20% TEA in water | 2019 | R | East Dunbartonshire Council |  | 12 | 25.6194 | 26.25 | -2.402% | G | 1.0246
Socotec Didcot | 50% TEA in acetone | 2019 | B | Gravesham Borough Council |  | 12 | 36.3861 | 29.2167 | 24.54% | G | 0.80296
Socotec Didcot | 50% TEA in acetone | 2019 | B | Gravesham Borough Council |  | 12 | 27.2694 | 24.5833 | 10.93% | G | 0.9015
Lambeth Scientific Services | 50% TEA in acetone | 2019 | R | London Borough of Islington |  | 12 | 45.5833 | 40.1 | 13.67% | P | 0.87971
Socotec Didcot | 50% TEA in acetone | 2019 | R | Slough Borough Council |  | 11 | 39.3485 | 32.1318 | 22.46% | G | 0.8166
Socotec Didcot | 50% TEA in acetone | 2019 | SU | Slough Borough Council |  | 11 | 31.953 | 21.7801 | 46.71% | G | 0.68163
Socotec Didcot | 50% TEA in acetone | 2019 | UB | Slough Borough Council |  | 10 | 38.3233 | 30.518 | 25.58% | G | 0.79633
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | R | City of Edinburgh Council |  | 12 | 32.2986 | 25.1583 | 28.38% | G | 0.77893
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | R | City of Edinburgh Council |  | 9 | 31.2444 | 27.2667 | 14.59% | G | 0.87269
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | R | City of Edinburgh Council |  | 12 | 44.8208 | 36.9667 | 21.25% | P | 0.82477
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | R | City of Edinburgh Council |  | 12 | 27.8264 | 24.1833 | 15.06% | G | 0.86908
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | KS | City of Edinburgh Council |  | 12 | 46.6333 | 41.85 | 11.43% | G | 0.89743
Socotec Didcot | 20% TEA in water | 2019 | R | London Borough of Ealing |  | 10 | 84.0167 | 64.2954 | 30.67% | G | 0.76527
Socotec Didcot | 20% TEA in water | 2019 | R | London Borough of ealing |  | 12 | 52.1444 | 42.0333 | 24.06% | G | 0.80609
Socotec Didcot | 20% TEA in water | 2019 | R | London Borough of ealing |  | 12 | 62.575 | 48.7074 | 28.47% | G | 0.77838
Gradko | 50% TEA in acetone | 2019 | R | Sandwell MBC |  | 10 | 39.6662 | 32.803 | 20.92% | G | 0.82698
Gradko | 50% TEA in acetone | 2019 | UB | Sandwell MBC |  | 11 | 16.6009 | 15.03 | 10.45% | S | 0.90537
Gradko | 50% TEA in acetone | 2019 | UB | Sandwell MBC |  | 12 | 25.4169 | 21.5067 | 18.18% | G | 0.84615
Gradko | 50% TEA in acetone | 2019 | R | Sandwell MBC |  | 12 | 30.0733 | 29.9358 | 0.4593% | S | 0.99543
Gradko | 20% TEA in water | 2019 | UB | Backpool Council |  | 12 | 16.1504 | 12.5433 | 28.76% | G | 0.77666
Socotec Didcot | 50% TEA in acetone | 2019 | R | Swansea Council |  | 12 | 32.3083 | 23.8208 | 35.63% | G | 0.7373
Socotec Didcot | 50% TEA in acetone | 2019 | UB | Swansea Council |  | 12 | 17.0167 | 12.9933 | 30.96% | G | 0.76357
Socotec Didcot | 50% TEA in acetone | 2019 | R | Knowsley MBC |  | 12 | 46.2167 | 37.4167 | 23.52% | G | 0.80959
Socotec Didcot | 50% TEA in acetone | 2019 | UI | North Lincolnshire Council |  | 12 | 22.2028 | 15.05 | 47.53% | G | 0.67784
Aberdeen Scientific Services | 20% TEA in water | 2020 | R | Aberdeen City Council |  | 9 | 15.9259 | 12.1278 | 31.32% | G | 0.76151
Aberdeen Scientific Services | 20% TEA in water | 2020 | UB | Aberdeen City Council |  | 11 | 20.1515 | 14.2545 | 41.37% | G | 0.70737
Aberdeen Scientific Services | 20% TEA in water | 2020 | R | Aberdeen City Council |  | 12 | 20.2222 | 16.5667 | 22.07% | G | 0.81923
Aberdeen Scientific Services | 20% TEA in water | 2020 | R | Aberdeen City Council |  | 12 | 30.1111 | 22.2917 | 35.08% | G | 0.74031
Aberdeen Scientific Services | 20% TEA in water | 2020 | R | Aberdeen City Council |  | 12 | 30.8056 | 24.55 | 25.48% | G | 0.79693
Aberdeen Scientific Services | 20% TEA in water | 2020 | R | Aberdeen City Council |  | 12 | 31.9444 | 26.0417 | 22.67% | G | 0.81522
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | East Suffolk Council |  | 12 | 29.9944 | 25.0693 | 19.65% | G | 0.8358
Staffordshire Scientific Services | 20% TEA in water | 2020 | R | Cannock Chase Council |  | 12 | 19.8861 | 15.2333 | 30.54% | G | 0.76603
SOCOTEC Didcot | 50% TEA in acetone | 2020 | UB | Canterbury City Council |  | 10 | 13.2317 | 10.33 | 28.09% | G | 0.7807
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Canterbury City Council |  | 9 | 26.1593 | 20.1778 | 29.64% | G | 0.77134
Gradko | 50% TEA in Acetone | 2020 | UC | Falkirk Council |  | 10 | 32.7253 | 26.2 | 24.91% | G | 0.8006
SOCOTEC Didcot | 50% TEA in acetone | 2020 | UB | Kingston upon Hull City Council |  | 12 | 24.4861 | 18.1667 | 34.79% | G | 0.74192
Gradko | 20% TEA in water | 2020 | R | Gedling Borough Council |  | 10 | 30.777 | 24.8 | 24.1% | G | 0.8058
Gradko | 50% TEA in Acetone | 2020 | UB | Falkirk Council |  | 11 | 16.0282 | 12 | 33.57% | G | 0.74868
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Ipswich Borough Council |  | 12 | 26.5472 | 20.655 | 28.53% | G | 0.77805
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Ipswich Borough Council |  | 12 | 35.7917 | 26.2675 | 36.26% | G | 0.7339
Gradko | 50% TEA in acetone | 2020 | UB | Middlesbrough Council |  | 10 | 17.5502 | 12.014 | 46.08% | G | 0.68455
Milton Keynes Council | 20% TEA in water | 2020 | UC | Milton Keynes Council |  | 12 | 16.3487 | 16.3033 | 0.2786% | G | 0.99722
Milton Keynes Council | 20% TEA in water | 2020 | R | Milton Keynes Council |  | 11 | 27.1045 | 24.6136 | 10.12% | G | 0.9081
Milton Keynes Council | 20% TEA in water | 2020 | R | Milton Keynes Council |  | 12 | 24.0239 | 17.6903 | 35.8% | G | 0.73636
SOCOTEC Didcot | 20% TEA in water | 2020 | R | Rhondda Cynon Taf CBC |  | 9 | 28.7778 | 23.4177 | 22.89% | G | 0.81374
Gradko | 50% TEA in acetone | 2020 | R | Royal Borough of Windsor and Maidenhead |  | 12 | 28.8678 | 24.6167 | 17.27% | G | 0.85274
Gradko | 50% TEA in acetone | 2020 | R | Royal Borough of Windsor and Maidenhead |  | 12 | 24.316 | 21.7667 | 11.71% | G | 0.89516
Gradko | 20% TEA in water | 2020 | R | SOUTHAMPTON CITY COUNCIL |  | 12 | 36.5596 | 26.675 | 37.06% | G | 0.72963
Gradko | 50% TEA in acetone | 2020 | SU | Redcar & Cleveland Borough Council |  | 11 | 16.0864 | 13.0373 | 23.39% | P | 0.81045
Staffordshire Scientific Services | 20% TEA in water | 2020 | UB | Stoke-on-Trent City Council |  | 9 | 22.9333 | 18.6701 | 22.83% | G | 0.8141
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Thanet District Council |  | 9 | 20.2 | 16.6667 | 21.2% | G | 0.82508
Lambeth Scientific Services | 50% TEA in acetone | 2020 | R | Elmbridge Borough Council |  | 12 | 24.0556 | 23.75 | 1.287% | G | 0.9873
Lambeth Scientific Services | 50% TEA in acetone | 2020 | R | Elmbridge Borough Council |  | 12 | 25.9444 | 26.4167 | -1.788% | G | 1.0182
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Medway Council |  | 12 | 26 | 18.3534 | 41.66% | G | 0.7059
SOCOTEC Didcot | 50% TEA in acetone | 2020 | B | Medway Council |  | 11 | 19.6576 | 10.0152 | 96.28% | G | 0.50949
Glasgow Scientific Services | 20% TEA in water | 2020 | R | East Dunbartonshire Council |  | 11 | 22.5273 | 19.4545 | 15.79% | P | 0.8636
Glasgow Scientific Services | 20% TEA in water | 2020 | R | East Dunbartonshire Council |  | 11 | 18.7091 | 18 | 3.939% | G | 0.9621
Glasgow Scientific Services | 20% TEA in water | 2020 | R | East Dunbartonshire Council |  | 10 | 14.8633 | 15.3 | -2.854% | P | 1.0294
SOCOTEC Didcot | 50% TEA in acetone | 2020 | B | Gravesham Borough Council |  | 12 | 23.1833 | 21.9583 | 5.579% | G | 0.94716
SOCOTEC Didcot | 50% TEA in acetone | 2020 | B | Gravesham Borough Council |  | 12 | 27.3861 | 23.5917 | 16.08% | G | 0.86145
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Monmouthshire County Concil |  | 10 | 32.27 | 23.8443 | 35.34% | G | 0.7389
SOCOTEC Didcot | 50% TEA in acetone | 2020 | UI | North Lincolnshire Council |  | 13 | 17.6949 | 13.9769 | 26.6% | G | 0.78989
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | City of York Council |  | 12 | 24.225 | 18.7836 | 28.97% | G | 0.77538
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | City of York Council |  | 11 | 22.4985 | 16.7475 | 34.34% | G | 0.74438
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | City of York Council |  | 12 | 32.85 | 23.3978 | 40.4% | G | 0.71226
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Cambridge City Council |  | 10 | 29.61 | 20.0644 | 47.57% | G | 0.67762
Gradko | 20% TEA in water | 2020 | R | Fareham Borough Council |  | 10 | 24.9677 | 14.0723 | 77.42% | G | 0.56362
Gradko | 20% TEA in water | 2020 | R | Fareham Borough Council |  | 12 | 29.845 | 22.09 | 35.11% | G | 0.74016
Gradko | 20% TEA in water | 2020 | R | Fareham Borough Council |  | 10 | 21.714 | 17.1717 | 26.45% | G | 0.79081
Staffordshire Scientific Services | 20% TEA in water | 2020 | KS | Manchester City Council |  | 12 | 39.8611 | 36.5046 | 9.195% | G | 0.91579
Staffordshire Scientific Services | 20% TEA in water | 2020 | UC | Manchester City Council |  | 12 | 26.2389 | 27.2981 | -3.88% | G | 1.0404
Staffordshire Scientific Services | 20% TEA in water | 2020 | SI | Manchester City Council |  | 12 | 16.0778 | 14.5878 | 10.21% | G | 0.90732
Gradko | 50% TEA in acetone | 2020 | R | Newham |  | 10 | 28.6093 | 24.21 | 18.17% | G | 0.84623
Staffordshire Scientific Services | 20% TEA in water | 2020 | R | Oldham Council |  | 11 | 28.55 | 22.6355 | 26.13% | G | 0.79284
Staffordshire Scientific Services | 20% TEA in water | 2020 | UB | Salford City Council |  | 12 | 21.5667 | 20.5279 | 5.06% | G | 0.95184
Staffordshire Scientific Services | 20% TEA in water | 2020 | B | Salford City Council |  | 12 | 13.0111 | 11.0187 | 18.08% | G | 0.84687
Staffordshire Scientific Services | 20% TEA in water | 2020 | R | Salford City Council |  | 12 | 37.0917 | 34.3235 | 8.065% | G | 0.92537
Gradko | 50% TEA in acetone | 2020 | R | Sandwell MBC |  | 12 | 33.9332 | 26.7442 | 26.88% | G | 0.78814
Gradko | 50% TEA in acetone | 2020 | B | Sandwell MBC |  | 9 | 13.6167 | 11.0719 | 22.98% | S | 0.81311
Gradko | 50% TEA in acetone | 2020 | R | Sandwell MBC |  | 11 | 24.9483 | 22.8091 | 9.379% | S | 0.91425
Gradko | 50% TEA in acetone | 2020 | UB | Sandwell Metropolitan Borough Council |  | 11 | 20.5012 | 18.7341 | 9.432% | G | 0.91381
Gradko | 20% TEA in water | 2020 | R | SOUTHAMPTON CITY COUNCIL |  | 11 | 32.3279 | 30.8091 | 4.93% | G | 0.95302
Lambeth Scientific Services | 50% TEA in acetone | 2020 | UB | Spelthorne Borough Counci |  | 12 | 21.8056 | 22.8625 | -4.623% | G | 1.0485
Lambeth Scientific Services | 50% TEA in acetone | 2020 | UB | Spelthorne Borough Council |  | 11 | 16.7273 | 17.2391 | -2.969% | P | 1.0306
Staffordshire Scientific Services | 20% TEA in water | 2020 | R | Stockport MBC |  | 12 | 31.1625 | 26.2897 | 18.54% | G | 0.84363
Staffordshire Scientific Services | 20% TEA in water | 2020 | R | Stockport MBC |  | 10 | 20.2233 | 16.8944 | 19.7% | G | 0.83539
Staffordshire Scientific Services | 20% TEA in water | 2020 | UB | Wigan Council |  | 12 | 21.0889 | 15.4288 | 36.68% | G | 0.73161
SOCOTEC Didcot | 50% TEA in acetone | 2020 | R | Wrexham County Borough Council |  | 9 | 16.9222 | 13.3667 | 26.6% | G | 0.78989
Aberdeen Scientific Services | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 11 | 54.0155 | 42.7596 | 26.32% | G | 0.79162
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | KS | Marylebone Road Intercomparison |  | 10 | 50.9653 | 43.4297 | 17.35% | G | 0.85214
Glasgow Scientific Services | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 11 | 53.1833 | 43.6895 | 21.73% | G | 0.82149
Gradko | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 12 | 56.8256 | 42.6339 | 33.29% | G | 0.75026
Gradko | 50% TEA in acetone | 2020 | KS | Marylebone Road Intercomparison |  | 12 | 56.6983 | 42.6339 | 32.99% | G | 0.75194
Lambeth Scientific Services | 50% TEA in acetone | 2020 | KS | Marylebone Road Intercomparison |  | 12 | 54.7417 | 42.6339 | 28.4% | G | 0.77882
Milton Keynes Council | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 12 | 58.9395 | 42.6339 | 38.25% | G | 0.72335
SOCOTEC Didcot | 50% TEA in acetone | 2020 | KS | Marylebone Road Intercomparison |  | 11 | 58.9924 | 42.7596 | 37.96% | G | 0.72483
SOCOTEC Didcot | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 11 | 56.7242 | 42.7596 | 32.66% | G | 0.75381
SOCOTEC Glasgow | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 11 | 54.4394 | 42.7596 | 27.32% | G | 0.78545
SOCOTEC Glasgow | 50% TEA in acetone | 2020 | KS | Marylebone Road Intercomparison |  | 10 | 53.7233 | 42.5418 | 26.28% | G | 0.79187
Somerset County Council | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 11 | 56.4658 | 42.7596 | 32.05% | G | 0.75726
Staffordshire Scientific Services | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 11 | 52.6788 | 44.3773 | 18.71% | G | 0.84241
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2020 | KS | Marylebone Road Intercomparison |  | 12 | 55.5861 | 42.6339 | 30.38% | G | 0.76699
Tayside Scientific Services | 20% TEA in water | 2020 | KS | Marylebone Road Intercomparison |  | 12 | 57.2139 | 42.6339 | 34.2% | G | 0.74517
Gradko | 20% TEA in water | 2020 | R | Bath & North East Somerset |  | 11 | 32.266 | 28.5606 | 12.97% | G | 0.88516
Staffordshire Scientific Services | 20% TEA in water | 2020 | R | East Staffordshire Borough Council |  | 10 | 38.6367 | 36.02 | 7.264% | G | 0.93228
SOCOTEC Didcot | 20% TEA in water | 2020 | R | Fife Council |  | 9 | 21.8481 | 13.2778 | 64.55% | G | 0.60773
SOCOTEC Didcot | 20% TEA in water | 2020 | R | Fife Council |  | 9 | 22.4148 | 17.0556 | 31.42% | G | 0.76091
Gradko | 20% TEA in water | 2020 | R | Gateshead Council |  | 12 | 22.3667 | 17.4667 | 28.05% | G | 0.78092
Gradko | 20% TEA in water | 2020 | R | Gateshead Council |  | 12 | 23.3519 | 20.9242 | 11.6% | G | 0.89604
Gradko | 20% TEA in water | 2020 | R | Gateshead Council |  | 10 | 26.396 | 24.777 | 6.534% | G | 0.93866
Gradko | 20% TEA in water | 2020 | R | Gateshead Council |  | 12 | 27.6672 | 21.2083 | 30.45% | G | 0.76655
Gradko | 20% TEA in water | 2020 | R | Gateshead Council |  | 12 | 30.5522 | 31.62 | -3.377% | G | 1.0349
Glasgow Scientific Services | 20% TEA in water | 2020 | R | Glasgow City Council |  | 12 | 25.7097 | 22.7417 | 13.05% | P | 0.88456
Glasgow Scientific Services | 20% TEA in water | 2020 | R | Glasgow City Council |  | 12 | 20.7639 | 19.8333 | 4.692% | P | 0.95518
Glasgow Scientific Services | 20% TEA in water | 2020 | R | Glasgow City Council |  | 11 | 21.6818 | 22.5 | -3.636% | P | 1.0377
Glasgow Scientific Services | 20% TEA in water | 2020 | KS | Glasgow City Council |  | 12 | 33.2347 | 36.3 | -8.444% | P | 1.0922
Glasgow Scientific Services | 20% TEA in water | 2020 | UB | Glasgow City Council |  | 12 | 18.6778 | 17.4667 | 6.934% | P | 0.93516
Gradko | 50% TEA in acetone | 2020 | R | London Borough of Richmond upon Thames |  | 12 | 22.0445 | 20.1417 | 9.447% | G | 0.91368
Gradko | 50% TEA in acetone | 2020 | B | London Borough of Richmond upon Thames |  | 9 | 19.089 | 15.8667 | 20.31% | G | 0.83119
Gradko | 20% TEA in water | 2020 | R | Luton Borough Council |  | 9 | 37.5251 | 28.044 | 33.81% | G | 0.74734
Gradko | 20% TEA in water | 2020 | R | Nottingham City Council |  | 12 | 31.1021 | 34 | -8.523% | G | 1.0932
Somerset County Council | 20% TEA in water | 2020 | R | South Gloucestershire Council |  | 12 | 18.7629 | 14.4417 | 29.92% | G | 0.76969
SOCOTEC Didcot | 20% TEA in water | 2020 | R | South Oxfordshire District Council |  | 11 | 32.3515 | 28.5118 | 13.47% | G | 0.88131
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | R | Stirling Council |  | 11 | 16.2333 | 14.5455 | 11.6% | G | 0.89602
Gradko | 20% TEA in water | 2020 | R | Dudley MBC |  | 13 | 33.434 | 27.8846 | 19.9% | G | 0.83402
Gradko | 20% TEA in water | 2020 | UB | Dudley MBC |  | 13 | 23.2896 | 14.4462 | 61.22% | G | 0.62028
Gradko | 20% TEA in water | 2020 | R | Dudley MBC |  | 13 | 43.8304 | 33.5538 | 30.63% | G | 0.76554
Socotec Didcot | 50% TEA in acetone | 2020 | R | Horsham District Council |  | 10 | 23.4033 | 22.9059 | 2.172% | G | 0.97874
Socotec Didcot | 50% TEA in acetone | 2020 | R | Horsham District Council |  | 12 | 22.1806 | 18.7078 | 18.56% | G | 0.84343
Socotec Didcot | 50% TEA in acetone | 2020 | R | Horsham District Council |  | 9 | 24.8815 | 17.517 | 42.04% | G | 0.70402
Socotec Didcot | 20% TEA in water | 2020 | R | New Forest DC |  | 9 | 26.9963 | 18.4567 | 46.27% | G | 0.68367
Staffordshire Scientific Services | 20% TEA in water | 2020 | R | Tameside MBC |  | 12 | 44.5472 | 30.7243 | 44.99% | G | 0.6897
Gradko | 20% TEA in water | 2020 | R | Ards and North Down Borough Council |  | 10 | 26.5667 | 19.82 | 34.04% | G | 0.74605
Gradko | 20% TEA in water | 2020 | R | Belfast City Council |  | 10 | 25.9373 | 21.119 | 22.82% | G | 0.81423
Gradko | 20% TEA in water | 2020 | R | Belfast City Council |  | 10 | 40.7467 | 36.188 | 12.6% | G | 0.88812
Gradko | 20% TEA in water | 2020 | R | Belfast City Council |  | 10 | 35.758 | 24.846 | 43.92% | G | 0.69484
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | KS | The City of Edinburgh Council |  | 9 | 30.9778 | 28.0778 | 10.33% | G | 0.90638
Somerset County Council | 20% TEA in water | 2020 | UB | Bristol City Council |  | 11 | 22.0553 | 18.6919 | 17.99% | G | 0.8475
Somerset County Council | 20% TEA in water | 2020 | R | Bristol City Council |  | 11 | 64.46 | 46.7373 | 37.92% | G | 0.72506
Somerset County Council | 20% TEA in water | 2020 | R | Bristol City Council |  | 11 | 29.846 | 28.1858 | 5.89% | G | 0.94438
Somerset County Council | 20% TEA in water | 2020 | UB | Bristol City Council |  | 12 | 18.6949 | 15.4953 | 20.65% | G | 0.82885
Somerset County Council | 20% TEA in water | 2020 | R | Bristol City Council |  | 12 | 30.6225 | 28.5623 | 7.213% | G | 0.93272
Somerset County Council | 20% TEA in water | 2020 | KS | Bristol City Council |  | 12 | 31.7577 | 28.0249 | 13.32% | G | 0.88246
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | R | The City of Edinburgh Council |  | 9 | 21.3963 | 16.3444 | 30.91% | P | 0.76389
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | R | The City of Edinburgh Council |  | 9 | 32.8407 | 27.8444 | 17.94% | G | 0.84786
Socotec Didcot | 50% TEA in acetone | 2020 | R | Dacorum Borough Council |  | 10 | 23.8033 | 19.01 | 25.21% | G | 0.79863
Gradko | 20% TEA in water | 2020 | R | Lancaster City Council |  | 11 | 27.3585 | 22.8091 | 19.95% | G | 0.83371
Gradko | 20% TEA in water | 2020 | R | Lancaster City Council |  | 10 | 31.9843 | 28.31 | 12.98% | G | 0.88512
Somerset County Council | 20% TEA in water | 2020 | R | Wiltshire Council |  | 11 | 30.5152 | 26.0673 | 17.06% | G | 0.85424
Somerset County Council | 20% TEA in water | 2020 | R | Wiltshire Council |  | 12 | 26.5833 | 25.3283 | 4.955% | G | 0.95279
Gradko | 50% TEA in acetone | 2020 | UB | Reading Borough Council |  | 12 | 14.3036 | 15.4932 | -7.678% | G | 1.0832
Gradko | 50% TEA in acetone | 2020 | R | Reading Borough Council |  | 12 | 29.8042 | 24.797 | 20.19% | G | 0.832
Gradko | 20% TEA in water | 2020 | R | Eastleigh Borough Council |  | 9 | 22.9885 | 20.2341 | 13.61% | G | 0.88018
Gradko | 20% TEA in water | 2020 | UB | Eastleigh Borough Council |  | 9 | 22.2581 | 18.8835 | 17.87% | G | 0.84839
Socotec Didcot | 50% TEA in acetone | 2020 | R | Huntingdonshire District Council |  | 12 | 36.4639 | 24.7863 | 47.11% | G | 0.67975
Gradko | 20% TEA in water | 2020 | R | Lisburn & Castlereagh City Council |  | 10 | 23.3 | 17.58 | 32.54% | G | 0.75451
Gradko | 50% TEA in acetone | 2020 | UB | Norwich City Council |  | 10 | 11.9593 | 10.45 | 14.44% | G | 0.87379
Gradko | 50% TEA in acetone | 2020 | SU | Reigate and Banstead BC (RG1) |  | 10 | 18.602 | 13.9558 | 33.29% | G | 0.75023
Lambeth Scientific Services | 50% TEA in acetone | 2020 | SU | Reigate and Banstead BC (RG1) |  | 12 | 15.6667 | 13.0028 | 20.49% | G | 0.82996
Lambeth Scientific Services | 50% TEA in acetone | 2020 | B | Reigate and Banstead BC (RG3) |  | 12 | 10.1111 | 9.7723 | 3.467% | P | 0.96649
Lambeth Scientific Services | 50% TEA in acetone | 2020 | SU | Reigate and Banstead BC RG2(6) |  | 12 | 15 | 14.4286 | 3.961% | G | 0.9619
Lambeth Scientific Services | 50% TEA in acetone | 2020 | R | Reigate and Banstead BC (RG7) |  | 11 | 36.0303 | 37.5819 | -4.129% | G | 1.0431
Gradko | 50% Tea in Acetone | 2020 | KS | Slough Borough Council |  | 12 | 33.9485 | 27.4845 | 23.52% | G | 0.80959
Gradko | 50% TEA in Acetone | 2020 | SU | Slough Borough Council |  | 11 | 21.3937 | 16.5625 | 29.17% | G | 0.77418
Gradko | 50% TEA in Acetone | 2020 | KS | Slough Borough Council |  | 12 | 29.0363 | 24.6337 | 17.87% | G | 0.84838
Lambeth Scientific Services | 50% TEA in acetone | 2020 | R | London Borough of Islington |  | 12 | 32.6389 | 30.8333 | 5.856% | G | 0.94468
Gradko | 50% TEA in acetone | 2020 | R | East Herts District Council |  | 11 | 24.8788 | 25.9747 | -4.219% | G | 1.0441
Aberdeen Scientific Services | 20% TEA in water | 2021 | R | Aberdeen City Council |  | 12 | 17.6944 | 12.6833 | 39.51% | G | 0.7168
Aberdeen Scientific Services | 20% TEA in water | 2021 | UB | Aberdeen City Council |  | 10 | 17.5667 | 14.12 | 24.41% | G | 0.8038
Aberdeen Scientific Services | 20% TEA in water | 2021 | R | Aberdeen City Council |  | 11 | 21.0606 | 16.6364 | 26.59% | G | 0.78993
Aberdeen Scientific Services | 20% TEA in water | 2021 | R | Aberdeen City Council |  | 13 | 34.7692 | 26.4154 | 31.62% | G | 0.75973
Aberdeen Scientific Services | 20% TEA in water | 2021 | R | Aberdeen City Council |  | 13 | 30.8718 | 24 | 28.63% | G | 0.77741
Aberdeen Scientific Services | 20% TEA in water | 2021 | R | Aberdeen City Council |  | 13 | 35.6154 | 27.1462 | 31.2% | G | 0.7622
Lambeth Scientific Services | 50% TEA in acetone | 2021 | R | Elmbridge Borough Council |  | 11 | 26.9394 | 26.1909 | 2.858% | P | 0.97222
Lambeth Scientific Services | 50% TEA in acetone | 2021 | R | Elmbridge Borough Council |  | 12 | 27.3889 | 25.2083 | 8.65% | G | 0.92039
Gradko | 50% TEA in acetone | 2021 | UC | Falkirk Council |  | 12 | 35.3753 | 34.1667 | 3.537% | G | 0.96583
Gradko | 50% TEA in acetone | 2021 | UB | Falkirk Council |  | 12 | 16.4392 | 13.4167 | 22.53% | G | 0.81614
Gradko | 20% TEA in water | 2021 | R | Gedling Borough Council |  | 12 | 31.7237 | 25.775 | 23.08% | G | 0.81248
Socotec Didcot | 50% TEA in acetone | 2021 | UB | Gravesham Borough Council |  | 12 | 22.6611 | 20.9983 | 7.919% | G | 0.92662
Socotec Didcot | 50% TEA in acetone | 2021 | UB | Gravesham Borough Council |  | 11 | 27.1879 | 23.0109 | 18.15% | G | 0.84637
Socotec Didcot | 50% TEA in acetone | 2021 | R | Horsham District Council |  | 12 | 27.2056 | 20.225 | 34.51% | G | 0.74341
Socotec Didcot | 50% TEA in acetone | 2021 | R | Ipswich Borough Council |  | 12 | 28.6944 | 23.1833 | 23.77% | G | 0.80794
Socotec Didcot | 50% TEA in acetone | 2021 | R | Ipswich Borough Council |  | 12 | 38.0028 | 28.5008 | 33.34% | G | 0.74997
Socotec Didcot | 50% TEA in acetone | 2021 | UB | Kingston upon Hull City Council |  | 11 | 23.6167 | 16.9091 | 39.67% | G | 0.71598
Socotec Didcot | 50% TEA in acetone | 2021 | R | Kingston upon Hull City Council |  | 12 | 30.4194 | 24.7417 | 22.95% | G | 0.81335
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | Newcastle under Lyme Borough Souncil |  | 9 | 28.1444 | 18.9 | 48.91% | G | 0.67154
Gradko | 50% TEA in acetone | 2021 | SU | Redcar & Cleveland Borough Council |  | 11 | 14.2567 | 11.0382 | 29.16% | G | 0.77425
Gradko | 50% TEA in acetone | 2021 | R | Royal Borough of Windsor and Maidenhead |  | 12 | 28.8639 | 26.4 | 9.333% | G | 0.91464
Gradko | 50% TEA in acetone | 2021 | R | Royal Borough of Windsor and Maidenhead |  | 11 | 26.3854 | 24.6182 | 7.178% | G | 0.93302
Staffordshire Scientific Services | 20% TEA in water | 2021 | UB | Salford City Council |  | 12 | 23.85 | 22.6368 | 5.359% | G | 0.94913
Staffordshire Scientific Services | 20% TEA in water | 2021 | B | Salford City Council |  | 12 | 13.3806 | 11.6272 | 15.08% | G | 0.86896
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | Salford City Council |  | 12 | 38.3444 | 34.5365 | 11.03% | G | 0.90069
Gradko | 50% TEA in Acetone | 2021 | R | Sandwell MBC |  | 12 | 36.8125 | 28.0167 | 31.4% | G | 0.76106
Gradko | 50% TEA in Acetone | 2021 | UB | Sandwell Metropolitan Borough Council |  | 11 | 22.6492 | 18.5364 | 22.19% | G | 0.81841
Somerset County Council | 20% TEA in water | 2021 | R | South Gloucestershire Council |  | 12 | 19.5671 | 14.9 | 31.32% | G | 0.76148
Lambeth Scientific Services | 50% TEA in acetone | 2021 | UB | Spelthorne Borough Council |  | 12 | 18.75 | 18.575 | 0.9421% | G | 0.99067
Lambeth Scientific Services | 50% TEA in acetone | 2021 | UB | Spelthorne Borough Council |  | 11 | 22.5303 | 23.0727 | -2.351% | G | 1.0241
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | Stoke-on-Trent City Council |  | 12 | 49.4639 | 48.9083 | 1.136% | G | 0.98877
Staffordshire Scientific Services | 20% TEA in water | 2021 | UB | Stoke-on-Trent City Council |  | 12 | 23.4139 | 18.85 | 24.21% | G | 0.80508
Gradko | 20% TEA in water | 2021 | UB | West Northamptonshire Council |  | 11 | 13.5715 | 10.2727 | 32.11% | G | 0.75693
Gradko | 20% TEA in water | 2021 | R | Ards and North Down Borough Council |  | 10 | 30.1333 | 21.77 | 38.42% | G | 0.72246
Gradko | 20% TEA in water | 2021 | R | Birmingham City Council |  | 10 | 33.2243 | 24.567 | 35.24% | G | 0.73943
Gradko | 20% TEA in water | 2021 | R | Cheshire West and Chester |  | 12 | 33.6088 | 29.4667 | 14.06% | G | 0.87675
Gradko | 20% TEA in water | 2021 | R | Cheshire West and Chester |  | 12 | 32.602 | 28.9583 | 12.58% | G | 0.88824
SOCOTEC Didcot | 50% TEA in acetone | 2021 | UB | City of York Council |  | 11 | 17.4955 | 12.6589 | 38.21% | G | 0.72355
SOCOTEC Didcot | 50% TEA in acetone | 2021 | R | City of York Council |  | 12 | 24.8611 | 19.5736 | 27.01% | G | 0.78732
SOCOTEC Didcot | 50% TEA in acetone | 2021 | R | City of York Council |  | 12 | 22.0889 | 17.1167 | 29.05% | G | 0.7749
SOCOTEC Didcot | 50% TEA in acetone | 2021 | R | City of York Council |  | 12 | 36.8056 | 25.2883 | 45.54% | G | 0.68708
Gradko | 20% TEA in water | 2021 | R | Lisburn & Castlereagh City Council |  | 12 | 24.7639 | 18.775 | 31.9% | G | 0.75816
Gradko | 50% TEA in acetone | 2021 | UB | Middlesbrough |  | 12 | 18.0714 | 13.625 | 32.63% | G | 0.75395
SOCOTEC Didcot | 50% TEA in acetone | 2021 | UI | North Lincolnshire Council |  | 12 | 17.1014 | 14.2583 | 19.94% | G | 0.83375
Gradko | 20% TEA in water | 2021 | R | Nottingham City Council |  | 12 | 31.9397 | 34.75 | -8.087% | G | 1.088
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | Oldham Borough Council |  | 10 | 33.9217 | 24.1414 | 40.51% | P | 0.71168
Gradko | 20% TEA in water | 2021 | R | SOUTHAMPTON CITY COUNCIL |  | 12 | 33.9456 | 32.275 | 5.176% | G | 0.95079
Gradko | 20% TEA in water | 2021 | R | SOUTHAMPTON CITY COUNCIL |  | 12 | 34.0786 | 26.5 | 28.6% | G | 0.77761
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | Trafford Council |  | 9 | 27.7383 | 23.7067 | 17.01% | G | 0.85465
Gradko | 20% TEA in water | 2021 | R | Bath & North East Somerset |  | 12 | 30.9221 | 26.8615 | 15.12% | G | 0.86868
Gradko | 20% TEA in water | 2021 | R | Bedford Borough Council |  | 11 | 33.7273 | 31.3455 | 7.599% | G | 0.92938
Gradko | 20% TEA in water | 2021 | R | Bedford Borough Council |  | 11 | 19.3333 | 17.3091 | 11.69% | G | 0.8953
Gradko | 20% TEA in water | 2021 | R | Blackburn with Darwen Borough Council |  | 12 | 26.7295 | 20.199 | 32.33% | G | 0.75568
Gradko | 20% TEA in water | 2021 | R | Brent Council |  | 12 | 50.814 | 46.2465 | 9.876% | G | 0.91011
Socotec Didcot | 50% TEA in acetone | 2021 | R | Bridgend Borough County Council / Shared Regulatory Services |  | 12 | 36.3764 | 25.45 | 42.93% | G | 0.69963
Socotec Didcot | 50% TEA in acetone | 2021 | UB | Derry City and Strabane District Council |  | 12 | 11.1333 | 8.6708 | 28.4% | G | 0.77882
Socotec Didcot | 50% TEA in acetone | 2021 | R | Derry City and Strabane District Council |  | 12 | 30.4139 | 29.7108 | 2.366% | G | 0.97688
Socotec Didcot | 50% TEA in acetone | 2021 | R | East Suffolk Council |  | 11 | 30.0909 | 24.6106 | 22.27% | P | 0.81787
Gradko | 20% TEA in water | 2021 | R | Gateshead Council |  | 10 | 23.162 | 18.71 | 23.79% | G | 0.80779
Gradko | 20% TEA in water | 2021 | R | Gateshead Council |  | 12 | 25.1722 | 22.1417 | 13.69% | G | 0.87961
Gradko | 20% TEA in water | 2021 | R | Gateshead Council |  | 11 | 27.3794 | 24.9455 | 9.757% | G | 0.9111
Gradko | 20% TEA in water | 2021 | R | Gateshead Council |  | 12 | 31.3431 | 24.765 | 26.56% | G | 0.79013
Gradko | 20% TEA in water | 2021 | R | Gateshead Council |  | 12 | 32.3222 | 33.6917 | -4.065% | G | 1.0424
Glasgow Scientific Services | 20% TEA in water | 2021 | R | Glasgow City Council |  | 12 | 25.9764 | 24.9583 | 4.079% | P | 0.96081
Glasgow Scientific Services | 20% TEA in water | 2021 | R | Glasgow City Council |  | 11 | 16.2136 | 20.8091 | -22.08% | P | 1.2834
Glasgow Scientific Services | 20% TEA in water | 2021 | R | Glasgow City Council |  | 12 | 17.5764 | 21.9167 | -19.8% | P | 1.2469
Glasgow Scientific Services | 20% TEA in water | 2021 | KS | Glasgow City Council |  | 12 | 37.4306 | 44.2167 | -15.35% | P | 1.1813
Glasgow Scientific Services | 20% TEA in water | 2021 | UB | Glasgow City Council |  | 12 | 13.9361 | 17.275 | -19.33% | P | 1.2396
Gradko | 50% TEA in acetone | 2021 | R | London Borough of Richmond upon Thames |  | 12 | 24.3961 | 21.2 | 15.08% | G | 0.86899
Gradko | 50% TEA in acetone | 2021 | B | London Borough of Richmond upon Thames |  | 9 | 16.2683 | 13.3889 | 21.51% | G | 0.823
Staffordshire Scientific Services | 20% TEA in water | 2021 | KS | Manchester City Council |  | 12 | 47.3583 | 43.1459 | 9.763% | G | 0.91105
Staffordshire Scientific Services | 20% TEA in water | 2021 | UC | Manchester City Council |  | 12 | 28.5361 | 29.2928 | -2.583% | G | 1.0265
Staffordshire Scientific Services | 20% TEA in water | 2021 | SI | Manchester City Council |  | 12 | 16.4083 | 15.4189 | 6.417% | G | 0.9397
Lambeth Scientific Services | 50% TEA in acetone | 2021 | KS | Marylebone Road Intercomparison |  | 11 | 45.5545 | 42.1163 | 8.164% | G | 0.92452
Socotec Glasgow | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 10 | 53.6348 | 41.8432 | 28.18% | P | 0.78015
Socotec Glasgow | 50% TEA in acetone | 2021 | KS | Marylebone Road Intercomparison |  | 10 | 52.3937 | 41.8432 | 25.21% | P | 0.79863
Socotec Didcot | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 10 | 56.7724 | 41.8432 | 35.68% | P | 0.73703
Socotec Didcot | 50% TEA in acetone | 2021 | KS | Marylebone Road Intercomparison |  | 10 | 55.6169 | 41.8432 | 32.92% | P | 0.75235
Milton Keynes Council | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 10 | 56.0848 | 41.5312 | 35.04% | G | 0.74051
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2021 | KS | Marylebone Road Intercomparison |  | 11 | 54.9542 | 42.1163 | 30.48% | G | 0.76639
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | KS | Marylebone Road Intercomparison |  | 11 | 48.3534 | 42.1163 | 14.81% | G | 0.87101
Glasgow Scientific Services | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 10 | 46.4 | 41.4757 | 11.87% | G | 0.89387
Gradko | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 11 | 52.6469 | 42.1163 | 25% | G | 0.79998
Gradko | 50% TEA in acetone | 2021 | KS | Marylebone Road Intercomparison |  | 10 | 51.5005 | 41.4757 | 24.17% | G | 0.80535
Aberdeen Scientific Services | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 11 | 55.283 | 42.1163 | 31.26% | G | 0.76183
Staffordshire Scientific Services | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 11 | 50.9273 | 42.1163 | 20.92% | G | 0.82699
Tayside Scientific Services | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 11 | 54.8398 | 42.1163 | 30.21% | G | 0.76799
Somerset County Council | 20% TEA in water | 2021 | KS | Marylebone Road Intercomparison |  | 9 | 54.539 | 42.0511 | 29.7% | G | 0.77103
Gradko | 20% TEA in Water | 2021 | R | Monmouthshire County Concil |  | 11 | 35.1464 | 28.8677 | 21.75% | G | 0.82136
Socotec Didcot | 20% TEA in water | 2021 | KS | New Forest District Council |  | 12 | 37.1861 | 24.7975 | 49.96% | G | 0.66685
Socotec Didcot | 20% TEA in water | 2021 | R | New Forest District Council |  | 12 | 29.0444 | 22.8292 | 27.23% | G | 0.78601
Socotec Didcot | 50% TEA in acetone | 2021 | R | North East Lincolnshire Council |  | 10 | 26.66 | 28.84 | -7.559% | G | 1.0818
Socotec Didcot | 50% TEA in acetone | 2021 | R | North East Lincolnshire Council |  | 9 | 44.7556 | 33.2667 | 34.54% | P | 0.7433
Gradko | 50% TEA in acetone | 2021 | R | Reading Borough Council |  | 12 | 30.2293 | 26.0825 | 15.9% | G | 0.86282
Staffordshire Scientific Services | 20% TEA in water | 2021 | UB | Wigan Council |  | 12 | 20.7028 | 16.6008 | 24.71% | G | 0.80187
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | Wigan Council |  | 12 | 27.8625 | 24.9867 | 11.51% | G | 0.89678
Gradko | 20% TEA in water | 2021 | R | Belfast City Council |  | 12 | 24.9497 | 20.6392 | 20.89% | G | 0.82723
Gradko | 20% TEA in water | 2021 | UC | Belfast City Council |  | 11 | 25.7483 | 20.5364 | 25.38% | G | 0.79758
Gradko | 20% TEA in water | 2021 | R | Belfast City Council |  | 12 | 42.4969 | 36.0917 | 17.75% | G | 0.84928
Gradko | 20% TEA in water | 2021 | R | Belfast City Council |  | 12 | 38.3189 | 27.4842 | 39.42% | G | 0.71725
Gradko | 20% TEA in water | 2021 | UB | Dudley MBC |  | 12 | 20.4008 | 15 | 36.01% | G | 0.73527
Gradko | 20% TEA in water | 2021 | R | Dudley MBC |  | 12 | 30.1297 | 28.9167 | 4.195% | G | 0.95974
Gradko | 20% TEA in water | 2021 | R | Dudley MBC |  | 12 | 41.7356 | 39.5667 | 5.482% | G | 0.94803
Gradko | 20% TEA in Water | 2021 | R | Lambeth |  | 10 | 91.365 | 62.33 | 46.58% | G | 0.68221
Gradko | 20% TEA in water | 2021 | R | Lancaster City Council |  | 13 | 38.4485 | 32.4615 | 18.44% | G | 0.84429
Gradko | 20% TEA in water | 2021 | R | Lancaster City Council |  | 13 | 27.858 | 26.5615 | 4.881% | G | 0.95346
Socotec Didcot | 50% TEA in acetone | 2021 | R | Leeds City Council |  | 13 | 39.9385 | 29.4769 | 35.49% | G | 0.73806
Socotec Didcot | 50% TEA in acetone | 2021 | KS | Leeds City Council |  | 12 | 34.0889 | 24.7167 | 37.92% | G | 0.72507
Socotec Didcot | 50% TEA in acetone | 2021 | R | Leeds City Council |  | 9 | 43.1852 | 30.6778 | 40.77% | G | 0.71038
Socotec Didcot | 50% TEA in acetone | 2021 | UC | Leeds City Council |  | 12 | 31.4333 | 22.8833 | 37.36% | G | 0.728
Gradko | 50% TEA in acetone | 2021 | R | Merton Council |  | 9 | 49.7593 | 32.0111 | 55.44% | G | 0.64332
Socotec Didcot | 20% TEA in water | 2021 | R | South Oxfordshire District Council |  | 11 | 25.0576 | 18.097 | 38.46% | G | 0.72222
Socotec Didcot | 20% TEA in water | 2021 | R | South Oxfordshire District Council |  | 11 | 37.3909 | 33.1258 | 12.88% | G | 0.88593
Gradko | 50% TEA in acetone | 2021 | UB | Wandsworth Council |  | 11 | 28.8182 | 26.25 | 9.784% | G | 0.91088
Somerset County Council | 20% TEA in water | 2021 | UB | Bristol City Council |  | 11 | 23.8545 | 20.0751 | 18.83% | G | 0.84156
Somerset County Council | 20% TEA in water | 2021 | R | Bristol City Council |  | 12 | 67.0667 | 50.2153 | 33.56% | G | 0.74874
Somerset County Council | 20% TEA in water | 2021 | R | Bristol City Council |  | 11 | 27.4576 | 30.3962 | -9.668% | G | 1.107
Somerset County Council | 20% TEA in water | 2021 | R | Bristol City Council |  | 12 | 33.1167 | 31.2625 | 5.931% | G | 0.94401
Somerset County Council | 20% TEA in water | 2021 | UB | Bristol City Council |  | 12 | 18.9194 | 17.296 | 9.386% | G | 0.91419
Somerset County Council | 20% TEA in water | 2021 | R | Bristol City Council |  | 11 | 35.9152 | 31.9783 | 12.31% | G | 0.89038
Somerset County Council | 20% TEA in water | 2021 | KS | Bristol City Council |  | 12 | 33.5597 | 24.0522 | 39.53% | G | 0.7167
Socotec Didcot | 50% TEA in acetone | 2021 | R | Cambridge City Council |  | 12 | 31.5486 | 21.2393 | 48.54% | G | 0.67323
Gradko | 20% TEA in water | 2021 | R | Cheltenham Borough Council |  | 12 | 28.6358 | 25.25 | 13.41% | G | 0.88176
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | East Staffordshire Borough Council |  | 12 | 38.8389 | 25.125 | 54.58% | G | 0.6469
Socotec Didcot | 20% TEA in water | 2021 | KS | Fife Council |  | 11 | 24.4333 | 19.47 | 25.49% | G | 0.79686
Socotec Didcot | 20% TEA in water | 2021 | R | Fife Council |  | 11 | 21.4394 | 15.7618 | 36.02% | G | 0.73518
Socotec Didcot | 20% TEA in water | 2021 | R | Fife Council |  | 12 | 20.9056 | 13.7583 | 51.95% | G | 0.65812
Socotec Didcot | 20% TEA in water | 2021 | R | Fife Council |  | 12 | 23.1361 | 19.2133 | 20.42% | G | 0.83045
Gradko | 50% TEA in acetone | 2021 | R | LB Newham |  | 11 | 28.9326 | 22.8455 | 26.64% | G | 0.78961
Gradko | 50% TEA in acetone | 2021 | KS | London Borough of Croydon |  | 12 | 47.8053 | 38.7392 | 23.4% | G | 0.81035
Socotec Didcot | 20% TEA in water | 2021 | R | London Borough of Ealing |  | 12 | 63.3667 | 49.8583 | 27.09% | G | 0.78682
Socotec Didcot | 20% TEA in water | 2021 | R | London Borough of Ealing |  | 12 | 40.8056 | 31.725 | 28.62% | G | 0.77747
Socotec Didcot | 20% TEA in water | 2021 | R | London Borough of Ealing |  | 11 | 46.9306 | 35.0364 | 33.95% | G | 0.74656
Gradko | 20% TEA in water | 2021 | R | Preston City Council |  | 12 | 23.8381 | 21.2417 | 12.22% | G | 0.89108
Lambeth Scientific Services | 50% TEA in acetone | 2021 | SU | Reigate and Banstead BC |  | 11 | 15.6364 | 15.2526 | 2.516% | G | 0.97545
Lambeth Scientific Services | 50% TEA in acetone | 2021 | B | Reigate and Banstead BC |  | 12 | 11.4167 | 9.6075 | 18.83% | G | 0.84154
Lambeth Scientific Services | 50% TEA in acetone | 2021 | SU | Reigate and Banstead BC |  | 12 | 15.4444 | 13.7263 | 12.52% | G | 0.88875
Lambeth Scientific Services | 50% TEA in acetone | 2021 | R | Reigate and Banstead BC |  | 11 | 38.4242 | 40.2623 | -4.565% | G | 1.0478
Socotec Didcot | 50% TEA in acetone | 2021 | UB | Torfaen County Borough Council |  | 9 | 11.8741 | 8.7982 | 34.96% | G | 0.74096
Socotec Didcot | 20% TEA in water | 2021 | R | Vale of White Horse DC |  | 11 | 22.3394 | 17.2286 | 29.66% | G | 0.77122
Somerset County Council | 20% TEA in water | 2021 | R | Wiltshire Council |  | 12 | 28.3056 | 27.0958 | 4.465% | G | 0.95726
Somerset County Council | 20% TEA in water | 2021 | R | Wiltshire Council |  | 12 | 32.4722 | 27.4608 | 18.25% | G | 0.84567
Glasgow Scientific Services | 20% TEA in water | 2021 | R | East Ayrshire Council |  | 10 | 18.1083 | 19.2515 | -5.938% | G | 1.0631
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | R | The City of Edinburgh Council |  | 12 | 21.7194 | 16.3789 | 32.61% | G | 0.75411
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | R | The City of Edinburgh Council |  | 11 | 20.8697 | 18.0035 | 15.92% | G | 0.86266
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | R | The City of Edinburgh Council |  | 12 | 33.5861 | 28.9988 | 15.82% | G | 0.86342
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | R | The City of Edinburgh Council |  | 12 | 24.5417 | 21.8954 | 12.09% | G | 0.89217
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | KS | The City of Edinburgh Council |  | 11 | 32.8818 | 28.5357 | 15.23% | G | 0.86783
Staffordshire Scientific Services | 20% TEA in water | 2021 | R | Bury Council |  | 12 | 22.1722 | 20.3338 | 9.041% | G | 0.91708
Socotec Didcot | 50% TEA in acetone | 2021 | UB | Canterbury City Council |  | 12 | 13.525 | 9.9083 | 36.5% | G | 0.73259
Socotec Didcot | 50% TEA in acetone | 2021 | R | Canterbury City Council |  | 12 | 29.3736 | 20.6667 | 42.13% | G | 0.70358
Socotec Didcot | 20% TEA in water | 2021 | R | Rhondda Cynon Taf CBC |  | 12 | 24.225 | 21.6643 | 11.82% | G | 0.8943
Milton Keynes Council | 20% TEA in water | 2021 | UC | Milton Keynes Council |  | 12 | 16.4139 | 16.5348 | -0.7314% | G | 1.0074
Socotec Didcot | 20% TEA in water | 2021 | R | Mid and East Antrim Borough Council |  | 12 | 22.3806 | 16.65 | 34.42% | G | 0.74395
Milton Keynes Council | 20% TEA in water | 2021 | R | Milton Keynes Council |  | 11 | 27.4303 | 30.405 | -9.784% | G | 1.1084
Milton Keynes Council | 20% TEA in water | 2021 | R | Milton Keynes Council |  | 11 | 24.4576 | 18.661 | 31.06% | G | 0.76299
Socotec Didcot | 50% TEA in acetone | 2021 | R | Wrexham County Borough Council |  | 12 | 17.9639 | 14.4667 | 24.17% | G | 0.80532
Socotec Didcot | 50% TEA in acetone | 2022 | UB | Torfaen County Borough Council |  | 13 | 13.4333 | 10.0678 | 33.43% | G | 0.74947
Gradko | 50% TEA in Acetone | 2022 | KS | Adur District Council |  | 10 | 29.5763 | 20.7 | 42.88% | G | 0.69988
Gradko | 20% TEA in water | 2022 | R | Blackburn With Darwen Bc |  | 12 | 25.7461 | 19.0667 | 35.03% | G | 0.74056
Socotec Didcot | 50% TEA in acetone | 2022 | R | Bridgend Council |  | 12 | 37.3472 | 26.5583 | 40.62% | G | 0.71112
Socotec Didcot | 50% TEA in Acetone | 2022 | R | Cardiff Council / Shared Regulatory Services |  | 11 | 42.1455 | 33.1182 | 27.26% | G | 0.78581
Socotec Didcot | 50% TEA in Acetone | 2022 | R | Dacorum Borough Council |  | 12 | 23.8306 | 18.225 | 30.76% | G | 0.76477
Gradko | 50% TEA in Acetone | 2022 | UC | Falkirk Council |  | 12 | 31.8031 | 25.9167 | 22.71% | G | 0.81491
Gradko | 50% TEA in Acetone | 2022 | UB | Falkirk Council |  | 9 | 14.6144 | 12.5556 | 16.4% | G | 0.85912
Gradko | 20% TEA in water | 2022 | R | Gedling Borough Council |  | 12 | 31.3167 | 26.1083 | 19.95% | G | 0.83369
Socotec Didcot | 50% TEA in Acetone | 2022 | UB | Gravesham Borough Council |  | 11 | 22.0727 | 18.4545 | 19.61% | G | 0.83608
Socotec Didcot | 50% TEA in Acetone | 2022 | UB | Gravesham Borough Council |  | 11 | 26.1576 | 22.3636 | 16.96% | G | 0.85496
Socotec Didcot | 50% TEA in acetone | 2022 | R | Kingston Upon Hull City Council |  | 12 | 29.7306 | 23.25 | 27.87% | G | 0.78202
Socotec Didcot | 50% TEA in acetone | 2022 | UB | Kingston Upon Hull City Council |  | 12 | 24.2917 | 18 | 34.95% | G | 0.74099
Gradko | 50% TEA in Acetone | 2022 | R | Lb Newham |  | 12 | 29.5789 | 22.9202 | 29.05% | G | 0.77488
Gradko | 50% TEA in acetone | 2022 | SU | Redcar & Cleveland Borough Council |  | 12 | 14.3242 | 9.8833 | 44.93% | G | 0.68998
Gradko | 50% TEA in Acetone | 2022 | R | Worthing Borough Council |  | 9 | 33.2885 | 23.0778 | 44.24% | G | 0.69327
Aberdeen Scientific Services | 20% TEA in water | 2022 | R | Aberdeen City Council |  | 12 | 15.3611 | 14.7333 | 4.261% | G | 0.95913
Aberdeen Scientific Services | 20% TEA in water | 2022 | UB | Aberdeen City Council |  | 12 | 21.1389 | 15.725 | 34.43% | G | 0.74389
Aberdeen Scientific Services | 20% TEA in water | 2022 | R | Aberdeen City Council |  | 11 | 18.8182 | 14.3727 | 30.93% | G | 0.76377
Aberdeen Scientific Services | 20% TEA in water | 2022 | R | Aberdeen City Council |  | 11 | 32.5758 | 22.3091 | 46.02% | G | 0.68484
Aberdeen Scientific Services | 20% TEA in water | 2022 | R | Aberdeen City Council |  | 12 | 34.6389 | 26.075 | 32.84% | G | 0.75277
Aberdeen Scientific Services | 20% TEA in water | 2022 | R | Aberdeen City Council |  | 12 | 33.8611 | 24.0417 | 40.84% | G | 0.71001
Gradko | 20% TEA in water | 2022 | R | Ards And North Down Borough Council |  | 12 | 32.8056 | 21.9583 | 49.4% | G | 0.66935
Gradko | 20% TEA in water | 2022 | R | Bath & North East Somerset |  | 12 | 29.5978 | 24.863 | 19.04% | G | 0.84003
Gradko | 20% TEA in water | 2022 | R | Birmingham City Council |  | 11 | 32.2244 | 23.5545 | 36.81% | G | 0.73095
SOCOTEC Didcot | 50% TEA in acetone | 2022 | UB | City Of York Council |  | 12 | 16.4972 | 12.5354 | 31.61% | G | 0.75985
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | City Of York Council |  | 12 | 24.85 | 19.3102 | 28.69% | G | 0.77707
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | City Of York Council |  | 11 | 22.8788 | 16.6748 | 37.21% | G | 0.72883
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | City Of York Council |  | 11 | 36.6576 | 26.6378 | 37.62% | G | 0.72666
Gradko | 20% TEA in water | 2022 | UB | East Devon District Council |  | 12 | 8.2254 | 6.6558 | 23.58% | G | 0.80918
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | East Suffolk Council |  | 11 | 31.7636 | 22.9129 | 38.63% | G | 0.72136
Gradko | 20% TEA in water | 2022 | R | Gateshead Council |  | 11 | 22.8336 | 19.9927 | 14.21% | G | 0.87558
Gradko | 20% TEA in water | 2022 | R | Gateshead Council |  | 12 | 23.4864 | 20.8375 | 12.71% | G | 0.88722
Gradko | 20% TEA in water | 2022 | R | Gateshead Council |  | 12 | 25.0825 | 22.7783 | 10.12% | G | 0.90814
Gradko | 20% TEA in water | 2022 | R | Gateshead Council |  | 11 | 30.2609 | 23.45 | 29.04% | G | 0.77493
Gradko | 20% TEA in water | 2022 | R | Gateshead Council |  | 9 | 30.9481 | 35.9778 | -13.98% | G | 1.1625
Glasgow Scientific Services | 20% TEA in Water | 2022 | R | Glasgow City Council |  | 12 | 30.4875 | 27.2333 | 11.95% | G | 0.89326
Glasgow Scientific Services | 20% TEA in Water | 2022 | R | Glasgow City Council |  | 11 | 14.2606 | 18.8273 | -24.26% | P | 1.3202
Glasgow Scientific Services | 20% TEA in Water | 2022 | KS | Glasgow City Council |  | 12 | 41.3042 | 38.7417 | 6.614% | G | 0.93796
Glasgow Scientific Services | 20% TEA in Water | 2022 | R | Glasgow City Council |  | 12 | 15.9833 | 21.3333 | -25.08% | P | 1.3347
Glasgow Scientific Services | 20% TEA in Water | 2022 | UB | Glasgow City Council |  | 12 | 13.9694 | 16.6 | -15.85% | P | 1.1883
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | Ipswich Borough Council |  | 11 | 41.9894 | 27.9109 | 50.44% | G | 0.66471
Gradko | 20% TEA in Water | 2022 | R | Lisburn & Castlereagh City Council |  | 12 | 23.6111 | 19.0833 | 23.73% | G | 0.80824
Staffordshire Scientific Services | 20% TEA in water | 2022 | KS | Manchester City Council |  | 12 | 48.9944 | 43.0516 | 13.8% | G | 0.8787
Staffordshire Scientific Services | 20% TEA in water | 2022 | UC | Manchester City Council |  | 12 | 29.0639 | 28.9434 | 0.4162% | G | 0.99585
Staffordshire Scientific Services | 20% TEA in water | 2022 | SI | Manchester City Council |  | 12 | 17.4639 | 15.5857 | 12.05% | G | 0.89245
Gradko | 20% TEA in Water | 2022 | R | Monmouthshire County Council |  | 12 | 34.6189 | 27.9744 | 23.75% | G | 0.80807
Aberdeen Scientific Services | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 11 | 56.2127 | 42.6023 | 31.95% | G | 0.75788
Edinburgh Scientific Services | 50% TEA in acetone | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 52.1836 | 42.4696 | 22.87% | G | 0.81385
Glasgow Scientific Services | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 54.5958 | 42.4696 | 28.55% | G | 0.77789
Gradko | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 52.1385 | 42.4696 | 22.77% | G | 0.81455
Gradko | 50% TEA in acetone | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 52.2325 | 42.4696 | 22.99% | G | 0.81309
Lambeth Scientific Services | 50% TEA in acetone | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 53.2611 | 42.4696 | 25.41% | G | 0.79738
Milton Keynes Council | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 11 | 54.6948 | 42.6023 | 28.38% | G | 0.77891
SOCOTEC Didcot | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 59.0005 | 42.4696 | 38.92% | G | 0.71982
SOCOTEC Didcot | 50% TEA in acetone | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 59.7531 | 42.4696 | 40.7% | G | 0.71075
SOCOTEC Glasgow | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 57.657 | 42.4696 | 35.76% | G | 0.73659
SOCOTEC Glasgow | 50% TEA in acetone | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 55.6778 | 42.4696 | 31.1% | G | 0.76277
Somerset County Council | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 54.7649 | 42.4696 | 28.95% | G | 0.77549
Staffordshire Scientific Services | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 51.1917 | 42.4696 | 20.54% | G | 0.82962
Tayside Scientific Services | 20% TEA in water | 2022 | KS | Marylebone Road Intercomparison |  | 12 | 56.4777 | 42.4696 | 32.98% | G | 0.75197
SOCOTEC Didcot | 20% TEA in water | 2022 | R | New Forest District Council |  | 9 | 27.2852 | 20.3786 | 33.89% | G | 0.74688
SOCOTEC Didcot | 20% TEA in water | 2022 | KS | Nfdc |  | 12 | 37.3444 | 25.4903 | 46.5% | G | 0.68257
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | North East Lincolnshire Council |  | 10 | 46.2667 | 30.96 | 49.44% | G | 0.66916
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | North East Lincolnshire Council |  | 10 | 28.0167 | 27.03 | 3.65% | G | 0.96478
Gradko | 20% TEA in Water | 2022 | UB | Plymouth City Council |  | 12 | 18.3044 | 17.7333 | 3.221% | G | 0.9688
Staffordshire Scientific Services | 20% TEA in water | 2022 | UB | Salford City Council |  | 12 | 23.2653 | 21.7667 | 6.885% | G | 0.93559
Staffordshire Scientific Services | 20% TEA in water | 2022 | B | Salford City Council |  | 10 | 12.8233 | 11.03 | 16.26% | G | 0.86015
Staffordshire Scientific Services | 20% TEA in water | 2022 | R | Salford City Council |  | 12 | 40.0083 | 34.0167 | 17.61% | G | 0.85024
Lambeth Scientific Services | 50% TEA in acetone | 2022 | R | Surrey Heath Borough Council |  | 11 | 21.7576 | 28.6273 | -24% | G | 1.3157
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | Wrexham County Borough Council |  | 12 | 16.3583 | 14.1667 | 15.47% | G | 0.86602
Gradko | 20% TEA in water | 2022 | UC | Belfast City Council |  | 12 | 25.9864 | 19.8775 | 30.73% | G | 0.76492
Gradko | 20% TEA in water | 2022 | R | Belfast City Council |  | 12 | 46.6853 | 36.4525 | 28.07% | G | 0.78081
Gradko | 20% TEA in water | 2022 | R | Belfast City Council |  | 12 | 25.0208 | 21.9517 | 13.98% | G | 0.87734
Gradko | 20% TEA in water | 2022 | R | Belfast City Council |  | 12 | 35.6211 | 27.615 | 28.99% | G | 0.77524
Gradko | 20% TEA in water | 2022 | R | Brighton & Hove City Council |  | 10 | 36.6533 | 22.51 | 62.83% | G | 0.61413
Staffordshire Scientific Services | 20% TEA in water | 2022 | R | Bury Council |  | 11 | 24.3879 | 21.0209 | 16.02% | G | 0.86194
Gradko | 50% TEA in acetone | 2022 | R | City Of London |  | 11 | 60.2618 | 53.9818 | 11.63% | G | 0.89579
Gradko | 50% TEA in acetone | 2022 | UB | City Of London |  | 12 | 28.1387 | 22.7546 | 23.66% | G | 0.80866
Staffordshire Scientific Services | 20% TEA in water | 2022 | R | East Staffordshire Borough Council |  | 10 | 38.8533 | 31.36 | 23.89% | G | 0.80714
Gradko | 20% TEA in water | 2022 | UB | Hertsmere Borough Council |  | 12 | 16.2401 | 15.1583 | 7.137% | G | 0.93339
SOCOTEC Didcot | 50% TEA in Acetone | 2022 | R | Horsham District Council |  | 11 | 25.1909 | 22.0273 | 14.36% | G | 0.87441
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | Leeds City Council |  | 12 | 40.2014 | 29.1667 | 37.83% | G | 0.72551
SOCOTEC Didcot | 50% TEA in acetone | 2022 | KS | Leeds City Council |  | 11 | 32.8848 | 22.7455 | 44.58% | G | 0.69167
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | Leeds City Council |  | 12 | 43.0389 | 34.15 | 26.03% | G | 0.79347
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | Leeds City Council |  | 11 | 40.6818 | 30.3182 | 34.18% | G | 0.74525
SOCOTEC Didcot | 50% TEA in acetone | 2022 | R | Leeds City Council |  | 12 | 30.2583 | 22.1083 | 36.86% | G | 0.73065
SOCOTEC Didcot | 50% TEA in acetone | 2022 | UC | Leeds City Council |  | 12 | 30.0903 | 22.4417 | 34.08% | G | 0.74581
Gradko | 50% TEA in Acetone | 2022 | KS | London Borough Of Croydon |  | 12 | 40.9334 | 36.8367 | 11.12% | G | 0.89992
Gradko | 50% TEA in Acetone | 2022 | R | Royal Borough Of Windsor And Maidenhead |  | 12 | 29.6167 | 25.9917 | 13.95% | G | 0.8776
Gradko | 50% TEA in Acetone | 2022 | R | Royal Borough Of Windsor And Maidenhead |  | 12 | 27.1926 | 27.475 | -1.028% | G | 1.0104
Gradko | 50% TEA in Acetone | 2022 | R | Sandwell Mbc |  | 12 | 33.9889 | 26.7517 | 27.05% | G | 0.78707
Gradko | 50% TEA in Acetone | 2022 | UB | Sandwell Mbc |  | 12 | 21.1629 | 18.9167 | 11.87% | G | 0.89386
Somerset County Council | 20% TEA in water | 2022 | R | South Gloucestershire Council |  | 9 | 31.647 | 26.0556 | 21.46% | G | 0.82332
Somerset County Council | 20% TEA in water | 2022 | R | South Gloucestershire Council |  | 12 | 24.7872 | 21.575 | 14.89% | G | 0.87041
Somerset County Council | 20% TEA in water | 2022 | R | South Gloucestershire Council |  | 12 | 19.9931 | 12.6333 | 58.26% | G | 0.63189
SOCOTEC Didcot | 20% TEA in water | 2022 | R | South Oxfordshire Distric Council |  | 12 | 24.6028 | 18.3925 | 33.77% | G | 0.74758
SOCOTEC Didcot | 20% TEA in water | 2022 | R | South Oxfordshire District Council |  | 12 | 35.5639 | 32.345 | 9.952% | G | 0.90949
Gradko | 20% TEA in water | 2022 | R | Southampton City Council |  | 12 | 36.2985 | 27.8 | 30.57% | G | 0.76587
Gradko | 20% TEA in water | 2022 | UC | Southampton City Council |  | 12 | 27.5841 | 23.9083 | 15.37% | G | 0.86674
Gradko | 20% TEA in water | 2022 | R | Southampton City Council |  | 12 | 33.6009 | 31.0083 | 8.361% | G | 0.92284
Lambeth Scientific Services | 50% TEA in acetone | 2022 | UB | Spelthorne Borough Council |  | 12 | 22.8194 | 19.6167 | 16.33% | G | 0.85965
Lambeth Scientific Services | 50% TEA in acetone | 2022 | UB | Spelthorne Borough Council |  | 10 | 26.1 | 24 | 8.75% | P | 0.91954
Staffordshire Scientific Services | 20% TEA in water | 2022 | UB | Stoke-On-Trent City Council |  | 11 | 23.1939 | 19.8068 | 17.1% | G | 0.85397
SOCOTEC Didcot | 50% TEA in Acetone | 2022 | R | Thanet District Council |  | 12 | 22.5639 | 17.4833 | 29.06% | G | 0.77484
Staffordshire Scientific Services | 20% TEA in water | 2022 | UB | Wigan Council |  | 12 | 20.5958 | 16.9417 | 21.57% | G | 0.82258
Staffordshire Scientific Services | 20% TEA in water | 2022 | R | Wigan Council |  | 12 | 27.2875 | 22.2583 | 22.59% | G | 0.8157
Somerset County Council | 20% TEA in water | 2022 | R | Wiltshire Council |  | 10 | 28.2667 | 28.185 | 0.2898% | G | 0.99711
Somerset County Council | 20% TEA in Water | 2022 | R | Wiltshire Council |  | 10 | 30.2667 | 26.457 | 14.4% | G | 0.87413
Gradko | 20% TEA in water | 2022 | R | Worcestershire |  | 11 | 13.0052 | 12.477 | 4.233% | G | 0.95939
Gradko | 20% TEA in water | 2022 | R | Lancaster City Council |  | 13 | 34.4014 | 27.3462 | 25.8% | G | 0.79491
Gradko | 20% TEA in water | 2022 | R | Lancaster City Council |  | 12 | 27.7624 | 24.0917 | 15.24% | G | 0.86778
Environmental Scientific Groups | 10% TEA in Water | 2001 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.0291
Environmental Scientific Groups | 10% TEA in Water | 2002 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.89524
Edinburgh Scientific Services | 50% TEA in Acetone | 2001 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.92971
Aberdeen CC | 50% TEA in Acetone | 2001 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.3357
Glasgow Scientific Services | 20% TEA in Water | 2002 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.91893
Gradko | 20% TEA in Water | 2002 |  | Overall Factor3 (14 studies) |  |  |  |  | Use |  | 1.0039
Gradko | 50% TEA in Acetone | 2000 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.1988
Gradko | 50% TEA in Acetone | 2001 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 1.4546
Gradko | 50% TEA in Acetone | 2002 |  | Overall Factor3 (12 studies) |  |  |  |  | Use |  | 1.2721
Gradko | 50% TEA in Water | 2000 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.3272
Gradko | 50% TEA in Water | 2001 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 1.3592
Harwell Scientific Services | 50% TEA in acetone | 2000 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.79306
Harwell Scientific Services | 50% TEA in acetone | 2001 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.7841
Harwell Scientific Services | 50% TEA in acetone | 2002 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.84441
Kent Scientific Services | 50% TEA in Acetone | 2001 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.0215
Kent Scientific Services | 50% TEA in Acetone | 2002 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.90323
Lambeth Scientific Services | 50% TEA in Acetone | 2000 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.96671
Lambeth Scientific Services | 50% TEA in Acetone | 2001 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 1.0933
Lambeth Scientific Services | 50% TEA in Acetone | 2002 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 1.1513
South Yorkshire Labs | 50% TEA in Acetone | 2002 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.90177
Jesmond Dene Laboratory | 20% TEA in water | 2000 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.0769
Jesmond Dene Laboratory | 50% TEA in acetone | 2002 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.84348
Gradko | 20% TEA in Water | 2003 |  | Overall Factor3 (12 studies) |  |  |  |  | Use |  | 0.95779
Gradko | 50% TEA in Water | 2003 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 1.0823
Kent Scientific Services | 50% TEA in Acetone | 2003 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.86291
South Yorkshire Labs | 50% TEA in Acetone | 2003 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.90001
Harwell Scientific Services | 50% TEA in Acetone | 2003 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.86502
West Yorkshire Analytical Services | 50% TEA in Acetone | 2003 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.96816
Jesmond Dene Laboratory | 50% TEA in Acetone | 2003 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.99722
Gradko | 50% TEA in Acetone | 2003 |  | Overall Factor3 (19 studies) |  |  |  |  | Use |  | 1.11
Bristol Scientific Services | 20% TEA in Water | 2003 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.89391
Environmental Scientific Groups | 10% TEA in Water | 2003 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 1.0099
Staffordshire Scientific Services | 50% TEA in Water | 2003 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 1.114
Walsall MBC | 50% TEA in Acetone | 2003 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.94032
West Yorkshire Analytical Services | 50% TEA in Acetone | 2002 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.86411
South Yorkshire Labs | 50% TEA in Acetone | 2000 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.85478
South Yorkshire Labs | 50% TEA in Acetone | 2001 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.8705
Lambeth Scientific Services | 50% TEA in Acetone | 2003 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.0524
Edinburgh Scientific Services | 50% TEA in Acetone | 2003 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.96102
West Yorkshire Analytical Services | 50% TEA in Acetone | 2001 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.0182
Gradko | 20% TEA in Water | 2004 |  | Overall Factor3 (11 studies) |  |  |  |  | Use |  | 0.912
Gradko | 50% TEA in Acetone | 2004 |  | Overall Factor3 (27 studies) |  |  |  |  | Use |  | 1.1024
Bristol Scientific Services | 20% TEA in Water | 2004 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.962
Environmental Scientific Groups | 10% TEA in Water | 2004 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.825
West Yorkshire Analytical Services | 50% TEA in Acetone | 2004 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.98258
Edinburgh Scientific Services | 50% TEA in Acetone | 2004 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.93519
Glasgow Scientific Services | 20% TEA in Water | 2004 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.893
Gradko | 50% TEA in Water | 2004 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.99071
Harwell Scientific Services | 50% TEA in Acetone | 2004 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.876
Jesmond Dene Laboratory | 50% TEA in Acetone | 2004 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 1.0152
Lambeth Scientific Services | 50% TEA in Acetone | 2004 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 1.193
Milton Keynes Council | 20% TEA in Water | 2004 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.78072
South Yorkshire Labs | 50% TEA in Acetone | 2004 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.902
Staffordshire Scientific Services | 50% TEA in Water | 2004 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.96
Walsall MBC | 50% TEA in Acetone | 2004 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.90748
Northampton BC | 20% TEA in Water | 2004 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.98348
Cardiff Scientific Services | 50% TEA in Acetone | 2003 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.99
West Yorkshire Analytical Services | 50% TEA in Acetone | 2005 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.789
Harwell Scientific Services | 50% TEA in Acetone | 2005 |  | Overall Factor3 (14 studies) |  |  |  |  | Use |  | 0.882
Gradko | 20% TEA in Water | 2005 |  | Overall Factor3 (14 studies) |  |  |  |  | Use |  | 0.97
Gradko | 50% TEA in Acetone | 2005 |  | Overall Factor3 (14 studies) |  |  |  |  | Use |  | 1.104
Kent Scientific Services | 50% TEA in Acetone | 2005 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.894
Environmental Scientific Groups | 10% TEA in Water | 2005 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 0.796
South Yorkshire Labs | 50% TEA in Acetone | 2005 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.908
Jesmond Dene Laboratory | 50% TEA in Acetone | 2005 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 1.029
Bristol Scientific Services | 20% TEA in Water | 2005 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 1.05
Glasgow Scientific Services | 20% TEA in Water | 2005 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.753
Lambeth Scientific Services | 50% TEA in Acetone | 2005 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 1.236
Staffordshire Scientific Services | 50% TEA in Water | 2005 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 1.054
Milton Keynes Council | 20% TEA in Water | 2005 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.808
Walsall MBC | 50% TEA in Acetone | 2005 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1
Cardiff Scientific Services | 50% TEA in Acetone | 2005 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.74
Northampton BC | 20% TEA in Water | 2006 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.143
Bristol Scientific Services | 20% TEA in Water | 2006 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.902
Cardiff Scientific Services | 50% TEA in Acetone | 2006 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.78
South Yorkshire Labs | 50% TEA in Acetone | 2006 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.915
Glasgow Scientific Services | 20% TEA in Water | 2006 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.959
Gradko | 20% TEA in Water | 2006 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.981
Gradko | 50% TEA in Acetone | 2006 |  | Overall Factor3 (18 studies) |  |  |  |  | Use |  | 1.011
Gradko | 50% TEA in Water | 2006 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.959
Harwell Scientific Services | 50% TEA in Acetone | 2006 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 0.79
Jesmond Dene Laboratory | 50% TEA in Acetone | 2006 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.862
Lambeth Scientific Services | 50% TEA in Acetone | 2006 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 1.276
Milton Keynes Council | 20% TEA in Water | 2006 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.775
Staffordshire Scientific Services | 50% TEA in Water | 2006 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 1.026
West Yorkshire Analytical Services | 50% TEA in Acetone | 2006 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.872
Environmental Scientific Groups | 10% TEA in Water | 2006 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.868
Tayside SS | 20% TEA in Water | 2006 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.782
Edinburgh Scientific Services | 50% TEA in Acetone | 2006 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.969
Kent Scientific Services | 20% TEA in Water | 2006 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.054
Kirklees Council Scientific Services | 50% TEA in Acetone | 2006 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.881
Lancashire CC | 50% TEA in Acetone | 2006 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 1.11
Bristol Scientific Services | 20% TEA in Water | 2007 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.765
Cardiff Scientific Services | 50% TEA in Acetone | 2007 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.843
Environmental Scientific Groups | 10% TEA in Water | 2007 |  | Overall Factor3 (17 studies) |  |  |  |  | Use |  | 0.88
Tayside SS | 20% TEA in Water | 2007 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.91
Edinburgh Scientific Services | 50% TEA in Acetone | 2007 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.95
Glasgow Scientific Services | 20% TEA in Water | 2007 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 1.046
Gradko | 20% TEA in Water | 2007 |  | Overall Factor3 (22 studies) |  |  |  |  | Use |  | 0.893
Gradko | 50% TEA in Acetone | 2007 |  | Overall Factor3 (17 studies) |  |  |  |  | Use |  | 0.99
Harwell Scientific Services | 50% TEA in Acetone | 2007 |  | Overall Factor3 (18 studies) |  |  |  |  | Use |  | 0.817
Jesmond Dene Laboratory | 50% TEA in Acetone | 2007 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.791
Kirklees Council Scientific Services | 50% TEA in Acetone | 2007 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.924
Lambeth Scientific Services | 50% TEA in Acetone | 2007 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 1.07
Lancashire CC | 50% TEA in Acetone | 2007 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 1.002
Milton Keynes Council | 20% TEA in water | 2007 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.718
Northampton BC | 20% TEA in Water | 2007 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 1.046
South Yorkshire Labs | 50% TEA in Acetone | 2007 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.884
Staffordshire Scientific Services | 50% TEA in Water | 2007 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.969
West Yorkshire Analytical Services | 50% TEA in Acetone | 2007 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.889
Gradko | 50% TEA in Water | 2007 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.925
Environmental Scientific Groups | 10% TEA in Water | 2008 |  | Overall Factor3 (15 studies) |  |  |  |  | Use |  | 0.83
Tayside SS | 20% TEA in Water | 2008 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.87
Glasgow Scientific Services | 20% TEA in Water | 2008 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.97
Gradko | 20% TEA in Water | 2008 |  | Overall Factor3 (21 studies) |  |  |  |  | Use |  | 0.92
Gradko | 50% TEA in Acetone | 2008 |  | Overall Factor3 (19 studies) |  |  |  |  | Use |  | 0.94
Gradko | 50% TEA in Water | 2008 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 1.05
Harwell Scientific Services | 50% TEA in Acetone | 2008 |  | Overall Factor3 (14 studies) |  |  |  |  | Use |  | 0.78
Lambeth Scientific Services | 50% TEA in Acetone | 2008 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.98
Milton Keynes Council | 20% TEA in Water | 2008 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.78
Jesmond Dene Laboratory | 50% TEA in Acetone | 2008 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.76
Northampton BC | 20% TEA in Water | 2008 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.85
South Yorkshire Labs | 50% TEA in Acetone | 2008 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.96
Staffordshire Scientific Services | 50% TEA in Water | 2008 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 1.03
West Yorkshire Analytical Services | 50% TEA in Acetone | 2008 |  | Overall Factor3 (11 studies) |  |  |  |  | Use |  | 0.83
Bristol Scientific Services | 20% TEA in water | 2008 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.86
Cardiff Scientific Services | 50% TEA in Acetone | 2008 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.86
Tayside SS | 20% TEA in Water | 2009 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.8
Cardiff Scientific Services | 50% TEA in Acetone | 2009 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.84
Edinburgh Scientific Services | 50% TEA in Acetone | 2008 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.9
Glasgow Scientific Services | 20% TEA in Water | 2009 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 1.23
Gradko | 20% TEA in Water | 2009 |  | Overall Factor3 (34 studies) |  |  |  |  | Use |  | 0.9
Gradko | 50% TEA in Acetone | 2009 |  | Overall Factor3 (16 studies) |  |  |  |  | Use |  | 0.97
Harwell Scientific Services | 50% TEA in Acetone | 2009 |  | Overall Factor3 (27 studies) |  |  |  |  | Use |  | 0.82
Northampton BC | 20% TEA in Water | 2009 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.72
South Yorkshire Labs | 50% TEA in Acetone | 2009 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.91
Staffordshire Scientific Services | 20% TEA in Water | 2009 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.81
West Yorkshire Analytical Services | 50% TEA in Acetone | 2009 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 0.85
Aberdeen CC | 20% TEA in Water | 2008 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.93
Bristol Scientific Services | 20% TEA in Water | 2009 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.79
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.85
Environmental Scientific Groups | 20% TEA in Water | 2009 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.82
Lambeth Scientific Services | 50% TEA in Acetone | 2009 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 1.02
Milton Keynes Council | 20% TEA in Water | 2009 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.79
Walsall MBC | 50% TEA in Acetone | 2009 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 1.17
Aberdeen CC | 20% TEA in Water | 2010 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.88
Bristol Scientific Services | 20% TEA in Water | 2010 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.85
Cardiff Scientific Services | 50% TEA in Acetone | 2010 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.85
Edinburgh Scientific Services | 50% TEA in Acetone | 2010 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.9
Environmental Scientific Groups | 20% TEA in Water | 2010 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.84
Environmental Scientific Groups | 50% TEA in Acetone | 2010 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.75
Glasgow Scientific Services | 20% TEA in Water | 2010 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 1.12
Gradko | 20% TEA in Water | 2010 |  | Overall Factor3 (42 studies) |  |  |  |  | Use |  | 0.92034
Gradko | 50% TEA in Acetone | 2010 |  | Overall Factor3 (16 studies) |  |  |  |  | Use |  | 1.03
Harwell Scientific Services | 50% TEA in Acetone | 2010 |  | Overall Factor3 (20 studies) |  |  |  |  | Use |  | 0.86
Harwell Scientific Services | 20% TEA in Water | 2010 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
Kent Scientific Services | 20% TEA in Water | 2010 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
Kirklees Council Scientific Services | 50% TEA in Acetone | 2010 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
Lambeth Scientific Services | 50% TEA in Acetone | 2010 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 1.06
Lancashire CC | 50% TEA in Acetone | 2010 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.9
Milton Keynes Council | 20% TEA in Water | 2010 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.84
Northampton BC | 20% TEA in Water | 2010 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.73
South Yorkshire Labs | 50% TEA in Acetone | 2010 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.88
Staffordshire Scientific Services | 20% TEA in Water | 2010 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.85
Tayside SS | 20% TEA in Water | 2010 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.79
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 |  | Overall Factor3 (12 studies) |  |  |  |  | Use |  | 0.9
Aberdeen Scientific Services | 20% TEA in water | 2011 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.83
Bristol Scientific Services | 20% TEA in water | 2011 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.82
Cardiff Scientific Services | 50% TEA in acetone | 2011 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.86
Edinburgh Scientific Services | 50% TEA in acetone | 2011 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.84
ESG Glasgow | 20% TEA in water | 2011 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.84
ESG Glasgow | 50% TEA in acetone | 2011 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.87
Exova | 20% TEA in water | 2011 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
Glasgow Scientific Services | 20% TEA in water | 2011 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.94
Gradko | 20% TEA in water | 2011 |  | Overall Factor3 (41 studies) |  |  |  |  | Use |  | 0.9
Gradko | 50% TEA in acetone | 2011 |  | Overall Factor3 (25 studies) |  |  |  |  | Use |  | 0.95
ESG Didcot | 20% TEA in water | 2011 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
ESG Didcot | 50% TEA in acetone | 2011 |  | Overall Factor3 (45 studies) |  |  |  |  | Use |  | 0.83
Kent Scientific Services | 20% TEA in water | 2011 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
Kirklees Council | 50% TEA in acetone | 2011 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.75
Lambeth Scientific Services | 50% TEA in acetone | 2011 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 1.06
Milton Keynes Council | 20% TEA in water | 2011 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.84
Northampton BC | 20% TEA in water | 2011 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.71
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.87
Staffordshire Scientific Services | 20% TEA in water | 2011 |  | Overall Factor3 (15 studies) |  |  |  |  | Use |  | 0.89
Tayside Scientific Services | 20% TEA in water | 2011 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.78
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 0.84
Aberdeen Scientific Services | 20% TEA in water | 2012 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.83
Edinburgh Scientific Services | 50% TEA in acetone | 2012 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.78
ESG Didcot | 20% TEA in water | 2012 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.69
ESG Didcot | 50% TEA in acetone | 2012 |  | Overall Factor3 (38 studies) |  |  |  |  | Use |  | 0.79
ESG Glasgow | 20% TEA in water | 2012 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.71
ESG Glasgow | 50% TEA in acetone | 2012 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.82
Exova | 20% TEA in water | 2012 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.89
Glasgow Scientific Services | 20% TEA in water | 2012 |  | Overall Factor3 (11 studies) |  |  |  |  | Use |  | 0.96
Gradko | 20% TEA in water | 2012 |  | Overall Factor3 (35 studies) |  |  |  |  | Use |  | 0.96
Gradko | 50% TEA in acetone | 2012 |  | Overall Factor3 (21 studies) |  |  |  |  | Use |  | 1.02
Kent Scientific Services | 20% TEA in water | 2012 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.82
Kirklees Council | 50% TEA in acetone | 2012 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.8
Lambeth Scientific Services | 50% TEA in acetone | 2012 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.87
Milton Keynes Council | 20% TEA in water | 2012 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.81
Northampton BC | 20% TEA in water | 2012 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.75
Somerset County Council | 20% TEA in water | 2012 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.95
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2012 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.8
Staffordshire Scientific Services | 20% TEA in water | 2012 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 0.86
Tayside Scientific Services | 20% TEA in water | 2012 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.84
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.78
Aberdeen Scientific Services | 20% TEA in water | 2013 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.83
Edinburgh Scientific Services | 50% TEA in acetone | 2013 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.75
ESG Didcot | 20% TEA in water | 2013 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.85
ESG Didcot | 50% TEA in acetone | 2013 |  | Overall Factor3 (44 studies) |  |  |  |  | Use |  | 0.81
ESG Glasgow | 20% TEA in water | 2013 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.72
ESG Glasgow | 50% TEA in acetone | 2013 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.73
Exova | 20% TEA in water | 2013 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.91
Glasgow Scientific Services | 20% TEA in water | 2013 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.99
Gradko | 20% TEA in water | 2013 |  | Overall Factor3 (36 studies) |  |  |  |  | Use |  | 0.95
Gradko | 50% TEA in acetone | 2013 |  | Overall Factor3 (20 studies) |  |  |  |  | Use |  | 1.01
Kent Scientific Services | 20% TEA in water | 2013 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
Kirklees Council | 50% TEA in acetone | 2013 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.74
Lambeth Scientific Services | 50% TEA in acetone | 2013 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.87
Milton Keynes Council | 20% TEA in water | 2013 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.84
Northampton BC | 20% TEA in water | 2013 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.73
Somerset County Council | 20% TEA in water | 2013 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.9
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2013 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.84
Staffordshire Scientific Services | 20% TEA in water | 2013 |  | Overall Factor3 (11 studies) |  |  |  |  | Use |  | 0.87
Tayside Scientific Services | 20% TEA in water | 2013 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.79
Aberdeen Scientific Services | 20% TEA in water | 2014 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.86
Edinburgh Scientific Services | 50% TEA in acetone | 2014 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.74
ESG Didcot | 20% TEA in water | 2014 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.79
ESG Didcot | 50% TEA in acetone | 2014 |  | Overall Factor3 (31 studies) |  |  |  |  | Use |  | 0.81
ESG Glasgow | 20% TEA in water | 2014 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.7
ESG Glasgow | 50% TEA in acetone | 2014 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.76
Glasgow Scientific Services | 20% TEA in water | 2014 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.83
Gradko | 20% TEA in water | 2014 |  | Overall Factor3 (22 studies) |  |  |  |  | Use |  | 0.92
Gradko | 50% TEA in acetone | 2014 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.98
Kirklees Council | 50% TEA in acetone | 2014 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.74
Lambeth Scientific Services | 50% TEA in acetone | 2014 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.88
Milton Keynes Council | 20% TEA in water | 2014 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.75
Northampton BC | 20% TEA in water | 2014 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.77
Somerset County Council | 20% TEA in water | 2014 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.89
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2014 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.67
Staffordshire Scientific Services | 20% TEA in water | 2014 |  | Overall Factor3 (15 studies) |  |  |  |  | Use |  | 0.83
Tayside Scientific Services | 20% TEA in water | 2014 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.79
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.75
Aberdeen Scientific Services | 20% TEA in water | 2015 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.84
Edinburgh Scientific Services | 50% TEA in acetone | 2015 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.76
ESG Didcot | 20% TEA in water | 2015 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.81
ESG Didcot | 50% TEA in acetone | 2015 |  | Overall Factor3 (29 studies) |  |  |  |  | Use |  | 0.79
ESG Glasgow | 20% TEA in water | 2015 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
ESG Glasgow | 50% TEA in acetone | 2015 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
Glasgow Scientific Services | 20% TEA in water | 2015 |  | Overall Factor3 (12 studies) |  |  |  |  | Use |  | 0.98
Gradko | 20% TEA in water | 2015 |  | Overall Factor3 (30 studies) |  |  |  |  | Use |  | 0.87
Gradko | 50% TEA in acetone | 2015 |  | Overall Factor3 (15 studies) |  |  |  |  | Use |  | 0.96
Kirklees Council | 50% TEA in acetone | 2015 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.76
Lambeth Scientific Services | 50% TEA in acetone | 2015 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 1
Milton Keynes Council | 20% TEA in water | 2015 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.76
Northampton BC | 20% TEA in water | 2015 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.66
Somerset County Council | 20% TEA in water | 2015 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.9
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.84
Staffordshire Scientific Services | 20% TEA in water | 2015 |  | Overall Factor3 (16 studies) |  |  |  |  | Use |  | 0.85
Tayside Scientific Services | 20% TEA in water | 2015 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.79
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.77
Aberdeen Scientific Services | 20% TEA in water | 2016 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.85
Edinburgh Scientific Services | 50% TEA in acetone | 2016 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.77
ESG Didcot | 20% TEA in water | 2016 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.83
ESG Didcot | 50% TEA in acetone | 2016 |  | Overall Factor3 (38 studies) |  |  |  |  | Use |  | 0.77
ESG Glasgow | 20% TEA in water | 2016 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.79
ESG Glasgow | 50% TEA in acetone | 2016 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
Glasgow Scientific Services | 20% TEA in water | 2016 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.97
Gradko | 20% TEA in water | 2016 |  | Overall Factor3 (32 studies) |  |  |  |  | Use |  | 0.92
Gradko | 50% TEA in acetone | 2016 |  | Overall Factor3 (19 studies) |  |  |  |  | Use |  | 1.01
Lambeth Scientific Services | 50% TEA in acetone | 2016 |  | Overall Factor3 (6 study) |  |  |  |  | Use |  | 1.06
Milton Keynes Council | 20% TEA in water | 2016 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.66
Northampton BC | 20% TEA in water | 2016 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.85
Somerset County Council | 20% TEA in water | 2016 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.88
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2016 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.83
Staffordshire Scientific Services | 20% TEA in water | 2016 |  | Overall Factor3 (13 studies) |  |  |  |  | Use |  | 0.91
Tayside Scientific Services | 20% TEA in water | 2016 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.82
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.75
Aberdeen Scientific Services | 20% TEA in water | 2017 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.79
Edinburgh Scientific Services | 50% TEA in acetone | 2017 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.81
ESG Didcot | 20% TEA in water | 2017 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.74
ESG Didcot | 50% TEA in acetone | 2017 |  | Overall Factor3 (30 studies) |  |  |  |  | Use |  | 0.77
ESG Glasgow | 20% TEA in water | 2017 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.8
ESG Glasgow | 50% TEA in acetone | 2017 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
Glasgow Scientific Services | 20% TEA in water | 2017 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.9
Gradko | 20% TEA in water | 2017 |  | Overall Factor3 (39 studies) |  |  |  |  | Use |  | 0.87
Gradko | 50% TEA in acetone | 2017 |  | Overall Factor3 (25 studies) |  |  |  |  | Use |  | 0.96
Lambeth Scientific Services | 50% TEA in acetone | 2017 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.93
Milton Keynes Council | 20% TEA in water | 2017 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.76
Somerset County Council | 20% TEA in water | 2017 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.77
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2017 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.87
Staffordshire Scientific Services | 20% TEA in water | 2017 |  | Overall Factor3 (19 studies) |  |  |  |  | Use |  | 0.88
Tayside Scientific Services | 20% TEA in water | 2017 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.72
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.77
Aberdeen Scientific Services | 20% TEA in water | 2018 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.82
Edinburgh Scientific Services | 50% TEA in acetone | 2018 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.9
Glasgow Scientific Services | 20% TEA in water | 2018 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.9
Gradko | 20% TEA in water | 2018 |  | Overall Factor3 (40 studies) |  |  |  |  | Use |  | 0.92
Gradko | 50% TEA in acetone | 2018 |  | Overall Factor3 (18 studies) |  |  |  |  | Use |  | 0.89
Lambeth Scientific Services | 50% TEA in acetone | 2018 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 1.04
Milton Keynes Council | 20% TEA in water | 2018 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.76
SOCOTEC Didcot | 20% TEA in water | 2018 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.75
SOCOTEC Didcot | 50% TEA in acetone | 2018 |  | Overall Factor3 (30 studies) |  |  |  |  | Use |  | 0.77
SOCOTEC Glasgow | 20% TEA in water | 2018 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.93
SOCOTEC Glasgow | 50% TEA in acetone | 2018 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.97
Somerset County Council | 20% TEA in water | 2018 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.91
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.94
Staffordshire Scientific Services | 20% TEA in water | 2018 |  | Overall Factor3 (16 studies) |  |  |  |  | Use |  | 0.88
Tayside Scientific Services | 20% TEA in water | 2018 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.8
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 |  | Overall Factor3 (8 studies) |  |  |  |  | Use |  | 0.8
Aberdeen Scientific Services | 20% TEA in water | 2019 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.81
Edinburgh Scientific Services | 50% TEA in acetone | 2019 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.85
Glasgow Scientific Services | 20% TEA in water | 2019 |  | Overall Factor3 (11 studies) |  |  |  |  | Use |  | 0.87
Gradko | 20% TEA in water | 2019 |  | Overall Factor3 (31 studies) |  |  |  |  | Use |  | 0.91
Gradko | 50% TEA in acetone | 2019 |  | Overall Factor3 (29 studies) |  |  |  |  | Use |  | 0.89
Lambeth Scientific Services | 50% TEA in acetone | 2019 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.91
Milton Keynes Council | 20% TEA in water | 2019 |  | Overall Factor3 (2 studies) |  |  |  |  | Use |  | 0.84
SOCOTEC Didcot | 20% TEA in water | 2019 |  | Overall Factor3 (12 studies) |  |  |  |  | Use |  | 0.77
SOCOTEC Didcot | 50% TEA in acetone | 2019 |  | Overall Factor3 (42 studies) |  |  |  |  | Use |  | 0.75
SOCOTEC Glasgow | 20% TEA in water | 2019 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.79
SOCOTEC Glasgow | 50% TEA in acetone | 2019 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.75
Somerset County Council | 20% TEA in water | 2019 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.83
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2019 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 1.01
Staffordshire Scientific Services | 20% TEA in water | 2019 |  | Overall Factor3 (17 studies) |  |  |  |  | Use |  | 0.93
Tayside Scientific Services | 20% TEA in water | 2019 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.8
West Yorkshire Analytical Services | 50% TEA in acetone | 2019 |  | Overall Factor3 (3 studies) |  |  |  |  | Use |  | 0.8
Aberdeen Scientific Services | 20% TEA in water | 2020 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.78
Edinburgh Scientific Services | 50% TEA in acetone | 2020 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.85
Glasgow Scientific Services | 20% TEA in water | 2020 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.95
Gradko | 20% TEA in water | 2020 |  | Overall Factor3 (27 studies) |  |  |  |  | Use |  | 0.81
Gradko | 50% TEA in acetone | 2020 |  | Overall Factor3 (22 studies) |  |  |  |  | Use |  | 0.84
Lambeth Scientific Services | 50% TEA in acetone | 2020 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.96
Milton Keynes Council | 20% TEA in water | 2020 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.83
SOCOTEC Didcot | 20% TEA in water | 2020 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.74
SOCOTEC Didcot | 50% TEA in acetone | 2020 |  | Overall Factor3 (24 studies) |  |  |  |  | Use |  | 0.76
SOCOTEC Glasgow | 20% TEA in water | 2020 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.79
SOCOTEC Glasgow | 50% TEA in acetone | 2020 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.79
Somerset County Council | 20% TEA in water | 2020 |  | Overall Factor3 (10 studies) |  |  |  |  | Use |  | 0.85
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2020 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
Staffordshire Scientific Services | 20% TEA in water | 2020 |  | Overall Factor3 (15 studies) |  |  |  |  | Use |  | 0.85
Tayside Scientific Services | 20% TEA in water | 2020 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.75
Aberdeen Scientific Services | 20% TEA in water | 2021 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.77
Edinburgh Scientific Services | 50% TEA in acetone | 2021 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.85
Glasgow Scientific Services | 20% TEA in water | 2021 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 1.11
Gradko | 20% TEA in water | 2021 |  | Overall Factor3 (34 studies) |  |  |  |  | Use |  | 0.84
Gradko | 50% TEA in acetone | 2021 |  | Overall Factor3 (16 studies) |  |  |  |  | Use |  | 0.82
Lambeth Scientific Services | 50% TEA in acetone | 2021 |  | Overall Factor3 (9 studies) |  |  |  |  | Use |  | 0.95
Milton Keynes Council | 20% TEA in water | 2021 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.89
SOCOTEC Didcot | 20% TEA in water | 2021 |  | Overall Factor3 (15 studies) |  |  |  |  | Use |  | 0.77
SOCOTEC Didcot | 50% TEA in acetone | 2021 |  | Overall Factor3 (28 studies) |  |  |  |  | Use |  | 0.77
SOCOTEC Glasgow | 20% TEA in water | 2021 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
SOCOTEC Glasgow | 50% TEA in acetone | 2021 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.8
Somerset County Council | 20% TEA in water | 2021 |  | Overall Factor3 (11 studies) |  |  |  |  | Use |  | 0.86
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2021 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
Staffordshire Scientific Services | 20% TEA in water | 2021 |  | Overall Factor3 (16 studies) |  |  |  |  | Use |  | 0.85
Tayside Scientific Services | 20% TEA in water | 2021 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.77
Aberdeen Scientific Services | 20% TEA in water | 2022 |  | Overall Factor3 (7 studies) |  |  |  |  | Use |  | 0.76
Edinburgh Scientific Services | 50% TEA in acetone | 2022 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.81
Glasgow Scientific Services | 20% TEA in water | 2022 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 1.05
Gradko | 20% TEA in water | 2022 |  | Overall Factor3 (27 studies) |  |  |  |  | Use |  | 0.83
Gradko | 50% TEA in acetone | 2022 |  | Overall Factor3 (14 studies) |  |  |  |  | Use |  | 0.82
Lambeth Scientific Services | 50% TEA in acetone | 2022 |  | Overall Factor3 (4 studies) |  |  |  |  | Use |  | 0.95
Milton Keynes Council | 20% TEA in water | 2022 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.78
SOCOTEC Didcot | 20% TEA in water | 2022 |  | Overall Factor3 (5 studies) |  |  |  |  | Use |  | 0.76
SOCOTEC Didcot | 50% TEA in acetone | 2022 |  | Overall Factor3 (26 studies) |  |  |  |  | Use |  | 0.76
SOCOTEC Glasgow | 20% TEA in water | 2022 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.74
SOCOTEC Glasgow | 50% TEA in acetone | 2022 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.76
Somerset County Council | 20% TEA in water | 2022 |  | Overall Factor3 (6 studies) |  |  |  |  | Use |  | 0.82
Staffordshire Scientific Services | 20% TEA in water | 2022 |  | Overall Factor3 (12 studies) |  |  |  |  | Use |  | 0.87
Tayside Scientific Services | 20% TEA in water | 2022 |  | Overall Factor3 (1 study) |  |  |  |  | Use |  | 0.75
1 For Casella Stanger/Bureau Veritas (NOT Bureau Veritas Labs) use Gradko 50% TEA in Acetone. For Casella Seal/GMSS/Casella CRE/Bureau Veritas Labs/Eurofins/ use Environmental Scientific Groups. From 2011 for Environmental Scientific Groups use ESG Glasgow. From 2011 for Harwell Scientific Services use ESG Didcot. For 2017 for SOCOTEC use ESG Didcot, as name changed mid year. For 2018 SOCOTEC entered as Didcot and Glasgow. Glasgow analysis lab moved to Didcot mid 2018. For Staffordshire CC SS/Staffordshire County Analyst use Staffordshire Scientific Services. For Bodycote Health Sciences and Clyde Analytical Laboratories use Exova. For Rotherham MBC use South Yorkshire Labs. For Dundee CC use Tayside SS. For Leicester Scientific Services use Staffordshire Scientific Services. For South Yorkshire Air Quality Samplers use South Yorkshire Labs. As of January 2010 sampler body changed. As of April 2010 sampler cap changed. Lancashire County Analysts withdrew from the Field intercomparison at the end of 2010. No submissions were supplied in 2011. Walsall MBC closed in March 2011. Bristol Scientific Services closed at the end of 2011. Somerset County Council did not start the Marylebone road intercomparison until June 2012. Exova stopped providing diffusion tubes at the end of 2013. Kent Scientific Services stopped providing diffusion tubes at the end of 2013. Kirklees Council stopped providing diffusion tubes in the middle of 2016. Northampton BC stopped providing diffusion tubes in 2017. 2018 - Gradko preparation method for 50% TEA in acetone tubes changed from pipetting to dipping of grids to coat them in TEA. West Yorkshire Analytical Services stopped providing diffusion tubes in 2020. South Yorkshire Air Quality Samplers stopped providing diffusion tubes in 2022.
2 In this situation it would be reasonable to use data from the nearest year.
3 Overall factors have been calculated using orthogonal regression to allow for uncertainty in both the automatic monitor and diffusion tube. The uncertainty of the diffusion tube has been assumed to be double that of the automatic monitor.
4 If you have your own co-location study, please send your data to us, so that it can be included here. If this is not possible, but you wish to combine these factors with your own, select and copy the relevant data from this spreadsheet and paste them into a new one (otherwise your calculations will include hidden data). Then add your own data and calculate the bias. To obtain a new correction factor that includes your data, average the bias (B) values, expressed as a factor, i.e. -16% is -0.16. Next add 1 to this value, e.g. -0.16 + 1.00 = 0.84 in this example, then take the inverse to give the bias adjustment factor 1/0.84 = 1.19. (This will not be exactly the same as the correction factor calculated using orthogonal regression as used in this spreadsheet, but will be reasonably close). | [5 empty] | To add data download a questionnaire
5 Where an annual data set falls into two years it has been ascribed to the year in which most of the data has fallen.
6 Tube precision is determined as follows: G = Good precision - coefficient of variation (CV) of diffusion tube replicates is considered G when the CV of eight or more periods is less than 20%, and the average CV of all monitoring periods is less than 10%; P = Poor precision - CV of four or more periods >20% and/or average CV >10%; S = Single tube, therefore not applicable; na = not available.

## Revisions
Diffusion Tube Bias Adjustment Factors 03/23 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/22) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2022 | 7 | 0.76
Edinburgh Scientific Services | 50% TEA in acetone | 2022 | 1 | 0.81
Glasgow Scientific Services | 20% TEA in water | 2022 | 6 | 1.05
Gradko | 20% TEA in water | 2022 | 27 | 0.83
Gradko | 50% TEA in acetone | 2022 | 14 | 0.82
Lambeth Scientific Services | 50% TEA in acetone | 2022 | 4 | 0.95
Milton Keynes Council | 20% TEA in water | 2022 | 1 | 0.78
SOCOTEC Didcot | 20% TEA in water | 2022 | 5 | 0.76
SOCOTEC Didcot | 50% TEA in acetone | 2022 | 26 | 0.76
SOCOTEC Glasgow | 20% TEA in water | 2022 | 1 | 0.74
SOCOTEC Glasgow | 50% TEA in acetone | 2022 | 1 | 0.76
Somerset County Council | 20% TEA in water | 2022 | 6 | 0.82
Staffordshire Scientific Services | 20% TEA in water | 2022 | 12 | 0.87
Tayside Scientific Services | 20% TEA in water | 2022 | 1 | 0.75
Number of Studies Included |  |  | 112
Factors for 2022 are very similar to 2021 with the exception of Milton Keynes Council, which only has 1 study in 2022 so far.
Glasgow Scientific Services factor should be treated with caution as there is not a good linear relationship between the diffusion tube concentrations and the automatic analyser concentrations.
The diffusion tubes under read the automatic analyser at concentrations below 21 mg/m3 and over read the automatic analyser at concentrations greater than 27 mg/m3.
There is also a large spread in bias factors (0.78 - 1.33).
2 studies were submitted for 2022 data that could not be included due to less than 9 valid data periods or poor data quality.
The National Diffusion Tube Bias Adjustment Factor Spreadsheet will be next updated at the end of March 2023
Local Air Quality Management Helpdesk
Telephone: | 0800 0327953
E-mail: | LAQMHelpdesk@bureauveritas.com
Previous Revisions Record:
Diffusion Tube Bias Adjustment Factors 09/22 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/22) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2021 | 7 | 0 | 7 | 0.77 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | 6 | 0 | 6 | 0.85 | 0
Glasgow Scientific Services | 20% TEA in water | 2021 | 7 | 0 | 7 | 1.11 | 0
Gradko | 20% TEA in water | 2021 | 34 | 0 | 34 | 0.84 | 0
Gradko | 50% TEA in acetone | 2021 | 16 | 0 | 16 | 0.82 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2021 | 9 | 0 | 9 | 0.95 | 0
Milton Keynes Council | 20% TEA in water | 2021 | 1 | 3 | 4 | 0.89 | 0.15
SOCOTEC Didcot | 20% TEA in water | 2021 | 13 | 2 | 15 | 0.77 | 0.01
SOCOTEC Didcot | 50% TEA in acetone | 2021 | 25 | 3 | 28 | 0.77 | -0.01
SOCOTEC Glasgow | 20% TEA in water | 2021 | 1 | 0 | 1 | 0.78 | 0
SOCOTEC Glasgow | 50% TEA in acetone | 2021 | 1 | 0 | 1 | 0.8 | 0
Somerset County Council | 20% TEA in water | 2021 | 11 | 0 | 11 | 0.86 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2021 | 1 | 0 | 1 | 0.77 | 0
Staffordshire Scientific Services | 20% TEA in water | 2021 | 15 | 1 | 16 | 0.85 | 0
Tayside Scientific Services | 20% TEA in water | 2021 | 1 | 0 | 1 | 0.77 | 0
Number of Studies Included |  |  | 148 | 9 | 157
The bias adjustment factor for Glasgow Scientific Services should be treated with caution as 4 of the 7 studies gave a bias adjustment factor greater than 1.17 when the factor is generally expected to be <1
The large change in factor for Milton Keynes Council was due to the number of studies going from 1 to 4.
0 studies were updated due to ratified data becoming available since the June 2022 release.
Diffusion Tube Bias Adjustment Factors 06/22 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/22) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2021 | 7 | 0 | 7 | 0.77 | 0.02
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | 1 | 5 | 6 | 0.85 | -0.02
Glasgow Scientific Services | 20% TEA in water | 2021 | 6 | 1 | 7 | 1.11 | -0.01
Gradko | 20% TEA in water | 2021 | 32 | 2 | 34 | 0.84 | 0
Gradko | 50% TEA in acetone | 2021 | 14 | 2 | 16 | 0.82 | -0.01
Lambeth Scientific Services | 50% TEA in acetone | 2021 | 5 | 4 | 9 | 0.95 | -0.02
Milton Keynes Council | 20% TEA in water | 2021 | 1 | 0 | 1 | 0.74 | 0
SOCOTEC Didcot | 20% TEA in water | 2021 | 5 | 8 | 13 | 0.76 | 0
SOCOTEC Didcot | 50% TEA in acetone | 2021 | 23 | 2 | 25 | 0.78 | 0
SOCOTEC Glasgow | 20% TEA in water | 2021 | 1 | 0 | 1 | 0.78 | 0
SOCOTEC Glasgow | 50% TEA in acetone | 2021 | 1 | 0 | 1 | 0.8 | 0
Somerset County Council | 20% TEA in water | 2021 | 2 | 9 | 11 | 0.86 | 0.09
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2021 | 1 | 0 | 1 | 0.77 | 0
Staffordshire Scientific Services | 20% TEA in water | 2021 | 14 | 1 | 15 | 0.85 | -0.01
Tayside Scientific Services | 20% TEA in water | 2021 | 1 | 0 | 1 | 0.77 | 0
Number of Studies Included |  |  | 114 | 34 | 148
The bias adjustment factor for Glasgow Scientific Services should be treated with caution as 4 of the 7 studies gave a bias adjustment factor greater than 1.17 when the factor is generally expected to be <1
6 studies were updated due to ratified data becoming available since the March 2022 release. This only affected the Aberdeen Scientific Services Factor.
2 additional studies were submitted but could not be included due to less than 9 valid data periods or poor data quality. In total 5 such studies were submitted in 2022.
Diffusion Tube Bias Adjustment Factors 03/22 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/22) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2021 | 7 | 0.75
Edinburgh Scientific Services | 50% TEA in acetone | 2021 | 1 | 0.87
Glasgow Scientific Services | 20% TEA in water | 2021 | 6 | 1.12
Gradko | 20% TEA in water | 2021 | 32 | 0.84
Gradko | 50% TEA in acetone | 2021 | 14 | 0.83
Lambeth Scientific Services | 50% TEA in acetone | 2021 | 5 | 0.97
Milton Keynes Council | 20% TEA in water | 2021 | 1 | 0.74
SOCOTEC Didcot | 20% TEA in water | 2021 | 5 | 0.76
SOCOTEC Didcot | 50% TEA in acetone | 2021 | 23 | 0.78
SOCOTEC Glasgow | 20% TEA in water | 2021 | 1 | 0.78
SOCOTEC Glasgow | 50% TEA in acetone | 2021 | 1 | 0.8
Somerset County Council | 20% TEA in water | 2021 | 2 | 0.77
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2021 | 1 | 0.77
Staffordshire Scientific Services | 20% TEA in water | 2021 | 14 | 0.86
Tayside Scientific Services | 20% TEA in water | 2021 | 1 | 0.77
Number of Studies Included |  |  | 114
The bias adjustment factor for Glasgow Scientific Services should be treated with caution as 4 of the 6 studies gave a bias adjustment factor greater than 1.17 when the factor is generally expected to be <1
3 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 09/21 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/21) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2020 | 7 | 0 | 7 | 0.78 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | 5 | 0 | 5 | 0.85 | 0
Glasgow Scientific Services | 20% TEA in water | 2020 | 9 | 0 | 9 | 0.95 | 0
Gradko | 20% TEA in water | 2020 | 27 | 0 | 27 | 0.81 | 0
Gradko | 50% TEA in acetone | 2020 | 21 | 1 | 22 | 0.84 | 0.01
Lambeth Scientific Services | 50% TEA in acetone | 2020 | 10 | 0 | 10 | 0.96 | 0.01
Milton Keynes Council | 20% TEA in water | 2020 | 4 | 0 | 4 | 0.83 | 0
SOCOTEC Didcot | 20% TEA in water | 2020 | 6 | 0 | 6 | 0.74 | 0
SOCOTEC Didcot | 50% TEA in acetone | 2020 | 24 | 0 | 24 | 0.76 | 0
SOCOTEC Glasgow | 20% TEA in water | 2020 | 1 | 0 | 1 | 0.79 | 0
SOCOTEC Glasgow | 50% TEA in acetone | 2020 | 1 | 0 | 1 | 0.79 | 0
Somerset County Council | 20% TEA in water | 2020 | 10 | 0 | 10 | 0.85 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2020 | 1 | 0 | 1 | 0.77 | 0
Staffordshire Scientific Services | 20% TEA in water | 2020 | 15 | 0 | 15 | 0.85 | 0
Tayside Scientific Services | 20% TEA in water | 2020 | 1 | 0 | 1 | 0.75 | 0
Number of Studies Included |  |  | 142 | 1 | 143
The Lambeth factor was incorrectly entered in the June release of the database and should have read 0.96, hence the chance on 0.01 in the table above.
2 studies were updated due to ratified data becoming available since the June 2021 release.
Although studies from both Medway and Fareham Councils had low bias adjustment factors, 0.51 & 0.56 respectively, these were included as they had passed the required QA checks
28 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 06/21 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/21) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2020 | 7 | 0 | 7 | 0.78 | 0.01
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | 2 | 3 | 5 | 0.85 | -0.03
Glasgow Scientific Services | 20% TEA in water | 2020 | 10 | -1 | 9 | 0.95 | -0.01
Gradko | 20% TEA in water | 2020 | 18 | 9 | 27 | 0.81 | 0
Gradko | 50% TEA in acetone | 2020 | 14 | 7 | 21 | 0.83 | 0.01
Lambeth Scientific Services | 50% TEA in acetone | 2020 | 5 | 5 | 10 | 0.95 | -0.01
Milton Keynes Council | 20% TEA in water | 2020 | 4 | 0 | 4 | 0.83 | 0
SOCOTEC Didcot | 20% TEA in water | 2020 | 6 | 0 | 6 | 0.74 | 0
SOCOTEC Didcot | 50% TEA in acetone | 2020 | 22 | 2 | 24 | 0.76 | -0.01
SOCOTEC Glasgow | 20% TEA in water | 2020 | 1 | 0 | 1 | 0.79 | 0
SOCOTEC Glasgow | 50% TEA in acetone | 2020 | 1 | 0 | 1 | 0.79 | 0
Somerset County Council | 20% TEA in water | 2020 | 2 | 8 | 10 | 0.85 | 0.09
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2020 | 1 | 0 | 1 | 0.77 | 0
Staffordshire Scientific Services | 20% TEA in water | 2020 | 15 | 0 | 15 | 0.85 | 0
Tayside Scientific Services | 20% TEA in water | 2020 | 1 | 0 | 1 | 0.75 | 0
Number of Studies Included |  |  | 109 | 33 | 142
12 studies were updated due to ratified data becoming available since the March 2021 release.
1 of the these ratified resubmissions (Glasgow Scientific Services) reduced the data capture of the automatic analyser to below 9 periods of data with a data capture of >=95%. Therefore this study was removed from the database.
The change in the Somerset County Council Bias Factor is due to the inclusion of 8 more studies, which brings the BF in line with 2019.
Although studies from both Medway and Fareham Councils had low bias adjustment factors, 0.51 & 0.56 respectively, these were included as they had passed the required QA checks
20 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 03/21 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/21) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2020 | 7 | 0.77
Edinburgh Scientific Services | 50% TEA in acetone | 2020 | 2 | 0.88
Glasgow Scientific Services | 20% TEA in water | 2020 | 10 | 0.96
Gradko | 20% TEA in water | 2020 | 18 | 0.81
Gradko | 50% TEA in acetone | 2020 | 14 | 0.82
Lambeth Scientific Services | 50% TEA in acetone | 2020 | 5 | 0.96
Milton Keynes Council | 20% TEA in water | 2020 | 4 | 0.83
SOCOTEC Didcot | 20% TEA in water | 2020 | 6 | 0.74
SOCOTEC Didcot | 50% TEA in acetone | 2020 | 22 | 0.77
SOCOTEC Glasgow | 20% TEA in water | 2020 | 1 | 0.79
SOCOTEC Glasgow | 50% TEA in acetone | 2020 | 1 | 0.79
Somerset County Council | 20% TEA in water | 2020 | 2 | 0.76
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2020 | 1 | 0.77
Staffordshire Scientific Services | 20% TEA in water | 2020 | 15 | 0.85
Tayside Scientific Services | 20% TEA in water | 2020 | 1 | 0.75
Number of Studies Included |  |  | 109
13 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 09/20 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/20) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2019 | 6 | 0 | 6 | 0.81 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | 6 | 0 | 6 | 0.85 | 0
Glasgow Scientific Services | 20% TEA in water | 2019 | 11 | 0 | 11 | 0.87 | 0
Gradko | 20% TEA in water | 2019 | 30 | 1 | 31 | 0.91 | -0.01
Gradko | 50% TEA in acetone | 2019 | 29 | 0 | 29 | 0.89 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2019 | 9 | 0 | 9 | 0.91 | 0
Milton Keynes Council | 20% TEA in water | 2019 | 2 | 0 | 2 | 0.84 | 0
SOCOTEC Didcot | 20% TEA in water | 2019 | 12 | 0 | 12 | 0.77 | 0
SOCOTEC Didcot | 50% TEA in acetone | 2019 | 38 | 4 | 42 | 0.75 | 0
SOCOTEC Glasgow | 20% TEA in water | 2019 | 1 | 0 | 1 | 0.79 | 0
SOCOTEC Glasgow | 50% TEA in acetone | 2019 | 1 | 0 | 1 | 0.75 | 0
Somerset County Council | 20% TEA in water | 2019 | 9 | 0 | 9 | 0.83 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2019 | 3 | 0 | 3 | 1.01 | 0
Staffordshire Scientific Services | 20% TEA in water | 2019 | 17 | 0 | 17 | 0.93 | 0
Tayside Scientific Services | 20% TEA in water | 2019 | 1 | 0 | 1 | 0.8 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2019 | 3 | 0 | 3 | 0.8 | 0
Number of Studies Included |  |  | 178 | 5 | 183
7 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
0 studies were updated due to ratified data becoming available since the June 2020 release.
1 study was not included due to unrepresentative low concentrations.
Diffusion Tube Bias Adjustment Factors 06/20 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/20) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2019 | 6 | 0 | 6 | 0.81 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | 1 | 5 | 6 | 0.85 | -0.02
Glasgow Scientific Services | 20% TEA in water | 2019 | 4 | 7 | 11 | 0.87 | 0.01
Gradko | 20% TEA in water | 2019 | 27 | 3 | 30 | 0.92 | -0.01
Gradko | 50% TEA in acetone | 2019 | 8 | 21 | 29 | 0.89 | 0.02
Lambeth Scientific Services | 50% TEA in acetone | 2019 | 1 | 8 | 9 | 0.91 | 0.06
Milton Keynes Council | 20% TEA in water | 2019 | 2 | 0 | 2 | 0.84 | 0
SOCOTEC Didcot | 20% TEA in water | 2019 | 6 | 6 | 12 | 0.77 | 0.01
SOCOTEC Didcot | 50% TEA in acetone | 2019 | 24 | 14 | 38 | 0.75 | 0
SOCOTEC Glasgow | 20% TEA in water | 2019 | 1 | 0 | 1 | 0.79 | 0
SOCOTEC Glasgow | 50% TEA in acetone | 2019 | 1 | 0 | 1 | 0.75 | 0
Somerset County Council | 20% TEA in water | 2019 | 2 | 7 | 9 | 0.83 | 0.05
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2019 | 1 | 2 | 3 | 1.01 | 0.23
Staffordshire Scientific Services | 20% TEA in water | 2019 | 17 | 0 | 17 | 0.93 | 0
Tayside Scientific Services | 20% TEA in water | 2019 | 1 | 0 | 1 | 0.8 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2019 | 3 | 0 | 3 | 0.8 | 0
Number of Studies Included |  |  | 105 | 73 | 178
7 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
17 studies were updated due to ratified data becoming available since the March 2010 release.
1 study was not included due to unrepresentative low concentrations.
Diffusion Tube Bias Adjustment Factors 03/20 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/20) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2019 | 6 | 0.81
Edinburgh Scientific Services | 50% TEA in acetone | 2019 | 1 | 0.87
Glasgow Scientific Services | 20% TEA in water | 2019 | 4 | 0.86
Gradko | 20% TEA in water | 2019 | 27 | 0.93
Gradko | 50% TEA in acetone | 2019 | 8 | 0.87
Lambeth Scientific Services | 50% TEA in acetone | 2019 | 1 | 0.85
Milton Keynes Council | 20% TEA in water | 2019 | 2 | 0.84
SOCOTEC Didcot | 20% TEA in water | 2019 | 6 | 0.76
SOCOTEC Didcot | 50% TEA in acetone | 2019 | 24 | 0.75
SOCOTEC Glasgow | 20% TEA in water | 2019 | 1 | 0.79
SOCOTEC Glasgow | 50% TEA in acetone | 2019 | 1 | 0.75
Somerset County Council | 20% TEA in water | 2019 | 2 | 0.78
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2019 | 1 | 0.78
Staffordshire Scientific Services | 20% TEA in water | 2019 | 17 | 0.93
Tayside Scientific Services | 20% TEA in water | 2019 | 1 | 0.8
West Yorkshire Analytical Services | 50% TEA in acetone | 2019 | 3 | 0.8
Number of Studies Included |  |  | 105
1 study was not included due to unrepresentative low concentrations.
Diffusion Tube Bias Adjustment Factors 09/19 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/19) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2018 | 7 | 0 | 7 | 0.82 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | 2 | 5 | 7 | 0.9 | -0.06
Glasgow Scientific Services | 20% TEA in water | 2018 | 10 | 0 | 10 | 0.9 | 0.01
Gradko | 20% TEA in water | 2018 | 37 | 3 | 40 | 0.92 | 0
Gradko | 50% TEA in acetone | 2018 | 18 | 0 | 18 | 0.89 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2018 | 7 | 0 | 7 | 1.04 | 0.01
Milton Keynes Council | 20% TEA in water | 2018 | 4 | 0 | 4 | 0.76 | -0.01
SOCOTEC Didcot | 20% TEA in water | 2018 | 3 | 1 | 4 | 0.75 | -0.02
SOCOTEC Didcot | 50% TEA in acetone | 2018 | 28 | 2 | 30 | 0.77 | 0.01
SOCOTEC Glasgow | 20% TEA in water | 2018 | 1 | 0 | 1 | 0.93 | 0
SOCOTEC Glasgow | 50% TEA in acetone | 2018 | 1 | 0 | 1 | 0.97 | 0
Somerset County Council | 20% TEA in water | 2018 | 3 | 6 | 9 | 0.91 | 0.02
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | 5 | 1 | 6 | 0.94 | -0.01
Staffordshire Scientific Services | 20% TEA in water | 2018 | 15 | 1 | 16 | 0.88 | -0.01
Tayside Scientific Services | 20% TEA in water | 2018 | 5 | 0 | 5 | 0.8 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | 8 | 0 | 8 | 0.8 | 0
Number of Studies Included |  |  | 154 | 19 | 173
5 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
1 study was not included due to unrepresentative low concentrations.
4 studies were updated due to ratified data becoming available since the June 2019 release.
Diffusion Tube Bias Adjustment Factors 06/19 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/19) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2018 | 7 | 0 | 7 | 0.82 | 0.01
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | 2 | 0 | 2 | 0.96 | 0
Glasgow Scientific Services | 20% TEA in water | 2018 | 9 | 1 | 10 | 0.89 | 0.03
Gradko | 20% TEA in water | 2018 | 30 | 7 | 37 | 0.92 | -0.01
Gradko | 50% TEA in acetone | 2018 | 8 | 10 | 18 | 0.89 | -0.03
Lambeth Scientific Services | 50% TEA in acetone | 2018 | 7 | 0 | 7 | 1.03 | 0
Milton Keynes Council | 20% TEA in water | 2018 | 4 | 0 | 4 | 0.77 | 0
SOCOTEC Didcot | 20% TEA in water | 2018 | 2 | 1 | 3 | 0.77 | 0.03
SOCOTEC Didcot | 50% TEA in acetone | 2018 | 21 | 7 | 28 | 0.76 | 0
SOCOTEC Glasgow | 20% TEA in water | 2018 | 1 | 0 | 1 | 0.93 | -0.02
SOCOTEC Glasgow | 50% TEA in acetone | 2018 | 1 | 0 | 1 | 0.97 | -0.01
Somerset County Council | 20% TEA in water | 2018 | 3 | 0 | 3 | 0.89 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | 4 | 1 | 5 | 0.95 | 0
Staffordshire Scientific Services | 20% TEA in water | 2018 | 13 | 2 | 15 | 0.89 | 0.02
Tayside Scientific Services | 20% TEA in water | 2018 | 5 | 0 | 5 | 0.8 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | 8 | 0 | 8 | 0.8 | 0
Number of Studies Included |  |  | 125 | 29 | 154
4 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
1 study was not included due to unrepresentative low concentrations.
16 studies were updated due to ratified data becoming available since the March 2019 release.
Diffusion Tube Bias Adjustment Factors 03/19 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/19) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2018 | 7 | 0.81
Edinburgh Scientific Services | 50% TEA in acetone | 2018 | 2 | 0.96
Glasgow Scientific Services | 20% TEA in water | 2018 | 9 | 0.86
Gradko | 20% TEA in water | 2018 | 30 | 0.93
Gradko | 50% TEA in acetone | 2018 | 8 | 0.92
Lambeth Scientific Services | 50% TEA in acetone | 2018 | 7 | 1.03
Milton Keynes Council | 20% TEA in water | 2018 | 4 | 0.77
SOCOTEC Didcot | 20% TEA in water | 2018 | 2 | 0.74
SOCOTEC Didcot | 50% TEA in acetone | 2018 | 21 | 0.76
SOCOTEC Glasgow | 20% TEA in water | 2018 | 1 | 0.95
SOCOTEC Glasgow | 50% TEA in acetone | 2018 | 1 | 0.98
Somerset County Council | 20% TEA in water | 2018 | 3 | 0.89
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2018 | 4 | 0.95
Staffordshire Scientific Services | 20% TEA in water | 2018 | 13 | 0.87
Tayside Scientific Services | 20% TEA in water | 2018 | 5 | 0.8
West Yorkshire Analytical Services | 50% TEA in acetone | 2018 | 8 | 0.8
Number of Studies Included |  |  | 125
4 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
1 study was not included due to unrepresentative low concentrations
Some Councils were unable to supply their studies as automatic data processing was incomplete by 19th March 2018.
Diffusion Tube Bias Adjustment Factors 09/18 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/18) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2017 | 7 | 0 | 7 | 0.79 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | 6 | 0 | 6 | 0.81 | 0
ESG Didcot | 20% TEA in water | 2017 | 5 | 0 | 5 | 0.74 | 0
ESG Didcot | 50% TEA in acetone | 2017 | 29 | 1 | 30 | 0.77 | 0
ESG Glasgow | 20% TEA in water | 2017 | 1 | 0 | 1 | 0.8 | 0
ESG Glasgow | 50% TEA in acetone | 2017 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2017 | 10 | 0 | 10 | 0.9 | 0
Gradko | 20% TEA in water | 2017 | 39 | 0 | 39 | 0.87 | 0
Gradko | 50% TEA in acetone | 2017 | 25 | 0 | 25 | 0.96 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2017 | 5 | 0 | 5 | 0.93 | 0
Milton Keynes Council | 20% TEA in water | 2017 | 4 | 0 | 4 | 0.76 | 0
Somerset County Council | 20% TEA in water | 2017 | 2 | 0 | 2 | 0.77 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2017 | 2 | 1 | 3 | 0.87 | -0.01
Staffordshire Scientific Services | 20% TEA in water | 2017 | 19 | 0 | 19 | 0.88 | 0
Tayside Scientific Services | 20% TEA in water | 2017 | 5 | 0 | 5 | 0.72 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | 9 | 0 | 9 | 0.77 | 0
Number of Studies Included |  |  | 169 | 2 | 171
A duplicate study was accidently included in the 06/19 addition for Kingston upon Hull City Council (ESG Didcot, 50% TEA in acetone), this has been deleted.
2 new ESG Didcot, 50% TEA in acetone studies were received between the June and September update.
1 study was updated due to ratified data becoming available since June 2018.
13 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 06/18 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/18) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2017 | 7 | 0 | 7 | 0.79 | 0.01
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | 2 | 4 | 6 | 0.81 | -0.08
ESG Didcot | 20% TEA in water | 2017 | 2 | 3 | 5 | 0.74 | 0.03
ESG Didcot | 50% TEA in acetone | 2017 | 27 | 2 | 29 | 0.77 | 0
ESG Glasgow | 20% TEA in water | 2017 | 1 | 0 | 1 | 0.8 | 0
ESG Glasgow | 50% TEA in acetone | 2017 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2017 | 6 | 4 | 10 | 0.9 | -0.01
Gradko | 20% TEA in water | 2017 | 34 | 5 | 39 | 0.87 | -0.02
Gradko | 50% TEA in acetone | 2017 | 22 | 3 | 25 | 0.96 | -0.01
Lambeth Scientific Services | 50% TEA in acetone | 2017 | 1 | 4 | 5 | 0.93 | 0.03
Milton Keynes Council | 20% TEA in water | 2017 | 1 | 3 | 4 | 0.76 | -0.13
Somerset County Council | 20% TEA in water | 2017 | 2 | 0 | 2 | 0.77 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2017 | 2 | 0 | 2 | 0.88 | 0
Staffordshire Scientific Services | 20% TEA in water | 2017 | 14 | 5 | 19 | 0.88 | -0.01
Tayside Scientific Services | 20% TEA in water | 2017 | 5 | 0 | 5 | 0.72 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | 4 | 5 | 9 | 0.77 | -0.01
Number of Studies Included |  |  | 131 | 38 | 169
9 studies were updated due to ratified data becoming available since the March 2018.
12 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 03/18 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/18) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2017 | 7 | 0.78
Edinburgh Scientific Services | 50% TEA in acetone | 2017 | 2 | 0.89
ESG Didcot | 20% TEA in water | 2017 | 2 | 0.71
ESG Didcot | 50% TEA in acetone | 2017 | 27 | 0.77
ESG Glasgow | 20% TEA in water | 2017 | 1 | 0.8
ESG Glasgow | 50% TEA in acetone | 2017 | 1 | 0.78
Glasgow Scientific Services | 20% TEA in water | 2017 | 6 | 0.91
Gradko | 20% TEA in water | 2017 | 34 | 0.89
Gradko | 50% TEA in acetone | 2017 | 22 | 0.97
Lambeth Scientific Services | 50% TEA in acetone | 2017 | 1 | 0.9
Milton Keynes Council | 20% TEA in water | 2017 | 1 | 0.89
Somerset County Council | 20% TEA in water | 2017 | 2 | 0.77
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2017 | 2 | 0.88
Staffordshire Scientific Services | 20% TEA in water | 2017 | 14 | 0.89
Tayside Scientific Services | 20% TEA in water | 2017 | 5 | 0.72
West Yorkshire Analytical Services | 50% TEA in acetone | 2017 | 4 | 0.78
Number of Studies Included |  |  | 131
6 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
1 study was not included due to unrepresentative low concentrations
Some Councils were unable to supply their studies as automatic data processing was incomplete by 22nd March 2018.
Diffusion Tube Bias Adjustment Factors 09/17 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/17) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2016 | 7 | 0 | 7 | 0.85 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | 6 | 0 | 6 | 0.77 | 0
ESG Didcot | 20% TEA in water | 2016 | 3 | 0 | 3 | 0.83 | 0
ESG Didcot | 50% TEA in acetone | 2016 | 38 | 0 | 38 | 0.77 | -0.01
ESG Glasgow | 20% TEA in water | 2016 | 1 | 0 | 1 | 0.79 | 0
ESG Glasgow | 50% TEA in acetone | 2016 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2016 | 9 | 0 | 9 | 0.97 | 0
Gradko | 20% TEA in water | 2016 | 27 | 5 | 32 | 0.92 | 0
Gradko | 50% TEA in acetone | 2016 | 18 | 1 | 19 | 1.01 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2016 | 5 | 1 | 6 | 1.06 | 0.02
Milton Keynes Council | 20% TEA in water | 2016 | 1 | 2 | 3 | 0.66 | -0.08
Northampton BC | 20% TEA in water | 2016 | 3 | 0 | 3 | 0.85 | 0
Somerset County Council | 20% TEA in water | 2016 | 3 | 0 | 3 | 0.88 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2016 | 2 | 0 | 2 | 0.83 | 0
Staffordshire Scientific Services | 20% TEA in water | 2016 | 12 | 1 | 13 | 0.91 | 0
Tayside Scientific Services | 20% TEA in water | 2016 | 5 | 0 | 5 | 0.82 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | 7 | 0 | 7 | 0.75 | 0
Number of Studies Included |  |  | 148 | 10 | 158
4 studies were updated due to ratified data becoming available since the June 2017 update.
19 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 06/17 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/17) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2016 | 1 | 6 | 7 | 0.85 | -0.01
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | 1 | 5 | 6 | 0.77 | -0.1
ESG Didcot | 20% TEA in water | 2016 | 2 | 1 | 3 | 0.83 | 0.08
ESG Didcot | 50% TEA in acetone | 2016 | 30 | 8 | 38 | 0.78 | 0.01
ESG Glasgow | 20% TEA in water | 2016 | 1 | 0 | 1 | 0.79 | 0
ESG Glasgow | 50% TEA in acetone | 2016 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2016 | 9 | 0 | 9 | 0.97 | 0
Gradko | 20% TEA in water | 2016 | 21 | 6 | 27 | 0.92 | -0.02
Gradko | 50% TEA in acetone | 2016 | 16 | 2 | 18 | 1.01 | -0.02
Lambeth Scientific Services | 50% TEA in acetone | 2016 | 1 | 4 | 5 | 1.04 | 0.1
Milton Keynes Council | 20% TEA in water | 2016 | 1 | 0 | 1 | 0.74 | 0
Northampton BC | 20% TEA in water | 2016 | 3 | 0 | 3 | 0.85 | 0
Somerset County Council | 20% TEA in water | 2016 | 3 | 0 | 3 | 0.88 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2016 | 2 | 0 | 2 | 0.83 | 0
Staffordshire Scientific Services | 20% TEA in water | 2016 | 11 | 1 | 12 | 0.91 | 0.03
Tayside Scientific Services | 20% TEA in water | 2016 | 1 | 4 | 5 | 0.82 | 0.05
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | 7 | 0 | 7 | 0.75 | 0
Number of Studies Included |  |  | 111 | 37 | 148
8 studies were updated due to ratified data becoming available since the March 2017 update.
17 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 03/17 V2 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (03/17 V2) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2016 | 1 | 0 | 1 | 0.86 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | 1 | 0 | 1 | 0.87 | 0
ESG Didcot | 20% TEA in water | 2016 | 2 | 0 | 2 | 0.75 | 0
ESG Didcot | 50% TEA in acetone | 2016 | 30 | 0 | 30 | 0.77 | 0
ESG Glasgow | 20% TEA in water | 2016 | 1 | 0 | 1 | 0.79 | 0
ESG Glasgow | 50% TEA in acetone | 2016 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2016 | 9 | 0 | 9 | 0.97 | 0
Gradko | 20% TEA in water | 2016 | 22 | -1 | 21 | 0.94 | -0.03
Gradko | 50% TEA in acetone | 2016 | 16 | 0 | 16 | 1.03 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2016 | 1 | 0 | 1 | 0.94 | 0
Milton Keynes Council | 20% TEA in water | 2016 | 1 | 0 | 1 | 0.74 | 0
Northampton BC | 20% TEA in water | 2016 | 3 | 0 | 3 | 0.85 | 0
Somerset County Council | 20% TEA in water | 2016 | 3 | 0 | 3 | 0.88 | 0
Somerset County Council | 50% TEA in acetone | 2016 | 1 | 0 | 1 | 0.77 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2016 | 2 | 0 | 2 | 0.83 | 0
Staffordshire Scientific Services | 20% TEA in water | 2016 | 11 | 0 | 11 | 0.88 | 0
Tayside Scientific Services | 20% TEA in water | 2016 | 1 | 0 | 1 | 0.77 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | 7 | 0 | 7 | 0.75 | 0
Number of Studies Included |  |  | 113 | -1 | 112
1 summary was removed from the Gradko 20% TEA in water data set due to possible sampling artefacts
9 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 03/17 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/17) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2016 | 1 | 0.86
Edinburgh Scientific Services | 50% TEA in acetone | 2016 | 1 | 0.87
ESG Didcot | 20% TEA in water | 2016 | 2 | 0.75
ESG Didcot | 50% TEA in acetone | 2016 | 30 | 0.77
ESG Glasgow | 20% TEA in water | 2016 | 1 | 0.79
ESG Glasgow | 50% TEA in acetone | 2016 | 1 | 0.78
Glasgow Scientific Services | 20% TEA in water | 2016 | 9 | 0.97
Gradko | 20% TEA in water | 2016 | 22 | 0.97
Gradko | 50% TEA in acetone | 2016 | 16 | 1.03
Lambeth Scientific Services | 50% TEA in acetone | 2016 | 1 | 0.94
Milton Keynes Council | 20% TEA in water | 2016 | 1 | 0.74
Northampton BC | 20% TEA in water | 2016 | 3 | 0.85
Somerset County Council | 20% TEA in water | 2016 | 3 | 0.88
Somerset County Council | 50% TEA in acetone | 2016 | 1 | 0.77
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2016 | 2 | 0.83
Staffordshire Scientific Services | 20% TEA in water | 2016 | 11 | 0.88
Tayside Scientific Services | 20% TEA in water | 2016 | 1 | 0.77
West Yorkshire Analytical Services | 50% TEA in acetone | 2016 | 7 | 0.75
Number of Studies Included |  |  | 113
9 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Some Councils were unable to supply their studies as automatic data processing was incomplete by 17th March 2017.
Diffusion Tube Bias Adjustment Factors 09/16 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/16) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2015 | 6 | 0 | 6 | 0.84 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | 9 | 0 | 9 | 0.76 | 0
ESG Didcot | 20% TEA in water | 2015 | 3 | 0 | 3 | 0.81 | 0
ESG Didcot | 50% TEA in acetone | 2015 | 26 | 3 | 29 | 0.79 | 0
ESG Glasgow | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.77 | 0
ESG Glasgow | 50% TEA in acetone | 2015 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2015 | 11 | 1 | 12 | 0.98 | 0
Gradko | 20% TEA in water | 2015 | 27 | 3 | 30 | 0.87 | -0.01
Gradko | 50% TEA in acetone | 2015 | 15 | 0 | 15 | 0.96 | 0
Kirklees Council | 50% TEA in acetone | 2015 | 3 | 0 | 3 | 0.76 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2015 | 6 | 1 | 7 | 1 | 0
Milton Keynes Council | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.76 | 0
Northampton BC | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.66 | 0
Somerset County Council | 20% TEA in water | 2015 | 10 | 0 | 10 | 0.9 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 | 2 | 0 | 2 | 0.84 | 0
Staffordshire Scientific Services | 20% TEA in water | 2015 | 16 | 0 | 16 | 0.85 | 0
Tayside Scientific Services | 20% TEA in water | 2015 | 5 | 5 | 10 | 0.79 | 0.02
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | 6 | 0 | 6 | 0.77 | 0
Number of Studies Included |  |  | 149 | 13 | 162
1 study was updated due to ratified data becoming available since the June 2016 revision
16 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 06/16 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/16) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2015 | 1 | 5 | 6 | 0.84 | 0.03
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | 3 | 6 | 9 | 0.76 | -0.05
ESG Didcot | 20% TEA in water | 2015 | 3 | 0 | 3 | 0.81 | 0
ESG Didcot | 50% TEA in acetone | 2015 | 26 | 0 | 26 | 0.79 | 0
ESG Glasgow | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.77 | 0
ESG Glasgow | 50% TEA in acetone | 2015 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2015 | 10 | 1 | 11 | 0.98 | 0
Gradko | 20% TEA in water | 2015 | 26 | 1 | 27 | 0.88 | -0.02
Gradko | 50% TEA in acetone | 2015 | 14 | 1 | 15 | 0.96 | 0.01
Kirklees Council | 50% TEA in acetone | 2015 | 3 | 0 | 3 | 0.76 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2015 | 3 | 3 | 6 | 1 | -0.07
Milton Keynes Council | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.76 | 0
Northampton BC | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.66 | 0
Somerset County Council | 20% TEA in water | 2015 | 10 | 0 | 10 | 0.9 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 | 2 | 0 | 2 | 0.84 | 0
Staffordshire Scientific Services | 20% TEA in water | 2015 | 16 | 0 | 16 | 0.85 | 0
Tayside Scientific Services | 20% TEA in water | 2015 | 5 | 0 | 5 | 0.77 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | 5 | 1 | 6 | 0.77 | 0.01
Number of Studies Included |  |  | 131 | 18 | 149
15 studies were updated due to ratified data becoming available since the April 2016 revision
15 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 04/16 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (04/16) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.81 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | 3 | 0 | 3 | 0.81 | 0
ESG Didcot | 20% TEA in water | 2015 | 3 | 0 | 3 | 0.81 | 0
ESG Didcot | 50% TEA in acetone | 2015 | 22 | 4 | 26 | 0.79 | -0.02
ESG Glasgow | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.77 | 0
ESG Glasgow | 50% TEA in acetone | 2015 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2015 | 10 | 0 | 10 | 0.98 | 0
Gradko | 20% TEA in water | 2015 | 25 | 1 | 26 | 0.9 | 0
Gradko | 50% TEA in acetone | 2015 | 15 | -1 | 14 | 0.95 | 0
Kirklees Council | 50% TEA in acetone | 2015 | 3 | 0 | 3 | 0.76 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2015 | 2 | 1 | 3 | 1.07 | 0
Milton Keynes Council | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.76 | 0
Northampton BC | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.66 | 0
Somerset County Council | 20% TEA in water | 2015 | 10 | 0 | 10 | 0.9 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 | 2 | 0 | 2 | 0.84 | 0
Staffordshire Scientific Services | 20% TEA in water | 2015 | 15 | 1 | 16 | 0.85 | 0.01
Tayside Scientific Services | 20% TEA in water | 2015 | 5 | 0 | 5 | 0.77 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | 5 | 0 | 5 | 0.76 | 0
Number of Studies Included |  |  | 125 | 6 | 131
1 study had previously been incorrectly attributed to Gradko 50% TEA in acetone, whereas they were actually analysed by Lambeth Scientific Services
14 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 03/16 V2 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (03/16 V2) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.81 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | 3 | 0 | 3 | 0.81 | 0
ESG Didcot | 20% TEA in water | 2015 | 3 | 0 | 3 | 0.81 | 0
ESG Didcot | 50% TEA in acetone | 2015 | 21 | 1 | 22 | 0.81 | 0
ESG Glasgow | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.77 | 0
ESG Glasgow | 50% TEA in acetone | 2015 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2015 | 6 | 4 | 10 | 0.98 | 0
Gradko | 20% TEA in water | 2015 | 29 | -4 | 25 | 0.9 | -0.01
Gradko | 50% TEA in acetone | 2015 | 15 | 0 | 15 | 0.95 | 0
Kirklees Council | 50% TEA in acetone | 2015 | 3 | 0 | 3 | 0.76 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2015 | 2 | 0 | 2 | 1.07 | 0
Milton Keynes Council | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.76 | 0
Northampton BC | 20% TEA in water | 2015 | 1 | 0 | 1 | 0.66 | 0
Somerset County Council | 20% TEA in water | 2015 | 3 | 7 | 10 | 0.9 | 0.03
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 | 2 | 0 | 2 | 0.84 | 0
Staffordshire Scientific Services | 20% TEA in water | 2015 | 15 | 0 | 15 | 0.84 | 0
Tayside Scientific Services | 20% TEA in water | 2015 | 5 | 0 | 5 | 0.77 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | 5 | 0 | 5 | 0.76 | 0
Number of Studies Included |  |  | 117 | 8 | 125
4 studies had previously been incorrectly attributed to Gradko 20% TEA in water, whereas they were actually analysed by Glasgow Scientific Services
14 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Some Councils were unable to supply their studies as automatic data processing was incomplete by 23rd March 2016.
Diffusion Tube Bias Adjustment Factors 03/16 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/16) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2015 | 1 | 0.81
Edinburgh Scientific Services | 50% TEA in acetone | 2015 | 3 | 0.81
ESG Didcot | 20% TEA in water | 2015 | 3 | 0.81
ESG Didcot | 50% TEA in acetone | 2015 | 21 | 0.81
ESG Glasgow | 20% TEA in water | 2015 | 1 | 0.77
ESG Glasgow | 50% TEA in acetone | 2015 | 1 | 0.78
Glasgow Scientific Services | 20% TEA in water | 2015 | 6 | 0.98
Gradko | 20% TEA in water | 2015 | 29 | 0.91
Gradko | 50% TEA in acetone | 2015 | 15 | 0.95
Kirklees Council | 50% TEA in acetone | 2015 | 3 | 0.76
Lambeth Scientific Services | 50% TEA in acetone | 2015 | 2 | 1.07
Milton Keynes Council | 20% TEA in water | 2015 | 1 | 0.76
Northampton BC | 20% TEA in water | 2015 | 1 | 0.66
Somerset County Council | 20% TEA in water | 2015 | 3 | 0.87
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2015 | 2 | 0.84
Staffordshire Scientific Services | 20% TEA in water | 2015 | 15 | 0.84
Tayside Scientific Services | 20% TEA in water | 2015 | 5 | 0.77
West Yorkshire Analytical Services | 50% TEA in acetone | 2015 | 5 | 0.76
Number of Studies Included |  |  | 117
14 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Some Councils were unable to supply their studies as automatic data processing was incomplete by 10th March 2016.
Diffusion Tube Bias Adjustment Factors 09/15 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/15) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2014 | 7 | 0 | 7 | 0.86 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | 7 | 0 | 7 | 0.74 | 0
ESG Didcot | 20% TEA in water | 2014 | 4 | 0 | 4 | 0.79 | 0
ESG Didcot | 50% TEA in acetone | 2014 | 30 | 1 | 31 | 0.81 | 0
ESG Glasgow | 20% TEA in water | 2014 | 1 | 0 | 1 | 0.7 | 0
ESG Glasgow | 50% TEA in acetone | 2014 | 1 | 0 | 1 | 0.76 | 0
Glasgow Scientific Services | 20% TEA in water | 2014 | 2 | 0 | 2 | 0.83 | 0
Gradko | 20% TEA in water | 2014 | 21 | 1 | 22 | 0.92 | 0.01
Gradko | 50% TEA in acetone | 2014 | 10 | 0 | 10 | 0.98 | 0
Kirklees Council | 50% TEA in acetone | 2014 | 1 | 0 | 1 | 0.74 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2014 | 2 | 3 | 5 | 0.88 | -0.01
Milton Keynes Council | 20% TEA in water | 2014 | 1 | 0 | 1 | 0.75 | 0
Northampton BC | 20% TEA in water | 2014 | 3 | 0 | 3 | 0.77 | 0
Somerset County Council | 20% TEA in water | 2014 | 8 | 0 | 8 | 0.89 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2014 | 1 | 0 | 1 | 0.67 | 0
Staffordshire Scientific Services | 20% TEA in water | 2014 | 15 | 0 | 15 | 0.83 | 0
Tayside Scientific Services | 20% TEA in water | 2014 | 10 | 0 | 10 | 0.79 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | 8 | 0 | 8 | 0.75 | 0
Number of Studies Included |  |  | 132 | 5 | 137
6 studies were updated due to ratified data becoming available since the June 2015 revision
11 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 06/15 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/15) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2014 | 1 | 6 | 7 | 0.86 | 0.11
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | 1 | 6 | 7 | 0.74 | -0.04
ESG Didcot | 20% TEA in water | 2014 | 5 | -1 | 4 | 0.79 | 0
ESG Didcot | 50% TEA in acetone | 2014 | 22 | 8 | 30 | 0.81 | 0
ESG Glasgow | 20% TEA in water | 2014 | 1 | 0 | 1 | 0.7 | 0
ESG Glasgow | 50% TEA in acetone | 2014 | 1 | 0 | 1 | 0.76 | 0
Glasgow Scientific Services | 20% TEA in water | 2014 | 2 | 0 | 2 | 0.83 | 0
Gradko | 20% TEA in water | 2014 | 16 | 5 | 21 | 0.91 | 0
Gradko | 50% TEA in acetone | 2014 | 9 | 1 | 10 | 0.98 | 0.01
Kirklees Council | 50% TEA in acetone | 2014 | 1 | 0 | 1 | 0.74 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2014 | 1 | 1 | 2 | 0.89 | 0.09
Milton Keynes Council | 20% TEA in water | 2014 | 1 | 0 | 1 | 0.75 | 0
Northampton BC | 20% TEA in water | 2014 | 3 | 0 | 3 | 0.77 | 0
Somerset County Council | 20% TEA in water | 2014 | 8 | 0 | 8 | 0.89 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2014 | 1 | 0 | 1 | 0.67 | 0
Staffordshire Scientific Services | 20% TEA in water | 2014 | 15 | 0 | 15 | 0.83 | 0
Tayside Scientific Services | 20% TEA in water | 2014 | 5 | 5 | 10 | 0.79 | 0.02
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | 8 | 0 | 8 | 0.75 | 0
Number of Studies Included |  |  | 101 | 31 | 132
11 studies were updated due to ratified data becoming available since the March 2015 revision
11 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Diffusion Tube Bias Adjustment Factors 03/15 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/15) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2014 | 1 | 0.75
Edinburgh Scientific Services | 50% TEA in acetone | 2014 | 1 | 0.78
ESG Didcot | 20% TEA in water | 2014 | 5 | 0.79
ESG Didcot | 50% TEA in acetone | 2014 | 22 | 0.81
ESG Glasgow | 20% TEA in water | 2014 | 1 | 0.7
ESG Glasgow | 50% TEA in acetone | 2014 | 1 | 0.76
Glasgow Scientific Services | 20% TEA in water | 2014 | 2 | 0.83
Gradko | 20% TEA in water | 2014 | 16 | 0.91
Gradko | 50% TEA in acetone | 2014 | 9 | 0.97
Kirklees Council | 50% TEA in acetone | 2014 | 1 | 0.74
Lambeth Scientific Services | 50% TEA in acetone | 2014 | 1 | 0.8
Milton Keynes Council | 20% TEA in water | 2014 | 1 | 0.75
Northampton BC | 20% TEA in water | 2014 | 3 | 0.77
Somerset County Council | 20% TEA in water | 2014 | 8 | 0.89
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2014 | 1 | 0.67
Staffordshire Scientific Services | 20% TEA in water | 2014 | 15 | 0.83
Tayside Scientific Services | 20% TEA in water | 2014 | 5 | 0.77
West Yorkshire Analytical Services | 50% TEA in acetone | 2014 | 8 | 0.75
Number of Studies Included |  |  | 101
8 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Some Councils were unable to supply their studies as automatic data processing was incomplete by 9th March 2015.
Diffusion Tube Bias Adjustment Factors 09/14 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/14) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.83 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | 7 | 0 | 7 | 0.75 | 0
ESG Didcot | 20% TEA in water | 2013 | 3 | 0 | 3 | 0.85 | 0
ESG Didcot | 50% TEA in acetone | 2013 | 37 | 7 | 44 | 0.81 | 0
ESG Glasgow | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.72 | 0
ESG Glasgow | 50% TEA in acetone | 2013 | 1 | 0 | 1 | 0.73 | 0
Exova | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.91 | 0
Glasgow Scientific Services | 20% TEA in water | 2013 | 5 | 0 | 5 | 0.99 | 0
Gradko | 20% TEA in water | 2013 | 35 | 1 | 36 | 0.95 | 0
Gradko | 50% TEA in acetone | 2013 | 19 | 1 | 20 | 1.01 | 0
Kent Scientific Services | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.77 | 0
Kirklees Council | 50% TEA in acetone | 2013 | 2 | 0 | 2 | 0.74 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2013 | 7 | 0 | 7 | 0.87 | 0
Milton Keynes Council | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.84 | 0
Northampton BC | 20% TEA in water | 2013 | 4 | 0 | 4 | 0.73 | 0
Somerset County Council | 20% TEA in water | 2013 | 3 | 0 | 3 | 0.9 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2013 | 3 | 0 | 3 | 0.84 | 0
Staffordshire Scientific Services | 20% TEA in water | 2013 | 11 | 0 | 11 | 0.87 | 0
Tayside Scientific Services | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.78 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | 7 | 0 | 7 | 0.79 | 0
Number of Studies Included |  |  | 150 | 9 | 159
Diffusion Tube Bias Adjustment Factors 06/14 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/14) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.83 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | 1 | 6 | 7 | 0.75 | -0.04
ESG Didcot | 20% TEA in water | 2013 | 2 | 1 | 3 | 0.85 | 0.09
ESG Didcot | 50% TEA in acetone | 2013 | 28 | 9 | 37 | 0.81 | 0.01
ESG Glasgow | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.72 | 0
ESG Glasgow | 50% TEA in acetone | 2013 | 1 | 0 | 1 | 0.73 | 0
Exova | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.91 | 0
Glasgow Scientific Services | 20% TEA in water | 2013 | 5 | 0 | 5 | 0.99 | 0
Gradko | 20% TEA in water | 2013 | 24 | 11 | 35 | 0.95 | 0
Gradko | 50% TEA in acetone | 2013 | 17 | 2 | 19 | 1.01 | 0.01
Kent Scientific Services | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.77 | 0
Kirklees Council | 50% TEA in acetone | 2013 | 2 | 0 | 2 | 0.74 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2013 | 1 | 6 | 7 | 0.87 | 0.04
Milton Keynes Council | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.84 | 0
Northampton BC | 20% TEA in water | 2013 | 4 | 0 | 4 | 0.73 | 0
Somerset County Council | 20% TEA in water | 2013 | 3 | 0 | 3 | 0.9 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2013 | 3 | 0 | 3 | 0.84 | 0
Staffordshire Scientific Services | 20% TEA in water | 2013 | 11 | 0 | 11 | 0.87 | 0
Tayside Scientific Services | 20% TEA in water | 2013 | 1 | 0 | 1 | 0.78 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | 7 | 0 | 7 | 0.79 | 0
Number of Studies Included |  |  | 115 | 35 | 150
Diffusion Tube Bias Adjustment Factors 03/14 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/14) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2013 | 1 | 0.83
Edinburgh Scientific Services | 50% TEA in acetone | 2013 | 1 | 0.79
ESG Didcot | 20% TEA in water | 2013 | 2 | 0.76
ESG Didcot | 50% TEA in acetone | 2013 | 28 | 0.8
ESG Glasgow | 20% TEA in water | 2013 | 1 | 0.72
ESG Glasgow | 50% TEA in acetone | 2013 | 1 | 0.73
Exova | 20% TEA in water | 2013 | 1 | 0.91
Glasgow Scientific Services | 20% TEA in water | 2013 | 5 | 0.99
Gradko | 20% TEA in water | 2013 | 24 | 0.95
Gradko | 50% TEA in acetone | 2013 | 17 | 1
Kent Scientific Services | 20% TEA in water | 2013 | 1 | 0.77
Kirklees Council | 50% TEA in acetone | 2013 | 2 | 0.74
Lambeth Scientific Services | 50% TEA in acetone | 2013 | 1 | 0.83
Milton Keynes Council | 20% TEA in water | 2013 | 1 | 0.84
Northampton BC | 20% TEA in water | 2013 | 4 | 0.73
Somerset County Council | 20% TEA in water | 2013 | 3 | 0.9
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2013 | 3 | 0.84
Staffordshire Scientific Services | 20% TEA in water | 2013 | 11 | 0.87
Tayside Scientific Services | 20% TEA in water | 2013 | 1 | 0.78
West Yorkshire Analytical Services | 50% TEA in acetone | 2013 | 7 | 0.79
Number of Studies Included |  |  | 115
7 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Some Councils were unable to supply their studies as automatic data processing was incomplete by 14th March 2014.
Diffusion Tube Bias Adjustment Factors 09/13 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/13) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.83 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | 7 | 0 | 7 | 0.78 | 0
ESG Didcot | 20% TEA in water | 2012 | 2 | 0 | 2 | 0.69 | 0
ESG Didcot | 50% TEA in acetone | 2012 | 38 | 0 | 38 | 0.79 | 0
ESG Glasgow | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.71 | 0
ESG Glasgow | 50% TEA in acetone | 2012 | 4 | 0 | 4 | 0.82 | 0
Exova | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.89 | 0
Glasgow Scientific Services | 20% TEA in water | 2012 | 11 | 0 | 11 | 0.96 | 0
Gradko | 20% TEA in water | 2012 | 34 | 1 | 35 | 0.96 | -0.01
Gradko | 50% TEA in acetone | 2012 | 21 | 0 | 21 | 1.02 | 0
Kent Scientific Services | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.82 | 0
Kirklees Council | 50% TEA in acetone | 2012 | 5 | 0 | 5 | 0.8 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2012 | 3 | 0 | 3 | 0.87 | 0
Milton Keynes Council | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.81 | 0
Northampton BC | 20% TEA in water | 2012 | 3 | 0 | 3 | 0.75 | 0
Somerset County Council | 20% TEA in water | 2012 | 2 | 0 | 2 | 0.95 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2012 | 3 | 0 | 3 | 0.8 | 0
Staffordshire Scientific Services | 20% TEA in water | 2012 | 13 | 0 | 13 | 0.86 | 0
Tayside Scientific Services | 20% TEA in water | 2012 | 7 | 0 | 7 | 0.84 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | 10 | 0 | 10 | 0.78 | 0
Number of Studies Included |  |  | 168 | 1 | 169
1 new study and 6 updated studies were added since the 07/13 release
Diffusion Tube Bias Adjustment Factors 07/13 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (07/13) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.83 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | 7 | 0 | 7 | 0.78 | 0
ESG Didcot | 20% TEA in water | 2012 | 2 | 0 | 2 | 0.69 | 0
ESG Didcot | 50% TEA in acetone | 2012 | 38 | 0 | 38 | 0.79 | 0
ESG Glasgow | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.71 | 0
ESG Glasgow | 50% TEA in acetone | 2012 | 4 | 0 | 4 | 0.82 | 0
Exova | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.89 | 0
Glasgow Scientific Services | 20% TEA in water | 2012 | 11 | 0 | 11 | 0.96 | 0
Gradko | 20% TEA in water | 2012 | 40 | -6 | 34 | 0.97 | 0.03
Gradko | 50% TEA in acetone | 2012 | 21 | 0 | 21 | 1.02 | 0
Kent Scientific Services | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.82 | 0
Kirklees Council | 50% TEA in acetone | 2012 | 5 | 0 | 5 | 0.8 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2012 | 3 | 0 | 3 | 0.87 | 0
Milton Keynes Council | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.81 | 0
Northampton BC | 20% TEA in water | 2012 | 3 | 0 | 3 | 0.75 | 0
Somerset County Council | 20% TEA in water | 2012 | 2 | 0 | 2 | 0.95 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2012 | 3 | 0 | 3 | 0.8 | 0
Staffordshire Scientific Services | 20% TEA in water | 2012 | 13 | 0 | 13 | 0.86 | 0
Tayside Scientific Services | 20% TEA in water | 2012 | 1 | 6 | 7 | 0.84 | -0.06
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | 10 | 0 | 10 | 0.78 | 0
Number of Studies Included |  |  | 168 | 0 | 168
In the June 13 version, Dundee City Council's results were mistakenly included under the Gradko results instead of Tayside Scientific Services
Diffusion Tube Bias Adjustment Factors 06/13 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/13) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen Scientific Services | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.83 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | 1 | 6 | 7 | 0.78 | -0.08
ESG Didcot | 20% TEA in water | 2012 | 2 | 0 | 2 | 0.69 | 0
ESG Didcot | 50% TEA in acetone | 2012 | 26 | 12 | 38 | 0.79 | 0
ESG Glasgow | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.71 | 0
ESG Glasgow | 50% TEA in acetone | 2012 | 5 | -1 | 4 | 0.82 | -0.02
Exova | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.89 | 0
Glasgow Scientific Services | 20% TEA in water | 2012 | 10 | 1 | 11 | 0.96 | 0.01
Gradko | 20% TEA in water | 2012 | 27 | 13 | 40 | 0.94 | -0.03
Gradko | 50% TEA in acetone | 2012 | 18 | 3 | 21 | 1.02 | 0.01
Kent Scientific Services | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.82 | 0
Kirklees Council | 50% TEA in acetone | 2012 | 5 | 0 | 5 | 0.8 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2012 | 2 | 1 | 3 | 0.87 | -0.04
Milton Keynes Council | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.81 | 0
Northampton BC | 20% TEA in water | 2012 | 3 | 0 | 3 | 0.75 | 0
Somerset County Council | 20% TEA in water | 2012 | 2 | 0 | 2 | 0.95 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2012 | 3 | 0 | 3 | 0.8 | 0
Staffordshire Scientific Services | 20% TEA in water | 2012 | 13 | 0 | 13 | 0.86 | 0
Tayside Scientific Services | 20% TEA in water | 2012 | 1 | 0 | 1 | 0.9 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | 10 | 0 | 10 | 0.78 | -0.01
Number of Studies Included |  |  | 133 | 35 | 168
In the April 13 version, Wrexham County Borough Council misreported their analysis lab as ESG Glasgow when it should have been ESG Didcot. Hence the -1 in number
Diffusion Tube Bias Adjustment Factors 03/13 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/13) Factor
|  |  | No. of Studies | Factor
Aberdeen Scientific Services | 20% TEA in water | 2012 | 1 | 0.83
Edinburgh Scientific Services | 50% TEA in acetone | 2012 | 1 | 0.86
ESG Didcot | 20% TEA in water | 2012 | 2 | 0.69
ESG Didcot | 50% TEA in acetone | 2012 | 26 | 0.79
ESG Glasgow | 20% TEA in water | 2012 | 1 | 0.71
ESG Glasgow | 50% TEA in acetone | 2012 | 5 | 0.84
Exova | 20% TEA in water | 2012 | 1 | 0.89
Glasgow Scientific Services | 20% TEA in water | 2012 | 10 | 0.95
Gradko | 20% TEA in water | 2012 | 27 | 0.97
Gradko | 50% TEA in acetone | 2012 | 18 | 1.01
Kent Scientific Services | 20% TEA in water | 2012 | 1 | 0.82
Kirklees Council | 50% TEA in acetone | 2012 | 5 | 0.8
Lambeth Scientific Services | 50% TEA in acetone | 2012 | 2 | 0.91
Milton Keynes Council | 20% TEA in water | 2012 | 1 | 0.81
Northampton BC | 20% TEA in water | 2012 | 3 | 0.75
Somerset County Council | 20% TEA in water | 2012 | 2 | 0.95
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2012 | 3 | 0.8
Staffordshire Scientific Services | 20% TEA in water | 2012 | 13 | 0.86
Tayside Scientific Services | 20% TEA in water | 2012 | 1 | 0.9
West Yorkshire Analytical Services | 50% TEA in acetone | 2012 | 10 | 0.79
Number of Studies Included |  |  | 133
13 studies were submitted but could not be included due to less than 9 valid data periods or poor data quality.
Gradko had a bad batch of 20% TEA in Water tubes in July 2013. This period of data has been deleted from all studies using this tube preparation.
Changes to Diffusion Tube Bias Adjustment Factors with 09/12 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/12) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen CC | 20% TEA in water | 2011 | 6 | 0 | 6 | 0.83 | 0
Bristol Scientific Services | 20% TEA in water | 2011 | 10 | 0 | 10 | 0.82 | 0
Cardiff Scientific Services | 50% TEA in acetone | 2011 | 3 | 0 | 3 | 0.86 | 0
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | 8 | 0 | 8 | 0.84 | 0
ESG Glasgow | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.84 | 0
ESG Glasgow | 50% TEA in acetone | 2011 | 1 | 3 | 4 | 0.87 | 0.05
Exova | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2011 | 7 | 0 | 7 | 0.94 | 0
Gradko | 20% TEA in water | 2011 | 40 | 1 | 41 | 0.9 | 0
Gradko | 50% TEA in acetone | 2011 | 24 | 1 | 25 | 0.95 | 0
ESG Didcot | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.77 | 0
ESG Didcot | 50% TEA in acetone | 2011 | 44 | 1 | 45 | 0.83 | 0
Kent Scientific Services | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.77 | 0
Kirklees Council | 50% TEA in acetone | 2011 | 5 | 0 | 5 | 0.75 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2011 | 6 | 0 | 6 | 1.06 | 0
Milton Keynes Council | 20% TEA in water | 2011 | 1 | 3 | 4 | 0.84 | 0.02
Northampton BC | 20% TEA in water | 2011 | 3 | 0 | 3 | 0.71 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | 6 | 0 | 6 | 0.87 | 0
Staffordshire Scientific Services | 20% TEA in water | 2011 | 15 | 0 | 15 | 0.89 | 0
Tayside Scientific Services | 20% TEA in water | 2011 | 8 | 0 | 8 | 0.78 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | 13 | 0 | 13 | 0.84 | 0
Number of Studies Included |  |  | 204 | 9 | 213
Changes to Diffusion Tube Bias Adjustment Factors with 07/12 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (07/12) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen CC | 20% TEA in water | 2011 | 6 | 0 | 6 | 0.83 | 0
Bristol Scientific Services | 20% TEA in water | 2011 | 10 | 0 | 10 | 0.82 | 0
Cardiff Scientific Services | 50% TEA in acetone | 2011 | 3 | 0 | 3 | 0.86 | 0.04
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | 8 | 0 | 8 | 0.84 | 0
ESG Glasgow | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.84 | 0
ESG Glasgow | 50% TEA in acetone | 2011 | 2 | -1 | 1 | 0.82 | -0.16
Exova | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.78 | 0
Glasgow Scientific Services | 20% TEA in water | 2011 | 7 | 0 | 7 | 0.94 | 0
Gradko | 20% TEA in water | 2011 | 38 | 2 | 40 | 0.9 | -0.01
Gradko | 50% TEA in acetone | 2011 | 23 | 1 | 24 | 0.95 | 0.01
ESG Didcot | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.77 | 0
ESG Didcot | 50% TEA in acetone | 2011 | 44 | 0 | 44 | 0.83 | 0
Kent Scientific Services | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.77 | 0
Kirklees Council | 50% TEA in acetone | 2011 | 5 | 0 | 5 | 0.75 | 0
Lambeth Scientific Services | 50% TEA in acetone | 2011 | 6 | 0 | 6 | 1.06 | 0
Milton Keynes Council | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.82 | 0
Northampton BC | 20% TEA in water | 2011 | 3 | 0 | 3 | 0.71 | 0
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | 6 | 0 | 6 | 0.87 | 0
Staffordshire Scientific Services | 20% TEA in water | 2011 | 12 | 3 | 15 | 0.89 | 0.01
Tayside Scientific Services | 20% TEA in water | 2011 | 8 | 0 | 8 | 0.78 | 0
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | 13 | 0 | 13 | 0.84 | 0
Number of Studies Included |  |  | 199 | 5 | 204
LB Newham misreported their analysis lab as ESG Glasgow when it should have been Gradko. Hence the -1 in number of studies added column for ESG Glasgow
Changes to Diffusion Tube Bias Adjustment Factors with 06/12 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (06/12) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen CC | 20% TEA in water | 2011 | 1 | 5 | 6 | 0.83 | -0.02
Bristol Scientific Services | 20% TEA in water | 2011 | 8 | 2 | 10 | 0.82 | -0.01
Bristol Scientific Services | 50% TEA in acetone | 2011 | 1 | -1 | 0 | N/A | N/A
Cardiff Scientific Services | 50% TEA in acetone | 2011 | 3 | 0 | 3 | 0.82 | -
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | 2 | 6 | 8 | 0.84 | -0.17
ESG Glasgow | 20% TEA in water | 2011 | 5 | -4 | 1 | 0.84 | 0.02
ESG Glasgow | 50% TEA in acetone | 2011 | 22 | -20 | 2 | 0.98 | 0.1
Exova | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.78 | -
Glasgow Scientific Services | 20% TEA in water | 2011 | 7 | 0 | 7 | 0.94 | -
Gradko | 20% TEA in water | 2011 | 25 | 13 | 38 | 0.91 | 0.02
Gradko | 50% TEA in acetone | 2011 | 20 | 3 | 23 | 0.94 | -
ESG Didcot | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.77 | -
ESG Didcot | 50% TEA in acetone | 2011 | 18 | 26 | 44 | 0.83 | -0.01
Kent Scientific Services | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.77 | -
Kirklees Council | 50% TEA in acetone | 2011 | 5 | 0 | 5 | 0.75 | -
Lambeth Scientific Services | 50% TEA in acetone | 2011 | 6 | 0 | 6 | 1.06 | -
Milton Keynes Council | 20% TEA in water | 2011 | 1 | 0 | 1 | 0.82 | -
Northampton BC | 20% TEA in water | 2011 | 3 | 0 | 3 | 0.71 | -
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | 4 | 2 | 6 | 0.87 | 0.08
Staffordshire Scientific Services | 20% TEA in water | 2011 | 11 | 1 | 12 | 0.88 | -
Tayside Scientific Services | 20% TEA in water | 2011 | 8 | 0 | 8 | 0.78 | -
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | 12 | 1 | 13 | 0.84 | -
Number of Studies Included |  |  | 165 | 34 | 199
Note:
The following Councils misreported their analysis laboratory for the 03/12 submission, this accounts for the negative numbers in the "No. Studies Added" column of the above table:
Castlereagh Borough Council | Canterbury City Council
Dumfries and Galloway Council | Down District Council
North Down Borough Council | Horsham District Council
North East Lincolnshire Council | Lisburn City Council
Suffolk Coastal District Council | Medway Council
Tunbridge Wells Borough Council | Rugby Borough Council
Wrexham County Borough Council
Wycombe District Council
Boston Borough Council's submission removed as it was not a collocation study
The following analysis laboratories have changed their names:
Environmental Scientific Groups is now ESG Glasgow
Harwell Scientific Services is now ESG Didcot
Changes to Diffusion Tube Bias Adjustment Factors with 04/12 Issue of the Spreadsheet
Analysis laboratory for Wigan incorrectly submitted. Bias factors for Environmental Scientific Groups and Staffordshire Scientific Services affected.
Analysis Laboratory | Method | Year | Old Bias Factor | New bias Factor
Environmental Scientific Groups | 20% TEA in water | 2011 | 0.88 | 0.87
Staffordshire Scientific Services | 20% TEA in water | 2011 | 0.82 | 0.81
Changes to Diffusion Tube Bias Adjustment Factors with 03/12 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (03/12) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Edinburgh Scientific Services | 50% TEA in Acetone | 2010 | 2 | 4 | 6 | 0.9 | -0.12
Aberdeen CC | 20% TEA in water | 2011 | 0 | 1 | 1 | 0.85
Bristol Scientific Services | 20% TEA in water | 2011 | 0 | 8 | 8 | 0.83
Bristol Scientific Services | 50% TEA in acetone | 2011 | 0 | 1 | 1 | 0.71
Cardiff Scientific Services | 50% TEA in acetone | 2011 | 0 | 3 | 3 | 0.82
Edinburgh Scientific Services | 50% TEA in acetone | 2011 | 0 | 2 | 2 | 1.01
Environmental Scientific Groups | 20% TEA in water | 2011 | 0 | 5 | 5 | 0.82
Environmental Scientific Groups | 50% TEA in acetone | 2011 | 0 | 22 | 22 | 0.88
Exova | 20% TEA in water | 2011 | 0 | 1 | 1 | 0.78
Glasgow Scientific Services | 20% TEA in water | 2011 | 0 | 7 | 7 | 0.94
Gradko | 20% TEA in water | 2011 | 0 | 26 | 26 | 0.89
Gradko | 50% TEA in acetone | 2011 | 0 | 20 | 20 | 0.94
Harwell Scientific Services | 20% TEA in water | 2011 | 0 | 1 | 1 | 0.77
Harwell Scientific Services | 50% TEA in acetone | 2011 | 0 | 18 | 18 | 0.84
Kent Scientific Services | 20% TEA in water | 2011 | 0 | 1 | 1 | 0.77
Kirklees Council | 50% TEA in acetone | 2011 | 0 | 5 | 5 | 0.75
Lambeth Scientific Services | 50% TEA in acetone | 2011 | 0 | 6 | 6 | 1.06
Milton Keynes Council | 20% TEA in water | 2011 | 0 | 1 | 1 | 0.82
Northampton BC | 20% TEA in water | 2011 | 0 | 3 | 3 | 0.71
South Yorkshire Air Quality Samplers | 50% TEA in acetone | 2011 | 0 | 4 | 4 | 0.79
Staffordshire Scientific Services | 20% TEA in water | 2011 | 0 | 11 | 11 | 0.88
Tayside Scientific Services | 20% TEA in water | 2011 | 0 | 8 | 8 | 0.78
West Yorkshire Analytical Services | 50% TEA in acetone | 2011 | 0 | 12 | 12 | 0.84
Number of Studies Included |  |  | 2 | 170 | 172
Changes to Diffusion Tube Bias Adjustment Factors with 09/11 Issue of the Spreadsheet
Laboratory | Method | Year | Previous Number of Studies | New (09/11) Update
|  |  |  | No. Studies Added | Total No. of Studies | Factor | Change in Factor
Aberdeen CC | 20% TEA in Water | 2010 | 1 | 4 | 5 | 0.88 | 0.06
Bristol Scientific Services | 20% TEA in Water | 2010 | 7 | 0 | 7 | 0.85 | -
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | 4 | 0 | 4 | 0.85 | -
Edinburgh Scientific Services | 50% TEA in Acetone | 2010 | 2 | 0 | 2 | 1.02 | -
Environmental Scientific Groups | 20% TEA in Water | 2010 | 10 | 0 | 10 | 0.84 | -
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | 4 | 0 | 4 | 0.75 | -
Glasgow Scientific Services | 20% TEA in Water | 2010 | 6 | 0 | 6 | 1.12 | -
Gradko | 20% TEA in Water | 2010 | 41 | 1 | 42 | 0.85 | -0.07
Gradko | 50% TEA in Acetone | 2010 | 16 | 0 | 16 | 1.03 | -
Harwell Scientific Services | 50% TEA in Acetone | 2010 | 19 | 1 | 20 | 0.85 | 0
Harwell Scientific Services | 20% TEA in Water | 2010 | 1 | 0 | 1 | 0.77 | -
Kent Scientific Services | 20% TEA in Water | 2010 | 1 | 0 | 1 | 0.78 | -
Kirklees Council Scientific Services | 50% TEA in Acetone | 2010 | 1 | 0 | 1 | 0.78 | -
Lambeth Scientific Services | 50% TEA in Acetone | 2010 | 4 | 0 | 4 | 1.06 | -
Lancashire CC | 50% TEA in Acetone | 2010 | 1 | 0 | 1 | 0.9 | -
Milton Keynes Council | 20% TEA in Water | 2010 | 6 | 0 | 6 | 0.84 | -
Northampton BC | 20% TEA in Water | 2010 | 3 | 0 | 3 | 0.73 | -
South Yorkshire Labs | 50% TEA in Acetone | 2010 | 5 | 0 | 5 | 0.88 | -
Staffordshire Scientific Services | 20% TEA in Water | 2010 | 9 | 0 | 9 | 0.85 | -
Tayside SS | 20% TEA in Water | 2010 | 4 | 0 | 4 | 0.79 | -
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | 12 | 0 | 12 | 0.9 | -
South Yorkshire Labs | 50% TEA in Acetone | 2009 | 9 | 0 | 9 | 0.91 | -
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | 5 | 0 | 5 | 0.85 | -
Number of Studies Included |  |  | 171 | 6 | 177
Changes to Diffusion Tube Bias Adjustment Factors with 06/11 Issue of the Spreadsheet
Laboratory | Method | Year | New (06/11) Factor
|  |  | No. of Studies | Factor
Aberdeen CC | 20% TEA in Water | 2010 | 1 | 0.82
Bristol Scientific Services | 20% TEA in Water | 2010 | 7 | 0.85
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | 4 | 0.85
Edinburgh Scientific Services | 50% TEA in Acetone | 2010 | 2 | 1.02
Environmental Scientific Groups | 20% TEA in Water | 2010 | 10 | 0.84
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | 4 | 0.75
Glasgow Scientific Services | 20% TEA in Water | 2010 | 6 | 1.12
Gradko | 20% TEA in Water | 2010 | 41 | 0.92
Gradko | 50% TEA in Acetone | 2010 | 16 | 1.03
Harwell Scientific Services | 50% TEA in Acetone | 2010 | 19 | 0.85
Harwell Scientific Services | 20% TEA in Water | 2010 | 1 | 0.77
Kent Scientific Services | 20% TEA in Water | 2010 | 1 | 0.78
Kirklees Council Scientific Services | 50% TEA in Acetone | 2010 | 1 | 0.78
Lambeth Scientific Services | 50% TEA in Acetone | 2010 | 4 | 1.06
Lancashire CC | 50% TEA in Acetone | 2010 | 1 | 0.9
Milton Keynes Council | 20% TEA in Water | 2010 | 6 | 0.84
Northampton BC | 20% TEA in Water | 2010 | 3 | 0.73
South Yorkshire Labs | 50% TEA in Acetone | 2010 | 5 | 0.88
Staffordshire Scientific Services | 20% TEA in Water | 2010 | 9 | 0.85
Tayside SS | 20% TEA in Water | 2010 | 4 | 0.79
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | 12 | 0.9
South Yorkshire Labs | 50% TEA in Acetone | 2009 | 9 | 0.91
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | 5 | 0.85
Number of Studies Included |  |  | 171
Changes to Diffusion Tube Bias Adjustment Factors with 04/11 Issue of the Spreadsheet
Laboratory | Method | Year | New (04/11) Factor
|  |  | No. of Studies | Factor
Aberdeen CC | 20% TEA in Water | 2010 | 1 | 0.82
Bristol Scientific Services | 20% TEA in Water | 2010 | 7 | 0.85
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | 4 | 0.85
Edinburgh Scientific Services | 50% TEA in Acetone | 2010 | 2 | 1.02
Environmental Scientific Groups | 20% TEA in Water | 2010 | 10 | 0.84
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | 3 | 0.83
Glasgow Scientific Services | 20% TEA in Water | 2010 | 1 | 1.1
Gradko | 20% TEA in Water | 2010 | 39 | 0.92
Gradko | 50% TEA in Acetone | 2010 | 17 | 0.99
Harwell Scientific Services | 50% TEA in Acetone | 2010 | 18 | 0.85
Harwell Scientific Services | 20% TEA in Water | 2010 | 1 | 0.77
Kent Scientific Services | 20% TEA in Water | 2010 | 1 | 0.78
Kirklees Council Scientific Services | 50% TEA in Acetone | 2010 | 1 | 0.78
Lambeth Scientific Services | 50% TEA in Acetone | 2010 | 3 | 1.07
Lancashire CC | 50% TEA in Acetone | 2010 | 1 | 0.9
Milton Keynes Council | 20% TEA in Water | 2010 | 6 | 0.84
Northampton BC | 20% TEA in Water | 2010 | 3 | 0.73
South Yorkshire Labs | 50% TEA in Acetone | 2010 | 5 | 0.88
Staffordshire Scientific Services | 20% TEA in Water | 2010 | 6 | 0.87
Tayside SS | 20% TEA in Water | 2010 | 4 | 0.78
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | 12 | 0.9
Edinburgh Scientific Services | 50% TEA in Acetone | 2009 | 5 | 0.85
Number of Studies Included |  |  | 150
Changes to Diffusion Tube Bias Adjustment Factors with 03/11 Issue of the Spreadsheet
Laboratory | Method | Year | New (03/11) Factor
|  |  | No. of Studies | Factor
Bristol Scientific Services | 20% TEA in Water | 2010 | 4 | 0.77
Cardiff Scientific Services | 50% TEA in Acetone | 2010 | 4 | 0.85
Environmental Scientific Groups | 50% TEA in Acetone | 2010 | 1 | 0.94
Gradko | 20% TEA in Water | 2010 | 9 | 0.95
Gradko | 50% TEA in Acetone | 2010 | 6 | 0.93
Harwell Scientific Services | 50% TEA in Acetone | 2010 | 3 | 0.78
South Yorkshire Labs | 50% TEA in Acetone | 2010 | 1 | 0.92
Staffordshire Scientific Services | 20% TEA in Water | 2010 | 2 | 0.91
West Yorkshire Analytical Services | 50% TEA in Acetone | 2010 | 11 | 0.92
Number of Studies Included |  |  | 41